## Sheet1
WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FIRM_NAME | MANUFACTURER_RECALL_REASON
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190025 | Z-0307-2022 | The Ocean Water Seal Chest Drain is a plastic, three bottle vacuum collection device. It consists of a collection chamber that holds shed mediastinal blood, a water seal chamber using a water seal that prevents air from entering into the pleural space and a suction control chamber, which incorporates a mechanical regulator to control suction pressures imposed upon the patient's pleural cavity. | Ocean Water Seal Chest Drain, Model 2002-400 | 2 | CDRH | 2021-11-25 |  |  | Atrium Medical Corporation | Inadequate/inappropriate Set-Up Instructions, which may result in delay in treatment and an unreasonable risk of harm to the patient
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190131 | Z-0298-2022 | Artis zee multi-purpose with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10094139 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190134 | Z-0301-2022 | Artis Q ceiling with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10848281 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190150 | Z-0291-2022 | Endurant II Stent Graft System REF/Description: ETBF2513C145EE STENT GRAFT ETBF2513C145EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETBF2820C166EE STENT GRAFT ETBF2820C166EE ENDUR II BIF ETCF2828C49EE STENT GRAFT ETCF2828C49EE ENDUR II CUFF ETCF2828C49EE STENT GRAFT ETCF2828C49EE ENDUR II CUFF ETCF2828C49EE STENT GRAFT ETCF2828C49EE ENDUR II CUFF ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETUF2514C102EE STENT GRAFT ETUF2514C102EE ENDUR II AUI ETBF2313C124EE STENT GRAFT ETBF2313C124EE ENDUR II BIF ETBF2513C145EE STENT GRAFT ETBF2513C145EE ENDUR II BIF ETBF2513C145EE STENT GRAFT ETBF2513C145EE ENDUR II BIF ETBF2513C145EE STENT GRAFT ETBF2513C145EE ENDUR II BIF ETBF2513C145EE STENT GRAFT ETBF2513C145EE ENDUR II BIF ETBF2516C166EE STENT GRAFT ETBF2516C166EE ENDUR II BIF ETBF2516C166EE STENT GRAFT ETBF2516C166EE ENDUR II BIF ETBF2516C166EE STENT GRAFT ETBF2516C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2813C166EE STENT GRAFT ETBF2813C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2816C166EE STENT GRAFT ETBF2816C166EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETBF2516C145EE STENT GRAFT ETBF2516C145EE ENDUR II BIF ETUF2514C102E STENT GRAFT ETUF2514C102E ENDUR II AUI ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETTF2828C70EE STENT GRAFT ETTF2828C70EE ENDUR II TUBE ETCF2525C49E STENT GRAFT ETCF2525C49E ENDUR II AOEXT ETCF2828C49E STENT GRAFT ETCF2828C49E ENDUR II AOEXT ETCF2828C49E STENT GRAFT ETCF2828C49E ENDUR II AOEXT ETCF2828C49E STENT GRAFT ETCF2828C49E ENDUR II AOEXT ETUF2814C102E STENT GRAFT ETUF2814C102E ENDUR II AUI ETUF2814C102E STENT GRAFT ETUF2814C102E ENDUR II AUI ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETTF2525C70EE STENT GRAFT ETTF2525C70EE ENDUR II TUBE ETBF2313C166EE STENT GRAFT ETBF2313C166EE ENDUR II BIF ETBF2513C124EE STENT GRAFT ETBF2513C124EE ENDUR II BIF ET | Endurant II Stent Graft System | 2 | CDRH | 2021-11-24 |  |  | Medtronic Vascular, Inc. | Due to, during stent graft deployment, the spindle may detach from the distal end of the spindle hypotube.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190297 | Z-0240-2022 | ProxiDiagnost N90 is multi-functional general R/F systems. | ProxiDiagnost N90 | 2 | CDRH | 2021-11-24 |  |  | Philips Healthcare | The Aleva monitor requires certain labels to be applied, some of which are required by FDA Electronic Product Radiation Control performance standards. This includes the Radiation Warning Label [21CFR 1020.30 (j)] and the X-ray control certification statement [21 CFR 1010.2]. Philips has determined that some systems at customer locations did not have these labels applied.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190291 | Z-0238-2022 | Philips Azurion systems with software releases 2.1(L1) and 2.1(L2) | Azurion R2.1 | 2 | CDRH | 2021-11-24 |  |  | Philips Healthcare | 3D-RA is a reconstruction software product that can be used with the Philips Azurin system. This product extends the functionality of compatible X-ray equipment with the capability of 3D X-ray imaging. Before starting a 3D-RA scan, the detector of the Philips Azurion system must be positioned in portrait or landscape orientation with an angle of 0 or +/-90 degrees relative to the 3D scan direction. Philips has discovered that Azurion releases 2.1(L1) and 2.1(L2) allow creation of a 3D-RA scan with a deviating detector orientation of up to 3 degrees from the exact portrait or landscape orientation (i.e., at 0 or +/-90 degrees), relative to the 3D scan direction. 3D-RA scans with a deviating detector orientation of ore than 1.0 degree from the exact portrait or landscape orientation (i.e., at 0 or +/-90 degrees) cannot be reconstructed with the 3D-RA reconstruction software.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190298 | Z-0241-2022 | CombiDiagnost R90 is multi-functional general R/F systems. | CombiDiagnost R90 | 2 | CDRH | 2021-11-24 |  |  | Philips Healthcare | The Aleva monitor requires certain labels to be applied, some of which are required by FDA Electronic Product Radiation Control performance standards. This includes the Radiation Warning Label [21CFR 1020.30 (j)] and the X-ray control certification statement [21 CFR 1010.2]. Philips has determined that some systems at customer locations did not have these labels applied.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190306 | Z-0242-2022 | Computed Tomography Systems Revolution CT, Revolution CT ES, Revolution Apex, Revolution CT with Apex edition | Revolution CT, Revolution CT ES, Revolution Apex, Revolution CT with Apex edition | 2 | CDRH | 2021-11-24 |  |  | GE Healthcare, LLC | The accumulated dose is incorrectly displayed in certain situations. Thisoccurs during an interventional exam using the SmartStep option if the table height is adjusted after the exam starts. In this situation, the SmartStep display will show a value that is higher (up to 100 times) than the actual dose delivered.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189955 | Z-0289-2022 | SOMATOM Edge Plus with software syngo.CT VB20 Model #10267000 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190133 | Z-0300-2022 | Artis Q floor with software VD12- angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10848280 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190151 | Z-0292-2022 | Endurant IIs Stent Graft System REF/Description: ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EJ STENT GRAFT ESBF2814C103EJ ENDUR IIS BIF ESBF2814C103EJ STENT GRAFT ESBF2814C103EJ ENDUR IIS BIF ESBF2814C103EJ STENT GRAFT ESBF2814C103EJ ENDUR IIS BIF ESBF2314C103E Stent Graft ESBF2314C103E Endur IIs BIF ESBF2314C103E Stent Graft ESBF2314C103E Endur IIs BIF ESBF2314C103E Stent Graft ESBF2314C103E Endur IIs BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2514C103EE STENT GRAFT ESBF2514C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2814C103EE STENT GRAFT ESBF2814C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF2314C103EE ENDUR IIS BIF ESBF2314C103EE STENT GRAFT ESBF231 | Endurant IIs Stent Graft System | 2 | CDRH | 2021-11-24 |  |  | Medtronic Vascular, Inc. | Due to, during stent graft deployment, the spindle may detach from the distal end of the spindle hypotube.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190096 | Z-0294-2022 | Medtronic Wireless Recharger Model WR9200 included in the RS6200 Recharger Kit used by the Activa RC (Model 37612) implantable neurostimulator. | Medtronic Wireless Recharger | 2 | CDRH | 2021-11-24 |  |  | Medtronic Neuromodulation | The wireless rechargers (WR; WR9200 and WR9220; included in the Recharger Kits) become unresponsive if the charging steps, indicated in the Recharger Patient User Guide, are not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190135 | Z-0302-2022 | Artis Q biplane with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10848282 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190137 | Z-0304-2022 | Artis Q.zen ceiling with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10848354 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190138 | Z-0305-2022 | Artis Q.zen biplane with software VD12- angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number:10848355 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190296 | Z-0239-2022 | DigitalDiagnost C90 | DigitalDiagnost C90 | 2 | CDRH | 2021-11-24 |  |  | Philips Healthcare | The Aleva monitor requires certain labels to be applied, some of which are required by FDA Electronic Product Radiation Control performance standards. This includes the Radiation Warning Label [21CFR 1020.30 (j)] and the X-ray control certification statement [21 CFR 1010.2]. Philips has determined that some systems at customer locations did not have these labels applied.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189949 | Z-0285-2022 | SOMATOM Definition Edge with software syngo.CT VB20 Model #10590000 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190089 | Z-0290-2022 | Endurant Stent Graft System REF/Description: ENTF2525C70EE STENT GRAFT ENTF2525C70EE ENDURANT TUBE | Endurant Stent Graft System | 2 | CDRH | 2021-11-24 |  |  | Medtronic Vascular, Inc. | During stent graft deployment, the spindle may detach from the distal end of the spindle hypotube
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190097 | Z-0295-2022 | Medtronic Wireless Recharger Model WR9200 included in the RS7200 Recharger Kit used by the RestoreUltra (Model 37712), RestoreSensor (Model 37714), Restore Ultra SureScan MRI (Model 97712), and RestoreSensor SureScan MRI (Model 97714) implantable neurostimulators. | Medtronic Wireless Recharger | 2 | CDRH | 2021-11-24 |  |  | Medtronic Neuromodulation | The wireless rechargers (WR; WR9200 and WR9220; included in the Recharger Kits) become unresponsive if the charging steps, indicated in the Recharger Patient User Guide, are not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189467 | Z-0306-2022 | BD Nexiva, 20 GA 1.00 IN (1.1 X 25 mm), 3660 ml/hr. 61ml/min; Closed IV Catheter System- Dual port -61 mL/min (3660 mL/hr)-0.5mL, REF 383536; BD Q-Style Luer Access Split Septum BD Vialon Material; BD Instaflash Needle Technology; Rx Only, Sterile EO. BD Nexiva closed IV catheter systems are over-the-needle, intravascular catheters. These devices have a radiopaque BD Vialon" catheter, needle, needle shield, septum, stabilization platform, integrated extension tubing, clamp, dual port adapter, vent plug and pre-attached BD Q-Syte Luer Access Split Septum. The needle and catheter are protected by a needle cover. A BD Q-Syte device or an end cap with protective cover is provided in the unit package. The closed system is designed to keep blood contained within the device throughout the insertion process. The septum is designed to wipe visible blood from the needle surface as the needle is withdrawn from the catheter, further reducing the risk of blood exposure. The needle tip is passively protected when the needle is removed, reducing the risk of accidental needlestick injury. These devices have BD Instaflash" Needle Technology, allowing for immediate visualization of blood along the catheter. Continuous blood return is seen in the extension tubing. The vent plug prevents blood leakage from the extension tubing during insertion. Both the stabilization platform and dual port adapter are color coded to indicate catheter gauge size (24GA (0.7 mm)=Yellow, 22GA (0.9 mm)=Blue, 20GA (1.1 mm)=Pink, 18GA (1.3 mm)=Green). BD Nexiva closed IV catheter systems are intended to be inserted into a patient s vascular system for short term use to sample blood, monitor blood pressure, or administer fluids. These catheters may be used for any patient population with consideration given to adequacy of vascular anatomy, procedure being performed, fluids being infused, and duration of therapy. The 22-18GA (0.9-1.3 mm) devices are suitable for use with power injectors set to a maximum pressure of 300 psi (2068 kPa) when access ports not suitable for use with power injectors are removed. | BD Nexiva 20 GA 1.00 (1.1 x 25mm) Dual Port | 2 | CDRH | 2021-11-24 |  |  | Becton Dickinson Infusion Therapy Systems Inc. | There is a breach in the product packaging that renders the product non-sterile.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189946 | Z-0283-2022 | SOMATOM Force with software syngo.CT VB20 Model #10742326 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189953 | Z-0287-2022 | SOMATOM Drive with software syngo.CT VB20 Model #10431700 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190130 | Z-0297-2022 | Artis zee ceiling with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number:10094137 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189950 | Z-0286-2022 | SOMATOM Definition Flash with software syngo.CT VB20 Model #10430603 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189977 | Z-0282-2022 | Medtronic CareLink SmartSync Device Manager application software (D00U005) used by the following Medtronic Cobalt and Crome implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds): Cobalt XT VR: DVPA2D1, DVPA2D4; Cobalt VR: DVPB3D1, DVPB3D4; Cobalt XT DR: DDPA2D1, DPA2D4; Cobalt DR: DDPB3D1, DDPB3D4; Cobalt XT HF: DTPA2D4, DTPA2D1; Cobalt XT HF; Quad: DTPA2QQ, DTPA2Q1; Cobalt HF: DTPB2D4, DTPB2D1; Cobalt HF Quad: DTPB2QQ, DTPB2Q1. Crome VR: DVPC3D1, DVPC3D4; Crome DR: DDPC3D1, DDPC3D4; Crome HF: DTPC2D4, DTPC2D1; Crome HF Quad: DTPC2QQ, DTPC2Q1. | Medtronic CareLink SmartSync | 2 | CDRH | 2021-11-24 |  |  | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | The processing of collected episode data may fail due to a software error.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190136 | Z-0303-2022 | Artis Q.zen floor with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10848353 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189947 | Z-0284-2022 | SOMATOM Definition AS with software syngo.CT VB20 Model #8098027 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189954 | Z-0288-2022 | SOMATOM Confidence with software syngo.CT VB20 Model #10590100 | Siemens | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | software syngo.CT VB20 in the installed base, with or without syngo.CT software versions VB20_SP1, SP2, SP3, or SP4, may result in workflow interruptions resulting in the delay in diagnosis, the necessity for patient rescans, and/or the need for additional contrast media may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190094 | Z-0293-2022 | Medtronic Wireless Recharger Model WR9220 included in the RS5200 Recharger Kit used by the InterStim Micro (Model 97810) implantable neurostimulator. | Medtronic Wireless Recharger | 2 | CDRH | 2021-11-24 |  |  | Medtronic Neuromodulation | The wireless rechargers (WR; WR9200 and WR9220; included in the Recharger Kits) become unresponsive if the charging steps, indicated in the Recharger Patient User Guide, are not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190129 | Z-0296-2022 | Artis zee floor with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10094135 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190132 | Z-0299-2022 | Artis zee biplane with software VD12-angiography systems developed for single and biplane diagnostic imaging and interventional procedure Model Number: 10094141 | Artis Zee/Zeego & Artis Q/Q.Zen | 2 | CDRH | 2021-11-24 |  |  | Siemens Medical Solutions USA, Inc | If SID (source-to-image distance) lift movement is activated and x-ray shall be released simultaneously; x-ray is not possible and the message No x-ray, try again is displayed, may result in a short delay in procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190363 | Z-0281-2022 | Continuous CT Fluoroscopy DVI Ceiling Mount - CCT Arm DVI Option - Gemini TF 64 System Number: 882471 Continuous CT (CCT) is a CT or PET/CT product option that enables physicians and radiologists to produce real-time CT images to facilitate image-guided surgery. | Continuous CT Fluoroscopy DVI Ceiling Mount | 2 | CDRH | 2021-11-23 20:23 |  |  | Philips North America Llc | Significant wear of the Monitor Bracket Assembly shaft may go unrecognized could lead to the Monitor Bracket Assembly detaching from the Suspension Arm, and resulting in injury to patient, operator, bystander, or service person located in the path of a falling part
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190358 | Z-0277-2022 | Continuous CT Fluoroscopy DVI Ceiling Mount - CCT Arm DVI Option - CT Vision ACQSIM CT System Number: 728209 Continuous CT (CCT) is a CT or PET/CT product option that enables physicians and radiologists to produce real-time CT images to facilitate image-guided surgery. | Continuous CT Fluoroscopy DVI Ceiling Mount | 2 | CDRH | 2021-11-23 20:23 |  |  | Philips North America Llc | Significant wear of the Monitor Bracket Assembly shaft may go unrecognized could lead to the Monitor Bracket Assembly detaching from the Suspension Arm, and resulting in injury to patient, operator, bystander, or service person located in the path of a falling part
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190362 | Z-0280-2022 | Continuous CT Fluoroscopy DVI Ceiling Mount - CCT Arm DVI Option - Gemini TF 16 System Number: 882470 Continuous CT (CCT) is a CT or PET/CT product option that enables physicians and radiologists to produce real-time CT images to facilitate image-guided surgery. | Continuous CT Fluoroscopy DVI Ceiling Mount | 2 | CDRH | 2021-11-23 20:23 |  |  | Philips North America Llc | Significant wear of the Monitor Bracket Assembly shaft may go unrecognized could lead to the Monitor Bracket Assembly detaching from the Suspension Arm, and resulting in injury to patient, operator, bystander, or service person located in the path of a falling part
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190361 | Z-0279-2022 | Continuous CT Fluoroscopy DVI Ceiling Mount - CCT Arm DVI Option - TruFlight Select PET/CT System Number: 882438 Continuous CT (CCT) is a CT or PET/CT product option that enables physicians and radiologists to produce real-time CT images to facilitate image-guided surgery. | Continuous CT Fluoroscopy DVI Ceiling Mount | 2 | CDRH | 2021-11-23 20:23 |  |  | Philips North America Llc | Significant wear of the Monitor Bracket Assembly shaft may go unrecognized could lead to the Monitor Bracket Assembly detaching from the Suspension Arm, and resulting in injury to patient, operator, bystander, or service person located in the path of a falling part
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190359 | Z-0278-2022 | Continuous CT Fluoroscopy DVI Ceiling Mount - CCT Arm DVI Option - Gemini GXL 16 Slice System Number: 882410 Continuous CT (CCT) is a CT or PET/CT product option that enables physicians and radiologists to produce real-time CT images to facilitate image-guided surgery. | Continuous CT Fluoroscopy DVI Ceiling Mount | 2 | CDRH | 2021-11-23 20:23 |  |  | Philips North America Llc | Significant wear of the Monitor Bracket Assembly shaft may go unrecognized could lead to the Monitor Bracket Assembly detaching from the Suspension Arm, and resulting in injury to patient, operator, bystander, or service person located in the path of a falling part
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190052 | Z-0276-2022 | HemosIL ReadiPlasTin (20 mL), Part number 0020301400 | HemosIL ReadiPlasTin | 2 | CDRH | 2021-11-23 14:02 |  |  | Instrumentation Laboratory | The firm has received customer reports of performance issues with the affected lot, including increased imprecision, out of range quality controls, and prolonged sample results. If quality controls are not performed or do not pass for each vial of reagent, there is a potential risk of reporting an erroneous result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189521 | Z-0275-2022 | The da Vinci Xi and X surgical systems (IS4000 and IS4200) each consist of three main components: the Surgeon Console (SC), the Vision Cart (VC), and the Patient Cart (PC). The Patient Cart is the operative component of the system and its primary function is to support the instrument arms and camera arm. The X/Xi Instrument Arm Drape is intended to be used with the instrument arms of the PC as a sterile barrier during clinical procedures, and makes the instrument arms sterile and suitable for surgery. X/Xi Instrument Arm Drapes are sterile unless the package is opened or damaged, and are indicated for a single use. The drapes are placed onto the instrument arms by two users (one sterile and one non-sterile)in tandem, with the non-sterile user leading the placement via drape cuff pockets. | da Vinci X/Xi Instrument Arm Drapes | 2 | CDRH | 2021-11-23 |  |  | Intuitive Surgical, Inc. | Incorrect expiration date on its package labeling (shipper box and inner carton label).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190021 | Z-0274-2022 | Freelite Human Lambda Free Kit (for use on the SPAPLUS) REF LK018.S Freelite Human Lambda Free Kit (for use on the SPAPLUS) (x10) REF LK018.10S | Freelite Human Lambda Free Kit for use on the SPAPLUS | 2 | CDRH | 2021-11-19 16:34 |  |  | The Binding Site Group, Ltd. | Due to calibration curves producing a higher activity than routinely expected and that there may be a decrease in separation between calibrator point 5 and calibrator point 6, which can be exacerbated with calibrator 6 activity reaches the optical ceiling of the SPA+ analyzer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189719 | Z-0273-2022 | Cardioblate CryoFlex Surgical Ablation Console | Cardioblate CryoFlex | 2 | CDRH | 2021-11-19 14:25 |  |  | Medtronic Perfusion Systems | There is potential for intermittent electrical connectivity between the console and probe.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189935 | Z-0271-2022 | ARCHITECT i1000SR REF 01L86-01/1L86 and 1L87; ARCHITECT i2000SR REF 03M74-02/3M74; ARCHITECT i2000 REF 08C89-01/1G17 and 8C89 | ARCHITECT i-Systems | 2 | CDRH | 2021-11-19 13:56 |  |  | Abbott Laboratories | Twelve software-related issues affecting software version 9.41 and earlier were identified. The issues include bypassing error codes, incorrect calibration, loss of module communication leading to overfilling wash buffer containers, missing QC flags, signal spikes, abnormal optics reads, and incorrect sample aspiration. The issues may lead to inaccurate results with an unknown bias direction and magnitude for assays measuring cardiac, metabolic (including diabetes), toxicology, prenatal, therapeutic drug monitoring, infectious disease, oncology, hepatic, pancreatic, hematological, endocrinologic, inflammatory analytes.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189936 | Z-0272-2022 | ARCHITECT c4000 REF 02P24-01/2P24, 1P86, 1R24, and 1R25; ARCHITECT c8000 REF 01G06-11/1G06; ARCHITECT c16000 REF 03L77-01/3L77; | ARCHITECT c-Systems | 2 | CDRH | 2021-11-19 13:56 |  |  | Abbott Laboratories | Twelve software-related issues affecting software version 9.41 and earlier were identified. The issues include bypassing error codes, incorrect calibration, loss of module communication leading to overfilling wash buffer containers, missing QC flags, signal spikes, abnormal optics reads, and incorrect sample aspiration. The issues may lead to inaccurate results with an unknown bias direction and magnitude for assays measuring cardiac, metabolic (including diabetes), toxicology, prenatal, therapeutic drug monitoring, infectious disease, oncology, hepatic, pancreatic, hematological, endocrinologic, inflammatory analytes.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189899 | Z-0269-2022 | Western Oxytote DTE Digital VIPR System (with D-size cylinder) Model: MNDS-602 | Digital VIPR System | 2 | CDRH | 2021-11-19 13:25 |  |  | Western/Scott Fetzer Company | VIPR (Valve Integrated Pressure Regulator) - battery rupture, causing the plastic VIPR shroud to break apart with emission of fumes and smoke
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189893 | Z-0267-2022 | Roberts Digital VIPR System (with E-size cylinder) Model: ROC-9840 | Digital VIPR System | 2 | CDRH | 2021-11-19 13:25 |  |  | Western/Scott Fetzer Company | VIPR (Valve Integrated Pressure Regulator) - battery rupture, causing the plastic VIPR shroud to break apart with emission of fumes and smoke
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189942 | Z-0270-2022 | Western Oxytote DTE Digital VIPR System (with E-size cylinder) Model: MNDS-603 | Digital VIPR System | 2 | CDRH | 2021-11-19 13:25 |  |  | Western/Scott Fetzer Company | VIPR (Valve Integrated Pressure Regulator) - battery rupture, causing the plastic VIPR shroud to break apart with emission of fumes and smoke
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189897 | Z-0268-2022 | Western Oxytote DTE Digital VIPR (head only) Model: MNDR-600 | Digital VIPR System | 2 | CDRH | 2021-11-19 13:25 |  |  | Western/Scott Fetzer Company | VIPR (Valve Integrated Pressure Regulator) - battery rupture, causing the plastic VIPR shroud to break apart with emission of fumes and smoke
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189567 | Z-0232-2022 | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device : 7FR Product Code: PT-65709-HFWC The Arrow-Trerotola" Percutaneous Thrombolytic Device (PTDTM) Catheter, in conjunction with the Arrow Rotator Drive Unit, permits mechanical declotting of native arterio-venousfistulae and synthetic dialysis grafts | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device: 7FR | 1 | CDRH | 2021-11-18 19:27 |  |  | Arrow International Inc | 7FR Percutaneous Thrombolytic Device (PTD) Inner lumen (orange) detached from the device basket result in embolization and retrieval with an additional intervention and consequent delay of therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189568 | Z-0233-2022 | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device : 7FR Product Code: PT-65709-W The Arrow-Trerotola" Percutaneous Thrombolytic Device (PTDTM) Catheter, in conjunction with the Arrow Rotator Drive Unit, permits mechanical declotting of native arterio-venousfistulae and synthetic dialysis grafts | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device: 7FR | 1 | CDRH | 2021-11-18 19:27 |  |  | Arrow International Inc | 7FR Percutaneous Thrombolytic Device (PTD) Inner lumen (orange) detached from the device basket result in embolization and retrieval with an additional intervention and consequent delay of therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189566 | Z-0231-2022 | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device : 7FR Product Code: PT-12709-WC The Arrow-Trerotola" Percutaneous Thrombolytic Device (PTDTM) Catheter, in conjunction with the Arrow Rotator Drive Unit, permits mechanical declotting of native arterio-venousfistulae and synthetic dialysis grafts | Arrow Trerotola Percutaneous Thrombolytic Device: 7Fr | 1 | CDRH | 2021-11-18 19:27 |  |  | Arrow International Inc | 7FR Percutaneous Thrombolytic Device (PTD) Inner lumen (orange) detached from the device basket result in embolization and retrieval with an additional intervention and consequent delay of therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189569 | Z-0234-2022 | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device : 7FR Product Code: PT-65709-WC The Arrow-Trerotola" Percutaneous Thrombolytic Device (PTDTM) Catheter, in conjunction with the Arrow Rotator Drive Unit, permits mechanical declotting of native arterio-venousfistulae and synthetic dialysis grafts | Arrow-Trerotola Over-The-Wire PTD Kit Percutaneous Thrombolytic Device: 7FR | 1 | CDRH | 2021-11-18 19:27 |  |  | Arrow International Inc | 7FR Percutaneous Thrombolytic Device (PTD) Inner lumen (orange) detached from the device basket result in embolization and retrieval with an additional intervention and consequent delay of therapy.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189933 | Z-0266-2022 | Cobas u 601 urinalysis test system, Catalog number 06334601001 The cobas U 601 urinalysis test system is comprised of the cobas u 601 urine analyzer and the cobas u pack. The cobas u 601 urine analyzer when used with the cobas u pack is a fully automated urinalysis system intended for the in vitro qualitative or semi-quantitative determination of urine analytes, including pH, leukocytes, nitrite, protein, glucose, ketones, urobilinogen, bilirubin, color and erythrocytes, as well as clarity. | Cobas u 601 urinalysis test system | 2 | CDRH | 2021-11-18 17:06 |  |  | Roche Diagnostics Operations, Inc. | A potential risk for false negative nitrite results exists when endogenous creatinine levels are 15,000 mg/L and above. No interference of Nitrite results was observed at creatinine levels up to 9,000 mg/L.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190161 | Z-0265-2022 | Smith and Nephew EVOS Self-Tapping Locking Screw SST: (1)REF 72413575, 3.5MM x 75MM, and (2) REF 72413570, 3.5MM X 70MM. Sterile. | Smith & Nephew EVOS screws | 2 | CDRH | 2021-11-18 14:04 |  |  | Smith & Nephew, Inc. | The labels on the packages of screws were switched.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189196 | Z-0263-2022 | Quantum 2000 Electrosurgical Generators Cooper Surgical Part Number: 909075-05W | Quantum 2000 Electrosurgical Generator | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | XXX
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188749 | Z-0255-2022 | REGARD Ultrasound Tray | Ultrasound Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189193 | Z-0260-2022 | LEEP PRECISION Generator 220V Cooper Surgical Part Number: LP-20-220 | LEEP PRECISION Generator 220V | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188746 | Z-0252-2022 | REGARD Ear Tray | Ear Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188747 | Z-0253-2022 | REGARD Newborn Tray | REGARD Newborn Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189183 | Z-0258-2022 | LEEP PRECISION Integrated System, 220V CooperSurgical part numbers: LP-10-220 | LEEP PRECISION Integrated System 220V | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189394 | Z-0264-2022 | Permobil TiLite, Models: Aero Z, and ZRA | Permobil TiLite | 2 | CDRH | 2021-11-18 |  |  | TiSport, Llc | Some affected wheelchairs, with transit tie downs, that are used as a passenger seat in a motor vehicle may not be configured in the right orientation to perform as intended if involved in a motor vehicle accident. Some affected wheelchairs can no longer be used as a passenger seat until transit tie downs are replaced with redesigned transit tie downs, verified, and released to the market.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189182 | Z-0257-2022 | LEEP PRECISION Integrated System, 120V CooperSurgical part number: LP-10-120 | LEEP PRECISION Integrated System 120V | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189195 | Z-0262-2022 | Quantum 2000 Electrosurgical Generators Cooper Surgical Part Number: 909075-05 | Quantum 2000 Electrosurgical Generator | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188750 | Z-0256-2022 | REGARD Vitrectomy Tray | Vitrectomy Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188745 | Z-0251-2022 | REGARD Core Biopsy Tray | Core Biopsy Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188743 | Z-0250-2022 | REGARD Basic Cataract tray | Basic Cataract tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188748 | Z-0254-2022 | REGARD Stereotactic Tray | Stereotactic Tray | 2 | CDRH | 2021-11-18 |  |  | ROi CPS LLC | Povidone-Iodine swabstick manufactured by PDI, Inc. recalled due to Out of Specifications was used as a component in some of the ROI CPS, LLC products.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189184 | Z-0259-2022 | LEEP PRECISION Generator 120V Cooper Surgical Part Number: LP-20-120 | LEEP PRECISION Generator 120V | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189194 | Z-0261-2022 | Quantum 2000 Electrosurgical Generators Cooper Surgical Part Number: 909075 | Quantum 2000 Electrosurgical Generator | 2 | CDRH | 2021-11-18 |  |  | CooperSurgical, Inc. | Coagulate function may not operate with the use of the Hand Switch, failure of the device to coagulate tissue may lead to a prolonged procedure requiring medical intervention to control any intraoperative bleeding and complete the case.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189397 | Z-0248-2022 | Vit Kit-Freeze is intended for use in assisted reproductive procedures for vitrification and storage of human oocytes (MII), pronuclear (PN) zygotes through day 3 cleavage stage embryos and blastocyst stage embryos. This kit is designed for use with CryoTip (Catalog #40709), and Vitrification Thaw Kit (Vit Kit-Thaw) for optimal recovery of specimens. Vitrification Freeze kit (which includes the VS Vitrification Solution) are cryopreservation media intended for use in the vitrification of oocytes (MII) and pronuclear (PN) zygotes, cleavage stage embryos, and blastocyst stage embryos. Catalog numbers: 90133-DCOS Vitrification Freeze kit, contained 9x1 mL. Vitrification solution, VS, Catalog #90132. Product form: Liquid. Equilibration Solution-ES is a HEPES buffered solution of Medium-199 containing gentamicin sulfate, 7.5% (v/v) of each DMSO and ethylene glycol and 20% (v/v) Dextran Serum Supplement (DSS). Vitrification Solution-VS is a HEPES buffered solution of Medium-199 containing gentamicin sulfate, 15% (v/v) of each DMSO and ethylene glycol, 20% (v/v) DSS and 0.5 M sucrose. DSS is a protein supplement consisting of 50 mg/mL therapeutic grade Human Serum Albumin (HSA) and 20 mg/mL Dextran. DSS is used at 20% (v/v) in Vit Kit-Freeze for a final concentration of 10 mg/mL HSA and 4 mg/mL Dextran. These two solutions are to be used in sequence according to the step-wise microdrop vitrification protocol. Vitrification Solution-VS Aseptically filtered. Test results reported on C of A Vit Kit -Freeze (Vitrification Freeze Kit) with Gentamycin and DSS Contains 9 x 1 mL Equilibration Solution- ES; 9 X 1 mL Vitrification Solution-VS ...90133-DCOS...RX Only...Sterile ...20YY-MM-DD | Vitrification Freeze kit, Catalog # 90133 DSOC | 2 | CDRH | 2021-11-17 17:45 |  |  | Fujifilm Irvine Scientific, Inc. | Due to a component in kit being labeled with the incorrect Expiration Date
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189398 | Z-0249-2022 | Vit Kit-Freeze is intended for use in assisted reproductive procedures for vitrification and storage of human oocytes (MII), pronuclear (PN) zygotes through day 3 cleavage stage embryos and blastocyst stage embryos. This kit is designed for use with CryoTip (Catalog #40709), and Vitrification Thaw Kit (Vit Kit-Thaw) for optimal recovery of specimens. Vitrification Freeze kit (which includes the VS Vitrification Solution) are cryopreservation media intended for use in the vitrification of oocytes (MII) and pronuclear (PN) zygotes, cleavage stage embryos, and blastocyst stage embryos. Catalog numbers: 90133-SO Vitrification Freeze kit, contained 9x1 mL. Vitrification solution, VS, Catalog #90132. Product form: Liquid. Equilibration Solution-ES is a HEPES buffered solution of Medium-199 containing gentamicin sulfate, 7.5% (v/v) of each DMSO and ethylene glycol and 20% (v/v) Dextran Serum Supplement (DSS). Vitrification Solution-VS is a HEPES buffered solution of Medium-199 containing gentamicin sulfate, 15% (v/v) of each DMSO and ethylene glycol, 20% (v/v) DSS and 0.5 M sucrose. DSS is a protein supplement consisting of 50 mg/mL therapeutic grade Human Serum Albumin (HSA) and 20 mg/mL Dextran. DSS is used at 20% (v/v) in Vit Kit-Freeze for a final concentration of 10 mg/mL HSA and 4 mg/mL Dextran. These two solutions are to be used in sequence according to the step-wise microdrop vitrification protocol. Vitrification Solution-VS Aseptically filtered. Test results reported on C of A Vit Kit -Freeze (Vitrification Freeze Kit) with Gentamycin and DSS Contains 2 x 1 mL Equilibration Solution- ES; 2 X 1 mL Vitrification Solution-VS ...90133-DCOS...RX Only...Sterile ...20YY-MM-DD | Vitrification Freeze kit | 2 | CDRH | 2021-11-17 17:45 |  |  | Fujifilm Irvine Scientific, Inc. | Due to a component in kit being labeled with the incorrect Expiration Date
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189881 | Z-0247-2022 | VITROS Immunodiagnostic Products Anti-HBc IgM Controls, Product Code 6801427 For use in monitoring the performance of the VITROS ECi/ECiQ/3600 Immunodiagnostic Systems and the VITROS 5600/XT 7600 Integrated Systems | VITROS Immunodiagnostic Products Anti-HBc IgM Controls | 3 | CDRH | 2021-11-17 16:06 |  |  | Ortho-Clinical Diagnostics, Inc. | Ortho Clinical Diagnostics confirmed a complaint that some VITROS Immunodiagnostic Products Anti-HBc IgM Controls Lot 1690 Controls Values booklet contained information for the incorrect product. While the accuracy of patient test results is not affected by this issue, there is a potential for delayed results due to the troubleshooting of Quality Control results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189449 | Z-0225-2022 | Scrip Creme All Purpose Lotion Intended for use in acoustic coupling of ultrasound transducer to tissue | Scrip | 1 | CDRH | 2021-11-17 15:03 |  |  | Scrip Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189448 | Z-0224-2022 | Scrip Ultrasound Gel Clear, 5L Intended for use in acoustic coupling of ultrasound transducer to tissue | Scrip | 1 | CDRH | 2021-11-17 15:03 |  |  | Scrip Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189526 | Z-0245-2022 | Ilizarov Wire Tensioner, Model 71070341, UDI (10)08033201842072(10)21HP02550(11)210831(21E434 | Ilizarov Wire Tensioner | 2 | CDRH | 2021-11-17 02:40 |  |  | Medicalplastic S.R.L. | Medicalplastic s.r.l. is conducting a voluntary field action to remove a group of Ilizarov Wire Tensioners due to a manufacturing error. An internal washer of the Ilizarov Wire Tensioner has been made in a wrong plastic material that could wear and break during the sterilization process in autoclave. Consequently this could cause malfunctioning of the tensioner.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189710 | Z-0243-2022 | API 50 CH, Model 50300 | API 50 CH | 2 | CDRH | 2021-11-17 |  |  | bioMerieux, Inc. | There is potential to provide an incorrect organism identification.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189934 | Z-0244-2022 | Torque Limiting Adaptor, AO Connect, 4 N-m, part #027-203, reusable, a component of the Bodycad Fine Osteotomy Implant and Disposable Instrumentation Tray, part #025-004. | Bodycad | 2 | CDRH | 2021-11-17 |  |  | Laboratoires Bodycad, Inc. | There is a risk of breakage of the torque limiting adaptor if used counterclockwise and it must not be used with a power tool due to no test data validating this use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189951 | Z-0237-2022 | OMNI K2 Hip Stem, Manufacturer's Product Code 130006 The OMNI K2 Hip Stem is intended for use as the femoral component of a primary total hip replacement. This femoral hip stem is intended for uncemented fixation and single use implantation. | OMNI K2 Hip Stem | 2 | CDRH | 2021-11-16 15:51 |  |  | OmniLife Science | The device is missing grit blasting on the proximal surface of the femoral stem, which may lead to the possibility of product revision in the event of undersizing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189880 | Z-0236-2022 | BIO-RAD BioPlex 2200 REF 12000651 Syphilis Total & RPR Calibrator Set CAL The BioPlex 2200 Syphilis Total & RPR Calibrator Set is intended for the calibration of the BioPlex 2200 Syphilis Total & RPR Reagent Pack. | BioPlex 2200 Syphilis Total & RPR | 2 | CDRH | 2021-11-16 15:15 |  |  | Bio-Rad Laboratories, Inc. | Due to two issues: 1) Customer complaints associated with greater than expected Rapid Plasma Reagin (RPR) reactivity which does not confirm with treponemal antibody test which may be due to a positive correlation between the mass COVID vaccinations and false reactive RPR results, and 2) During routine stability monitoring, an increase in mean bias above specification was observed which can cause elevated RPR reactivity potentially due to early shelf-life failure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189879 | Z-0235-2022 | BIO-RAD BioPlex 2200 REF 12000650 SYPH T PACK Syphilis Total & RPR Reagent Pack The BioPlex Syphilis Total & RPR kit is a multiplex flow immunoassay intended for the qualitative detection of total (IgG/IgM) antibodies to Treponema pallidum and the qualitative detection and/or titer determination of nontreponemal reagin antibodies in human serum or plasma | BioPlex 2200 Syphilis Total & RPR | 2 | CDRH | 2021-11-16 15:15 |  |  | Bio-Rad Laboratories, Inc. | Due to two issues: 1) Customer complaints associated with greater than expected Rapid Plasma Reagin (RPR) reactivity which does not confirm with treponemal antibody test which may be due to a positive correlation between the mass COVID vaccinations and false reactive RPR results, and 2) During routine stability monitoring, an increase in mean bias above specification was observed which can cause elevated RPR reactivity potentially due to early shelf-life failure.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189683 | Z-0229-2022 | bvi CustomEyes Kit-Lafayette General Medical Center: CustomEyes Cataract Surgical Kits that contain Beaver Xstar Crescent Knife (REF#373809) as a component REF# 58000572 REF#: 373809 |  | 2 | CDRH | 2021-11-16 13:08 |  |  | Beaver Visitec International, Inc. | Product packaged as Beaver Xstar Crescent Knife, 2.5mm, 55 degrees bevel-down blades contain bevel-up blades
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189684 | Z-0230-2022 | bvi CustomEyes Kit-CustomEyes Cataract Surgical Kits that contain Beaver Xstar Crescent Knife (REF#373809) as a component-non-sterile sample for evaluation purposes REF# BVI 58000587 |  | 2 | CDRH | 2021-11-16 13:08 |  |  | Beaver Visitec International, Inc. | Product packaged as Beaver Xstar Crescent Knife, 2.5mm, 55 degrees bevel-down blades contain bevel-up blades
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189681 | Z-0227-2022 | Beaver Xstar Crescent Knife, 2.5mm, 55 degrees, bevel down. Intended for or cataract surgery. REF#: 373809 | BVI Beaver Xstar Crescent Knife | 2 | CDRH | 2021-11-16 13:08 |  |  | Beaver Visitec International, Inc. | Product packaged as Beaver Xstar Crescent Knife, 2.5mm, 55 degrees bevel-down blades contain bevel-up blades
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189682 | Z-0228-2022 | bvi CustomEyes Kit-AMO Germany GmbH Katarakt-set - Dr. Wagner: CustomEyes Cataract Surgical Kits that contain Beaver Xstar¿ Crescent Knife (REF#373809) as a component a) REF# 588221 REF#: 373809 |  | 2 | CDRH | 2021-11-16 13:08 |  |  | Beaver Visitec International, Inc. | Product packaged as Beaver Xstar Crescent Knife, 2.5mm, 55 degrees bevel-down blades contain bevel-up blades
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189845 | Z-0226-2022 | Sensis/ Sensis Vibe systems with software VD1X as follows: (1)Sensis Model 10764561 (2) Sensis Vibe Hemo Model 11007641 (3)Sensis Vibe Combo Model 11007642 Intended Use: A diagnostic and administrative tool supporting hemodynamic catheterizations and/or electrophysiology studies | Siemens Sensis Vibe recording system | 2 | CDRH | 2021-11-16 12:06 |  |  | Siemens Medical Solutions USA, Inc | System operator manual states that the system should be rebooted once, every 7 days. In some cases, users are not following these instructions and the system is running for more than 7 days. Under rare circumstances, this may result in a partial freeze of the user interface which would no longer update the vital signs
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189633 | Z-0217-2022 | SMART FLEX 6x120 BIL, 120cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189636 | Z-0220-2022 | SMART FLEX 10x100 BIL, 80cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189637 | Z-0221-2022 | SMART FLEX 10x100 BIL, 120cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189634 | Z-0218-2022 | SMART FLEX 6x150 BIL, 120cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189635 | Z-0219-2022 | SMART FLEX 8x60 BIL, 80cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189632 | Z-0216-2022 | SMART FLEX 9x60 BIL, 80cm The S.M.A.R.T.¿ Flex Biliary Stent System is indicated for use in the palliation of malignant strictures in the biliary tree. | S.M.A.R.T. Flex | 2 | CDRH | 2021-11-12 14:52 |  |  | Cordis Corporation | There is a potential for distal tip dislodgement or separation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189598 | Z-0222-2022 | Transseptal Needle, Trocar. RPN TSNC-18-71.0, TSNC-19-56.0. GPN G02364, G02365. Product consists of a transseptal needle and obturator. | Transseptal Needle | 2 | CDRH | 2021-11-12 |  |  | Cook Inc. | Transseptal needles may contain rust on the interior and/or exterior of the needle, which may not be detected by the user. Potential adverse events if an affected product is used ranges from increased procedural time to inflammatory reactions, which at their most severe include systemic reactions potentially leading to permanent impairment or death.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189599 | Z-0223-2022 | Transseptal Needle with Catheter. RPN TSN-17-75.0-ENDRYS. GPN G19261. Coaxial set consisting of an outer catheter, curved-tip metal cannula, and a tapered-tip inner needle | Transseptal Needle with Catheter | 2 | CDRH | 2021-11-12 |  |  | Cook Inc. | Transseptal needles may contain rust on the interior and/or exterior of the needle, which may not be detected by the user. Potential adverse events if an affected product is used ranges from increased procedural time to inflammatory reactions, which at their most severe include systemic reactions potentially leading to permanent impairment or death.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189349 | Z-0215-2022 | Hemodialysis Delivery System, Software Version 2.x. | PrixMax System | 2 | CDRH | 2021-11-11 |  |  | Baxter Healthcare Corporation | If the operator initiates therapy with a saved prescription profile and makes a change to the prescription after a disposable filter change using the Same Patient button, the system may suggest values from the original prescription profile, rather than the current prescription.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189843 | Z-0214-2022 | T4 EIA Kit, Cat No: BC-1007, Store at 2-8C, For in vitro diagnostic use only This product s intended use is For the quantitative determination of total thyroxine (T4) concentration in human serum. The test is useful in the diagnosis and treatment of thyroid disorders. | Total Thyroxine (T4) Enzyme Immunoassay Test Kit | 2 | CDRH | 2021-11-10 14:45 | 2021-11-10 20:07 |  | BioCheck, Inc. | Immunoassay TEST Kit do not meet product shelf-life standards resulting in failed controls.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189590 | Z-0183-2022 | Ellume COVID-19 Home Test | Ellume COVID-19 Home Test | 1 | CDRH | 2021-11-10 |  |  | ELLUME LTD | Due to a higher rate of false positive test results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189727 | Z-0213-2022 | Philips Azurion Interventional Fluoroscopic X-Ray System, Software version 2.1.x, Model numbers 722063, 722064, 722067, 722068, 722078, 722079, 722221, 722222, 722223, 722224, 722225, 722226, 722227, 722228. (Azurion 3 M12, Azurion 3 M15, Azurion 5 M12, Azurion 5 M20, Azurion 7 B12, Azurion 7 B20, Azurion 7 M12, Azurion 7 M20) | Philips Azurion Interventional Fluoroscopic X-Ray System | 2 | CDRH | 2021-11-09 |  |  | Philips North America Llc | When the user presses or releases both the APC (accept) button and the Float Tabletop (panning) button at the same time, the geometry may stop reacting on movement requests (table lock-up) and X-ray imaging becomes unavailable.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189725 | Z-0212-2022 | THERAKOS CELLEX Photopheresis Procedural Kit PRODUCT CODE: CLXUSA | CELLEX Procedural Kit | 2 | CDRH | 2021-11-05 13:55 |  |  | Therakos, Inc. | Product was released prior to all testing being completed. If used and a drive tube leak were to occur, may cause illness or injury.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189622 | Z-0210-2022 | Covidien Emprint Ablation Catheter with Thermosphere Technology Bronchoscopic Ablation, CFN CA108L1 | Covidien Emprint Ablation Catheter with Thermosphere Technology Bronchoscopic Ablation | 2 | CDRH | 2021-11-05 11:45 |  |  | Covidien Llc | The firm identified an error in the device temperature monitoring circuitry which is designed to interrupt energy delivery in the event of a primary cooling system failure. Use of an affected device in conjunction with a primary cooling system failure could lead to the applicator tip overheating and associated heat damage which could result in an allergic/toxic reaction, unintended tissue ablation, tissue burn, and/or a delay of treatment while an alternate device is located.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189456 | Z-0205-2022 | McKesson PREVENT HT Safety Blood Collection Needles, 22G 1 1/4 inch, MFR # 16-NBC22GS | McKesson PREVENT HT Safety Blood Collection Needle | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189457 | Z-0206-2022 | McKesson PREVENT HT Safety Winged Blood Collection Needle with Tube Holder, MFR # 16-NBCH22GS | McKesson PREVENT HT Safety Winged Blood Collection Needle with Tube Holder | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189461 | Z-0209-2022 | HENRY SCHEIN Safety Blood Collection Needle 22Gx 1 1/4 Inch, REF 570-2068, 100pcs/box 1000pcs/ctn 350*210*350mm | Henry Schein Safety Blood Collection Needle | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189630 | Z-0211-2022 | QuikClot Combat Gauze-For Temporary External Use To Control Traumatic Bleeding Model / Part# 200 | Z-Medica | 2 | CDRH | 2021-11-05 |  |  | Z-Medica, LLC | Lack of Packaging seal integrity may result in a sterile barrier breach.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189458 | Z-0207-2022 | McKesson PREVENT HT Safety Winged Blood Collection Needle with Tube Holder, MFR # 16-NBCH21GS | McKesson PREVENT HT Safety Winged Blood Collection Needle with Tube Holder | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189455 | Z-0204-2022 | McKesson PREVENT HT Safety Blood Collection Needles, 21G 1 1/4 inch, MFR # 16-NBC21GS | McKesson PREVENT HT Safety Blood Collection Needle | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189459 | Z-0208-2022 | HENRY SCHEIN Safety Blood Collection Needle 21Gx 1 1/4 Inch, REF 570-2067, 100pcs/box 1000pcs/ctn 350*210*350mm | Henry Schein Safety Blood Collection Needle | 2 | CDRH | 2021-11-05 |  |  | Zhejiang Kindly Medical Devices Co., Ltd. | Potential for difficulty in safety mechanism activation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189956 | Z-0202-2022 | Aspiration Syringe Kit, REF: MVSK60 | Syringe Kit | 2 | CDRH | 2021-11-04 |  |  | MICROVENTION INC. | Syringe Kits labeled for distribution OUS were inadvertently shipped to US customers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190005 | Z-0203-2022 | GM Helix Acqua Implant 3.5x10, Article No. 140.944 | Neodent Implant Systems - GM Helix Acqua Implant | 2 | CDRH | 2021-11-04 |  |  | Jjgc Industria E Comercio De Materials Dentarios Sa | 3.5x11.5mm implants were mislabeled as 3.5x10mm implants. In borderline cases where the bone quality is low and the bone ridge has a maximum indication for the implant height (10mm), if the incorrect implant is installed, it could result in damage to the nerve or adjacent bone structures/cavities.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189466 | Z-0201-2022 | Clinician Programmer Application (CPA) model A610 Clinician Software Application | Clinician Programmer Application (CPA) | 2 | CDRH | 2021-11-04 |  |  | Medtronic Neuromodulation | A software anomaly may occur with the clinician programmer application.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189631 | Z-0200-2022 | Artis zeego, Fluoroscopic X-Ray System Models: 10280959,10502505, 10848283 in combination with the following workplace/system application software models: Product/Model #: syngo X Workplace, 10281061 syngo X Workplace, 10281299 syngo X Workplace, 10502522 Syngo Application Software, 10848815 | Siemens Artis zeego | 2 | CDRH | 2021-11-04 |  |  | Siemens Medical Solutions USA, Inc | Software error, the software can produce an incorrect interpretation of the table rotation if the table has been rotated by more than 4degrees in either direction. This can cause the 3D image reconstruction to be shown rotated by up to 5 degrees relative to the patient and to be used inappropriately by the system and result in incorrect treatment of the patient
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189778 | Z-0194-2022 | IPL Coupling Gel, I L (AX1009013), 12 pack of 0.25 L (AX1009018), 6 pack of 1 L (AX1009019 or AX1009063); and IPL Coupling Gel or EcoGel 200, 0.25 L (AX1009012) manufactured by Eco-Med Pharmaceutical | IPL Coupling Gel, EcoGel 200 | 1 | CDRH | 2021-11-04 |  |  | Lumenis, Inc. | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189597 | Z-0199-2022 | The MEDRAD Twist & Go Disposable Syringe 150 mL with Quick Fill Tube is a syringe made from a clear polycarbonate material and is an accessory to the MEDRAD Mark 7 Arterion Injection System. It is designed to be loaded into the injector head of the injection system. | MEDRAD Twist & Go Disposable Syringe 150 mL with Quick Fill Tube | 2 | CDRH | 2021-11-03 |  |  | Bayer Medical Care, Inc. | Labeling error: MEDRAD Mark 7 Arterion Injection System Syringes (ART 700 SYR) labeled as Twist & Go Syringes (TAG 150 SYR)
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189491 | Z-0195-2022 | Atellica IM SARS-CoV-2 Antigen Assay (100 Test Kit), SMN # 11207861 | Atellica IM SARS-CoV-2 Antigen (CoV2Ag) Assay | 2 | CDRH | 2021-11-03 |  |  | Siemens Healthcare Diagnostics, Inc | Data provided in the Emergency Use Authorization (EUA) submission was insufficient to determine that the assays were adequately validated for the intended use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189748 | Z-0198-2022 | ImplantDirect, REF 834713, Legacy 2 Implant 4.7mmD x 13mmL, 4.5mmD Platform, | ImplantDirect | 2 | CDRH | 2021-11-03 |  |  | Implant Direct Sybron Manufacturing LLC | Packaged dental implant contains a different size then the size declared on the labeling.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189492 | Z-0196-2022 | ADVIA Centaur SARS-CoV-2 Antigen Assay (100 Test Kit), SMN # 11207866 | ADVIA Centaur SARS-CoV-2 Antigen Assay | 2 | CDRH | 2021-11-03 |  |  | Siemens Healthcare Diagnostics, Inc | Data provided in the Emergency Use Authorization (EUA) submission was insufficient to determine that the assays were adequately validated for the intended use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189746 | Z-0197-2022 | ImplantDirect, REF 834711, Legacy 2 Implant 4.7mmD x 11.5mmL, 4.5mmD Platform, | ImplantDirect | 2 | CDRH | 2021-11-03 |  |  | Implant Direct Sybron Manufacturing LLC | Packaged dental implant contains a different size then the size declared on the labeling.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189705 | Z-0190-2022 | Arietta 850 Ultrasound- Software Version: V1.0.0 through V4.1.3 | Arietta 850 Ultrasound- Software Version: V1.0.0 through V4.1.3 | 2 | CDRH | 2021-11-02 13:30 |  |  | FujiFilm Healthcare Americas Corporation | Diagnostic ultrasound system with the specified software version and used in conjunction with a Fujifilm (formally Hitachi) transducer does not display the measurement results correctly with two cardiology measurement functions: Mitral Regurgitation Flow (MR Flow) Measurement; Mitral Valve (MV) Measurement. In addition there is an error in the instruction manual.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189706 | Z-0191-2022 | Lisendo 880 Ultrasound Software Version: V1.0 through V4.1.3 | Lisendo 880 Ultrasound Software Version: V1.0 through V4.1.3 | 2 | CDRH | 2021-11-02 13:30 |  |  | FujiFilm Healthcare Americas Corporation | Diagnostic ultrasound system with the specified software version and used in conjunction with a Fujifilm (formally Hitachi) transducer does not display the measurement results correctly with two cardiology measurement functions: Mitral Regurgitation Flow (MR Flow) Measurement; Mitral Valve (MV) Measurement. In addition there is an error in the instruction manual.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189704 | Z-0189-2022 | Arietta 750 Ultrasound Software Version: V1.0.0 through V2.1.3 | Arietta 750 Ultrasound Software Version: V1.0.0 through V2.1.3 | 2 | CDRH | 2021-11-02 13:30 |  |  | FujiFilm Healthcare Americas Corporation | Diagnostic ultrasound system with the specified software version and used in conjunction with a Fujifilm (formally Hitachi) transducer does not display the measurement results correctly with two cardiology measurement functions: Mitral Regurgitation Flow (MR Flow) Measurement; Mitral Valve (MV) Measurement. In addition there is an error in the instruction manual.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189722 | Z-0192-2022 | Arietta 65 Ultrasound Software Versions: V1.0 through V4.0.0 | Fujifilm | 2 | CDRH | 2021-11-02 |  |  | FujiFilm Healthcare Americas Corporation | When the ultrasound system is used in combination with a Fujifilm (formally Hitachi) transducer, there is an error in the calculation formula of Effective Regurgitant Orifice Area (EROA) when you perform MR Flow combined with Left Ventricular Outflow Tract Flow (LVOT Flow) after performing an echocardiogram. The incorrect value is displayed. In addition, the error appears in the instruction manual
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189788 | Z-0193-2022 | Aptio Automation Interface Module to ADVIA Centaur XP/XPT. Aptio by Siemens Centaur Interface Firmware (FW) CENa_3-1-0-1 and lower. SMN 10703038 | Aptio AutomationAptio Automation Interface Module to ADVIA Centaur XP/XPT | 2 | CDRH | 2021-11-02 |  |  | Siemens Healthcare Diagnostics, Inc. | The Aptio Automation Firmware for the ADVIA Centaur XP/XPT Interface Module may lead, in some cases, to an incorrect association of test results to sample ID. Incorrect association of test results could impact clinical interpretations and clinical decisions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189545 | Z-0187-2022 | COVENA PICC Sleeve | COVENA PICC Sleeve | 2 | CDRH | 2021-11-01 |  |  | Queen Comfort Products lLC | Marketed without FDA clearance
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189471 | Z-0188-2022 | 3M Red Dot Electrodes, Radiolucent Monitoring Electrode with Foam Tape, Catalog No. 2244, (3M ID (SKU): 70-2011-8506-6, SAP ID: 7100212989), packaged 10 electrodes per card, 5 cards per pouch, 20 pouches per shipper box, 1,000 electrodes total per shipper box. The 3M Red Dot Radiolucent Monitoring Electrode with Foam Tape 2244 is intended to be used by healthcare professionals for ECG monitoring. This electrode is disposable, intended for single use, and has been tested for up to 3 days wear. | 3M Red Dot | 2 | CDRH | 2021-11-01 |  |  | 3M Company - Health Care Business | Potential for breakage of the ECG electrode carbon stud from the eyelet.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189592 | Z-0119-2022 | Monoject 0.9% Sodium Chloride, Flush Prefilled Syringe, 10 mL, REF: 8881570121 inside the following Aligned Medical Solutions Custom Convenience Kits: AMS-9041CP Leaderflex Insertion Kit with Ultrasound, AMS-9046CP-1 Insertion Tray-RX, AMS8939A Universal Procedure Pack w/Split Drape, AMS9957A Port Insertion Pack, and AMS12149 Procedure Pack | Custom Convenience Kits | 1 | CDRH | 2021-10-29 16:05 |  |  | Windstone Medical Packaging, Inc. | Monoject 0.9% Sodium Chloride Flush Syringe Manufactured by Cardinal Health and placed into kits. Potential for the plunger to draw back after air has been expelled and reintroduce air back into the syringe.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189530 | Z-0186-2022 | Flexor Check-Flo Introducer - Raabe Modification, Introducer Set, RPN KCFW-5.0-38-90-RB-RAABE, GPN G13069 The introducer sheath incorporates a hydrophilic coated Flexor shaft with a hemostasis valve and is provided with a dilator. | Flexor Check-Flo Introducer - Raabe Modification | 2 | CDRH | 2021-10-29 13:01 |  |  | Cook Inc. | Affected lots may be manufactured incorrectly with the radiopaque marker band located just below the Check-Flo proximal fitting instead of the distal tip. The issue may not be identified until the device is placed in the patient and viewed under fluoroscopy, leading to increased procedural time.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189529 | Z-0185-2022 | Flexor Check-Flo Introducer, Introducer Set, RPN KCFW-6.0-35-45-RB, GPN G09908 The introducer sheath incorporates a hydrophilic coated Flexor shaft with a hemostasis valve and is provided with a dilator. | Flexor Check-Flo Introducer | 2 | CDRH | 2021-10-29 13:01 |  |  | Cook Inc. | Affected lots may be manufactured incorrectly with the radiopaque marker band located just below the Check-Flo proximal fitting instead of the distal tip. The issue may not be identified until the device is placed in the patient and viewed under fluoroscopy, leading to increased procedural time.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189315 | Z-0184-2022 | FIVE S 3.5x65, Flexible Intubation Videoscope, REF: 091361-01 (1 Unit), and 091361-06 (6 Units) | Karl Storz FIVE S Flexible Intubation Videoscope | 2 | CDRH | 2021-10-29 10:24 |  |  | Karl Storz Endoscopy | During a regular requalification review, affected flexible intubation video endoscopes did not achieve the required sterility assurance level. There is a risk that the patient may be exposed to a higher risk of infection.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189744 | Z-0182-2022 | Oncentra Brachy users using version 4.0 and above with Applicator Modeling license 170283-00 | Oncentra Brachy | 2 | CDRH | 2021-10-28 13:38 |  |  | Elekta, Inc. | Under certain circumstances, there can be a difference between the planned and delivered dose of the two ovoid channels.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189522 | Z-0180-2022 | Atellica CH Urinary/Cerebrospinal Fluid Protein (UCFP), SMN 11097543 | Atellica CH Urinary/Cerebrospinal Fluid Protein (UCFP) | 2 | CDRH | 2021-10-28 |  |  | Siemens Healthcare Diagnostics, Inc. | The firm has confirmed the potential for Atellica CH Urinary/Cerebrospinal Fluid Protein reagent carryover impacting Enzymatic Hemoglobin A1c (A1c_E and A1c_EM) results. Falsely depressed Enzymatic Hemoglobin A1c results may affect consideration of intervention.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189523 | Z-0181-2022 | Advia Chemistry Urinary/Cerebrospinal Fluid Protein (UCFP), SMN 11319151 | Advia Chemistry Urinary/Cerebrospinal Fluid Protein (UCFP) | 2 | CDRH | 2021-10-28 |  |  | Siemens Healthcare Diagnostics, Inc. | The firm has confirmed the potential for ADVIA Chemistry Urinary/Cerebrospinal Fluid Protein reagent carryover impacting Enzymatic Hemoglobin A1c (A1c_E and A1c_EM) results. Falsely depressed Enzymatic Hemoglobin A1c results may affect consideration of intervention.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189595 | Z-0179-2022 | BD Syringe Tip Cap Bulk Sterile Convenience Pak-1000 units/shelf carton. Used to seal syringes and maintain container closure integrity throughout the shelf life of the BD WWD syringe. Catalog Number: 305822 | BD Syringe Tip Cap Bulk Sterile Convenience Pak | 2 | CDRH | 2021-10-27 13:28 | 2021-10-27 21:07 |  | Becton Dickinson & Company | Holes in product packaging impacts package integrity and may compromise sterility.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188884 | Z-0178-2022 | Bivona Uncuffed Wire Endotracheal Tubes, | Oral/Nasal Tracheal Tube | 2 | CDRH | 2021-10-27 |  |  | Smiths Medical ASD Inc. | Inadequate pouch seal leading to a compromised sterile barrier on products with a shaft length of 325mm.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188883 | Z-0177-2022 | Bivona Aire-Cuf Endotracheal Tube | Oral/Nasal Tracheal Tube | 2 | CDRH | 2021-10-27 |  |  | Smiths Medical ASD Inc. | Inadequate pouch seal leading to a compromised sterile barrier on products with a shaft length of 325mm.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189439 | Z-0176-2022 | ASC PneumoLiner device, Part No. WA90500US (US market only) The Pneumoliner device is intended for use as a multiple instrument port and tissue containment system during minimally invasive gynecologic laparoscopic surgery to enable the isolation and containment of tissue, considered benign, resected during single-port or multi- site laparoscopic surgery during power morcellation and removal. | PneumoLiner | 2 | CDRH | 2021-10-26 |  |  | Olympus Corporation of the Americas | Due to manufacturing error, the Pneumoliner Bag Distal Tab that exits the Introducer shaft is in the wrong orientation. This results in the user deploying the bag upside down, which will make tissue encapsulation and bag closure more difficult, introducing the risk of trapping the small bowel/viscera in the bag at closure resulting in patient injury.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189414 | Z-0174-2022 | Concerto User Interface Software provided with the following systems: (1) Xstrahl 100 Electronic Brachytherapy - 100kV Superficial X-Ray Therapy System; (2) Xstrahl 150 Electronic Brachytherapy - 150kV Superficial X-Ray Therapy System; (3) Xstrahl 200 - 225kV Superficial / Orthovoltage X-Ray Therapy System; (4) Xstrahl 300 - 300kV Orthovoltage X-Ray Therapy System; (5) Xstrahl X 80 Photoelectric Therapy System. | Concerto | 2 | CDRH | 2021-10-26 |  |  | Xstrahl Limited | If a saved treatment plan with 2 opposing beams is edited prior to approval, then Beam 2 is not updated with the changed parameters upon selecting save, resulting in error messages during the treatment and possible mis-treatment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189437 | Z-0175-2022 | Cliquid MD version 3.4 software used in conjunction with Analyst MD software on SCIEX Citrine & 4500MD series liquid chromatography-tandem mass spectrometry (LC-MS/MS) systems. Model No. 5088288 | Cliquid MD | 2 | CDRH | 2021-10-26 |  |  | AB SCIEX | The values of the Internal Standard (IS) concentrations are incorrectly derived when the user builds a customized test, then set the IS column to hidden and left the column empty. This may lead to the user making incorrect conclusions using incorrect results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189497 | Z-0118-2022 | ROSA One 3.1 Brain application The device is intended for the spatial positioning and orientation of instruments holders or tool guides to be used by trained surgeons to guide standard surgical instruments during brain and spine surgeries. | ROSA One 3.1 Brain application | 1 | CDRH | 2021-10-25 |  |  | MEDTECH SAS | The firm has become aware of a software anomaly affecting ROSA One 3.1 Brain application which led to the inaccurate placement of an electrode during surgery. The firm has received 3 global complaints related to the issue. An incorrect trajectory could result in serious injury or death if undetected during surgery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189295 | Z-0173-2022 | C315-HIS Delivery Catheter | C315-HIS Delivery Catheter | 2 | CDRH | 2021-10-22 14:31 |  |  | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Medtronic is recalling three lots of unused C315HIS Delivery Catheters due to reports of an inability to pass a lead through the catheter.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189354 | Z-0169-2022 | REF/Catalogue Number GSXE0030, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617661 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189355 | Z-0170-2022 | REF/Catalogue Number GSXE0025B, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620111 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189350 | Z-0165-2022 | REF/Catalogue Number GSX0020A, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631018 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189352 | Z-0167-2022 | REF/Catalogue Number GSXE0020, GORE CARDIOFORM SEPTAL OCCLUDER 20 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617630 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189356 | Z-0171-2022 | REF/Catalogue Number GSXE0030B, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132620128 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189353 | Z-0168-2022 | REF/Catalogue Number GSXE0025, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132617647 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189357 | Z-0172-2022 | REF/Catalogue Number GSX0030H, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132648948 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189351 | Z-0166-2022 | REF/Catalogue Number GSX0025A, GORE CARDIOFORM SEPTAL OCCLUDER 25 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631025 *Not distributed within the US | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189284 | Z-0164-2022 | REF/Catalogue Number GSX0030A, GORE CARDIOFORM SEPTAL OCCLUDER 30 mm Diameter, Catheter Working Length 75 cm, Delivery Profile 10 Fr, STERILE EO, Rx Only, UDI: 00733132631032 | GORE CARDIOFORM SEPTAL OCCLUDER | 2 | CDRH | 2021-10-22 13:56 |  |  | W L Gore & Associates, Inc. | Septal Occluders contain the incorrect expiration date. Product is labeled with a 3 yr. expiration date instead of 2 yrs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189265 | Z-0163-2022 | Pneupac paraPac plus 300 ventilator kit, REF P300NGB | Pneupac paraPac plus 300 ventilator kit | 3 | CDRH | 2021-10-22 13:14 |  |  | Smiths Medical ASD Inc. | Two devices were labeled with the same serial number
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189418 | Z-0001-2022 | RAPTAR Laser Range Finders | RAPTAR Laser Range Finders | 2 | CDRH | 2021-10-22 10:08 |  |  | Potomac River Group, Corp. | RAPTAR units are capable of being field modified to enable the use of the higher power visible and IR laser pointers and the IR Illuminator.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189465 | Z-0161-2022 | Easi-Care Gait Belt Metal Buckle (54", 60", and 72" lengths) | Easi-Care Gait Belt | 2 | CDRH | 2021-10-21 11:26 |  |  | Kinsman Enterprises Inc | The belt may fail at lower loads than the labeled 600lbs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189272 | Z-0158-2022 | Medtronic Sterile single-use blades that are components of the StraightShot M Series Handpiece and Fusion ENT Navigation Software, and the Midas Rex Microsaw and Triton System with the following product numbers: (1) REF 40A2500, Nucleus Removal Angled Blade, 25 cm x 4.0 mm; (2) REF 40R2000, Nucleus Removal Set, 20 cm x 4.0 mm; (3) REF 40R2500, Nucleus Removal Set, 25 cm x 4.0 mm; (4) REF 40S2000, Nucleus Removal Straight Blade, 20 cm x 4.0 mm; (5) REF 40S2500, Nucleus Removal Straight Blade, 25 cm x 4.0 mm; (6) REF 45A2000, Nucleus Removal Angled Blade, 20 cm x 4.5 mm; (7) REF 45A2500, Nucleus Removal Angled Blade, 25 cm x 4.5 mm; (8) REF 45R2000, Nucleus Removal Set, 20 cm x 4.5 mm; (9) REF 45R2500, Nucleus Removal Set, 25 cm x 4.5 mm; (10) REF 45S2000, Nucleus Removal Straight Blade, 20 cm x 4.5 mm; (11) REF 45S2500, Nucleus Removal Straight Blade, 25 cm x 4.5 mm. | Medtronic | 2 | CDRH | 2021-10-21 |  |  | Medtronic Xomed, Inc. | There was an increase in complaints related to tip breaks and wobble/vibration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189702 | Z-0162-2022 | SIMPLEXA COVID-19 Positive Control Pack, REF MOL4160, IVD, Rx Only, UDI: (01) 20816101027051 | Simplexa Covid-19 Positive Control Pack | 2 | CDRH | 2021-10-21 |  |  | DiaSorin Molecular LLC | Due to unapproved raw materials used in the manufacturing of Covid-19 control pack, their is the potential for delayed amplification signals.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189035 | Z-0156-2022 | Medtronic sterile single-use blades that are components of the StraightShot M Series Handpiece and Fusion ENT Navigation Software, and the Midas Rex Microsaw and Triton System with the following product numbers: (1) REF 1882916, RAD 60 Sinus Blade, 11 cm x 2.9 mm; (2) REF 1883514, RAD 12 Curved Sinus Blade, 11 cm x 3.5 mm; (3) REF 1884013, Powered Tonsillectomy Blade, 11 cm x 4 mm; (4) REF 1884020, Tricut Laryngeal Blade, 22.5 cm x 4 mm; (5) REF 1884032, Tracheal Blade, 37 cm x 4 mm; (6) REF 1884033, Tracheal Blade, Angle Tip; (7) REF 1882902HRE, Serrated Blade, 11 cm X 2.9 mm, Rotatable; (8) REF 1882905HRE, Silver Bullet Blade, 11 cm x 2.9 mm, Rotatable; (9) REF 1882924HRE, Skimmer Blade, 27 cm x 2.9 mm, Rotatable; (10) REF 1882936E, Angle-Tip Serrated Blade, 18 cm x 2.9 mm; (11) REF 1882937E, Angle-Tip Serrated Blade, 22 cm x 2.9 mm; (12) REF 1882979HRE, Skimmer Blade, 13 cm x 2.9 mm, Rotatable; (13) REF 1883080HRE, Quadcut Blade, 3.0 mm, Rotatable; (14) REF 1883480HRE, Quadcut Blade, 3.4 mm, Rotatable; (15) REF 1883502HRE, Serrated Blade, 11 cm x 3.5 mm, Rotatable; (16) REF 1883506HRE, RAD 40 Blades, 11 cm x 3.5 mm, Rotatable; (17) REF 1883512HRE, RAD 12 Blade, 11 cm x 3.5 mm, Rotatable; (18) REF 1883516HRE, RAD 60 Blade, 11 cm x 3.5 mm, Rotatable; (19) REF 1884005HRE, Silver Bullet Blade, 11 cm x 4 mm, Rotatable; (20) REF 1884030HRE, Tricut Airway Blade, Angled Tip, 22 cm x 4 mm, Rotatable; (21) REF 1884031HRE, Tricut Airway Blade, Angled Tip, 27 cm x 4 mm, Rotatable; (22) REF 1884033HRE, Tricut Airway Blade, Angled Tip, 37 cm x 4 mm, Rotatable; (23) REF 1884380HR, Quadcut Blade, 4.3 mm, Rotatable; (24) REF 1884380HRE, Quadcut Blade, 4.3 mm, Rotatable. | Medtronic | 2 | CDRH | 2021-10-21 |  |  | Medtronic Xomed, Inc. | There was an increase in complaints related to tip breaks and wobble/vibration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189607 | Z-0160-2022 | REF 8420.15.010, Bone Screw/Vite - Ti6AI4V, Self Tapping/Autofilettante, Dia=6.5mm, h=20mm, STERILE R, 2100099, UDI: (01)08033390018845 | Bone Screws 6.5mm | 2 | CDRH | 2021-10-21 |  |  | Limacorporate S.p.A | There is a potential that the length of bone screws identified on labeling may not correspond to the actual length of the screw included.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189269 | Z-0157-2022 | Medtronic sterile single-use blades that are components of the StraightShot M Series Handpiece and Fusion ENT Navigation Software, and the Midas Rex Microsaw and Triton System with the following product numbers: (1) REF 1883080EM, Rotatable Fusion Blade, Quadcut 3.0 mm x 13 cm; (2) REF 1883480EM, Rotatable Fusion Blade, Quadcut, 3.4 mm x 13 cm; (3) REF 1884006EM, Rotatable Fusion Blade, RAD 40; (4) REF 1884012EM, Rotatable Fusion Blade, RAD 12; (5) REF 1884080EM, Rotatable Fusion Blade, Tricut, 13 cm; (6) REF 1884380EM, Rotatable Fusion Blade, Quadcut, 4.3 mm x 13 cm. | Medtronic | 2 | CDRH | 2021-10-21 |  |  | Medtronic Xomed, Inc. | There was an increase in complaints related to tip breaks and wobble/vibration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189606 | Z-0159-2022 | REF 8420.15.020, Bone Screw/Vite - Ti6AI4V, Self Tapping/Autofilettante, Dia=6.5mm, h=25mm, STERILE R, 2100099, UDI: (01)08033390018852 | Bone Screws 6.5mm | 2 | CDRH | 2021-10-21 |  |  | Limacorporate S.p.A | There is a potential that the length of bone screws identified on labeling may not correspond to the actual length of the screw included.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189393 | Z-0091-2022 | Cardiosave Li-Ion Battery Pack, Part No. 0146-00-0097. Component of the Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) | Cardiosave Li-Ion Battery Pack | 1 | CDRH | 2021-10-20 14:26 |  |  | Datascope Corp. | A Cardiosave Lithium-Ion Battery Pack used during protocol testing failed to meet the minimum runtime requirement per Getinge internal Product Specification. These nonconforming batteries were inadvertently released to customers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189401 | Z-0155-2022 | medfusion Stopcock, 1-way and 4-way, Model Numbers (REF): a. 081-101BE b. 081-101GE c. 081-101GEE d. 081-101NE e. 081-102NE f. 081-102RE g. 089-101E h. 888-101NE i. 081-102GEE j. 888-101BE k. 888-101GE l. 888-101GEE m. 081-101RE n. 081-102GE o. 081-102BE p. 888-101RE | Infusion Stopcock | 2 | CDRH | 2021-10-20 |  |  | Smiths Medical ASD Inc. | Product has a lack of sterility assurance
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189372 | Z-0107-2022 | CADD-Solis VIP ambulatory infusion pump, Model Numbers: a. 21-2120-0105-02L b. 21-2120-0105-03L c. 21-2120-0105-06L d. 21-2120-0105-07L e. 21-2120-0105-08L f. 21-2120-0105-12L g. 21-2120-0105-13L h. 21-2120-0105-14L i. 21-2120-0105-15L j. 21-2120-0105-17L k. 21-2120-0105-50L l. 21-2127-0105-02L m. 21-2127-0105-06L n. 21-2127-0105-50L | CADD-Solis VIP ambulatory infusion pump | 1 | CDRH | 2021-10-19 13:17 |  |  | Smiths Medical ASD Inc. | Specific Pumps may exhibit intermittent performance in the AILD (Air in Line Detector) function. If an AILD does not sufficiently discern fluid from air in line, an air in line event may not be recognized by the pump and may not alarm to notify the clinician.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188958 | Z-0144-2022 | Herniatome, 17G x 15cm, Lumbar Type with with 20ga-30cm Introducer, Model Number 5091715 | Herniatome | 2 | CDRH | 2021-10-19 08:48 |  |  | ADRIA SRL | The recalling firm was notified by the contract sterilizer that the process parameters of the sterilization cycles were not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188959 | Z-0145-2022 | Herniatome, 17G x 15cm, Lumbar Type, Model Number 5091745 | Herniatome | 2 | CDRH | 2021-10-19 08:48 |  |  | ADRIA SRL | The recalling firm was notified by the contract sterilizer that the process parameters of the sterilization cycles were not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188960 | Z-0146-2022 | Herniatome, 17G x 9cm, Cervical Type, Model Number 5091749 | Herniatome | 2 | CDRH | 2021-10-19 08:48 |  |  | ADRIA SRL | The recalling firm was notified by the contract sterilizer that the process parameters of the sterilization cycles were not followed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189506 | Z-0143-2022 | 6.5MMX165MM CANNULATED SCREW | none | 2 | CDRH | 2021-10-19 |  |  | Smith & Nephew, Inc. | The product label incorrectly indicated that the screws are fully threaded, but the screws contained in the package were partially threaded.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189122 | Z-0151-2022 | HemosIL Liquid Anti-Xa- Automated chromogenic assay for the quantitative determination of unfractionated heparin (UFH) and low molecular weight heparin (LMWH) activity in human citrated plasma on IL Coagulation Systems, (ACLTOP¿ Family, ACL TOP Family 50 Seriest and ACL Elite¿/Elite Pro) Part Number: 0020302601 | HemosIL Liquid Anti-Xa PN 0020302601: not FDA-cleared; not sold in the US | 2 | CDRH | 2021-10-19 |  |  | Instrumentation Laboratory | Labeled On-board instrument stability issue for current and future lots, reduced On-board Instrument Stability from 7 days to 5 days
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189424 | Z-0152-2022 | smiths medical portex Loss of Resistance Device, REF 100/398/000 | smiths medical portex Loss of Resistance Device | 2 | CDRH | 2021-10-19 |  |  | Smiths Medical ASD Inc. | The labeling was missing information on sterilization and prohibition of reuse.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189119 | Z-0149-2022 | HemosIL Liquid Anti-Xa- Automated chromogenic assay for the quantitative determination of unfractionated heparin (UFH) and low molecular weight heparin (LMWH) activity in human citrated plasma on IL Coagulation Systems, (ACL TOP¿ Family, ACL TOP Family 50 Seriest and ACL Elite¿/Elite Pro/8/9/10000) Part Number: 0020302600 | HemosIL Liquid Anti-Xa | 2 | CDRH | 2021-10-19 |  |  | Instrumentation Laboratory | Labeled On-board instrument stability issue for current and future lots, reduced On-board Instrument Stability from 7 days to 5 days
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189300 | Z-0142-2022 | Canon PET-CT SCANNER, Cartesion Prime, MODEL PCD-1000A | Canon PET-CT SCANNER, Cartesion Prime | 2 | CDRH | 2021-10-19 |  |  | Canon Medical System, USA, INC. | A software problem has been identified which could result in the diagnostic imaging system not proceeding to the next actual scan even though automatic start of the next scan is specified. This could result in the diagnostic imaging system failing resulting in rescanning and reinjection of contrast medium.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189275 | Z-0154-2022 | Catalog # BEK3411, Ecomomy Jamshidi Bone Marrow Biopsy Tray, 11 G X 4"Biopsy/Aspiration needle, sterile, Qty: 10, Single Use Only, RX only. Exp. Date Lidocaine Vendor lot. CareFusion Intended use to aspirate marrow from either the sternum or the iliac crest. Intraosseous Infusion is recommended only in life-threatening emergencies where vascular access is critical and conventional forms of access are not available or are unobtainable. | Economy Jamshidi Bone Marrow Biopsy and Aspiration Tray; 2) Economy Jamshidi¿ Bone Marrow Biopsy Tra | 2 | CDRH | 2021-10-19 |  |  | Bard Peripheral Vascular Inc | Due to misbranded products shipped to customers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189408 | Z-0148-2022 | CADD Administration Sets, Model #21-7346-24 | CADD Administration Set | 2 | CDRH | 2021-10-19 |  |  | Smiths Medical ASD Inc. | Product may have been manufactured with an air filter assembled in an incorrect orientation.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189294 | Z-0147-2022 | Kit: Central Line Dres Change 20/Cs | Central Line Dressing Change Kit | 2 | CDRH | 2021-10-19 |  |  | Medical Action Industries, Inc. 306 | The ChloraPrep 3ml kit component has a shorter expiry date than the labeled kit expiry date.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189256 | Z-0153-2022 | Catalog # BEK4511, Qty: 10, Ecomomy Jamshidi Bone Marrow Biopsy and Aspiration Tray, 11 G X 4"Biopsy/Aspiration needle, 15 G Aspiration Needle (Adjustable Length 15/16" -1-7/8") (24mm-48mm) sterile, For Single Use Only, RX only. Exp. Date Lidocaine Vendor lot. CareFusion. Original Jamshidi" Bone Marrow Biopsy/Aspiration Needle: - also referred to as Jamshidi needle Intended use for the posterior iliac crest biopsy technique | Economy Jamshidi Bone Marrow Biopsy and Aspiration Tray | 2 | CDRH | 2021-10-19 |  |  | Bard Peripheral Vascular Inc | Due to misbranded products shipped to customers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189121 | Z-0150-2022 | HemosIL Liquid Anti-Xa- automated chromogenic assay for in vitro diagnostic use by laboratory professionals in clinical laboratories. The assay provides quantitative results on 3.2% citrated human plasma Part Number: 0020302602 | HemosIL Liquid Anti-Xa | 2 | CDRH | 2021-10-19 |  |  | Instrumentation Laboratory | Labeled On-board instrument stability issue for current and future lots, reduced On-board Instrument Stability from 7 days to 5 days
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189189 | Z-0130-2022 | Foundation Kit | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189205 | Z-0132-2022 | SIGNA Premier nuclear magnetic resonance imaging system | SIGNA Premier | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189209 | Z-0136-2022 | Discovery MR750 3.0T nuclear magnetic resonance imaging system | Discovery MR750 3.0T | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189208 | Z-0135-2022 | Discovery MR750w 3.0T nuclear magnetic resonance imaging system | Discovery MR750w 3.0T | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189212 | Z-0139-2022 | SIGNA Artist nuclear magnetic resonance imaging system | SIGNA Artist | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189542 | Z-0124-2022 | Atellica IM 1300 Analyzer, SMN 11066001 | Atellica IM 1300 Analyzer | 2 | CDRH | 2021-10-16 |  |  | Siemens Healthcare Diagnostics, Inc | Ancillary reagent packs used on Atellica IM 1300 and Atellica IM 1600 analyzers may overflow when pierced while on an analyzer located at an altitude of greater than 350m (1148ft) above sea level. As a result, customers at these altitudes may observe an increased number of Reagent Volume Check Errors.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189187 | Z-0128-2022 | IUD Insertion Kit | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189211 | Z-0138-2022 | SIGNA Voyager nuclear magnetic resonance imaging system | SIGNA Voyager | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189185 | Z-0126-2022 | Tenaculum, Plastic | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189214 | Z-0141-2022 | 1.5T SIGNA HDxt MR System nuclear magnetic resonance imaging system | 1.5T SIGNA HDxt MR System | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189186 | Z-0127-2022 | IUD Insertion Pack | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189213 | Z-0140-2022 | Optima MR450w 1.5T nuclear magnetic resonance imaging system | Optima | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189206 | Z-0133-2022 | SIGNA Architect nuclear magnetic resonance imaging system | SIGNA Architect | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189207 | Z-0134-2022 | SIGNA Pioneer nuclear magnetic resonance imaging system | SIGNA Pioneer | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189210 | Z-0137-2022 | SIGNA Creator, SIGNA Explorer nuclear magnetic resonance imaging system | SIGNA Creator, SIGNA Explorer | 2 | CDRH | 2021-10-16 | 2021-10-16 21:06 |  | GE Healthcare, LLC | Under certain conditions, some slices may be missing which can lead to a gap of anatomy in the 3D Volume images.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189543 | Z-0125-2022 | Atellica IM 1600 Analyzer, SMN 11066000 | Atellica IM 1600 Analyzer | 2 | CDRH | 2021-10-16 |  |  | Siemens Healthcare Diagnostics, Inc | Ancillary reagent packs used on Atellica IM 1300 and Atellica IM 1600 analyzers may overflow when pierced while on an analyzer located at an altitude of greater than 350m (1148ft) above sea level. As a result, customers at these altitudes may observe an increased number of Reagent Volume Check Errors.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189188 | Z-0129-2022 | Endometrial Biopsy Kit | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189190 | Z-0131-2022 | Standard IUD Insertion Kit | Stradis | 2 | CDRH | 2021-10-16 |  |  | Stradis Medical, LLC dba Stradis Healthcare | Firm unable to confirm that the tenaculum devices were produced in accordance with required design and manufacturing standards.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189120 | Z-0121-2022 | P2131 Automated Field Steam Sterilizer NSN: 6530-01-641-4641 | P2131 Automated Field Steam Sterilizer | 2 | CDRH | 2021-10-15 |  |  | Fort Defiance Industries, LLC | The rear heater plate and cover assembly that contains the electrical wiring for the heaters does not have sufficient bonding to ground.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189033 | Z-0122-2022 | MPR Slotted Mallet, labeling etched onto devices as follows: Medical Products Resource 800.524.5194 www.m-p-r.com MPR-1120-1000 (01)00840001405868 | Slotted Mallet | 2 | CDRH | 2021-10-15 |  |  | Implant Resource Inc | Potential for mallet head to become separated from mallet handle. There is also significant deformations on striking surface that were not expected
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189420 | Z-0120-2022 | VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Quantitative Controls (IVD)- For use in monitoring the performance of the VITROS ECi/ECiQ/3600 Immunodiagnostic Systems and the VITROS 5600/XT 7600 Integrated Systems when used with the VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Quantitative Reagent Pack used in combination with the VITROS Immunodiagnostic Products Anti- SARS-CoV-2 IgG Quantitative Calibrator for the determination of IgG antibodies to SARS-CoV-2. Product Code: 6199966 | VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Quantitative Controls | 2 | CDRH | 2021-10-15 |  |  | Ortho Clinical Diagnostics Inc | VITROS Immunodiagnostic Products Anti-SARS-CoV-2 IgG Quantitative Controls Outside of Expected Control Range, potential for delayed results
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189406 | Z-0123-2022 | EPIQ CVxi Diagnostic Ultrasound Systems, Product Code: 795232 | EPIQ CVxi Diagnostic Ultrasound Systems | 2 | CDRH | 2021-10-15 |  |  | Philips Ultrasound Inc | Ultrasound system software issue can cause an EchoNavigator error notification, leading to an unresponsive system, necessitating system restart, which could lead to therapy/treatment delay and/or unnecessary therapy/treatment. Can occur during ultrasound-guided TEE with EchoNavigator, if secondary screen capture is enabled and user presses other system buttons before secondary capture completed.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189115 | Z-0006-2022 | Alinity m SARS-CoV-2 Amp Kit (Non-US Assay) | Alinity m SARS-CoV-2 Amp Kit (Non-US Assay) | 1 | CDRH | 2021-10-14 |  |  | Abbott Molecular, Inc. | There is a potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189118 | Z-0008-2022 | Alinity m Resp-4-Plex AMP Kit (Non-US Assay) | Alinity m Resp-4-Plex AMP Kit (Non-US Assay) | 1 | CDRH | 2021-10-14 |  |  | Abbott Molecular, Inc. | There is a potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189093 | Z-0117-2022 | Power Processor 3K and 3K High Speed Stockyards Module Nameplate: PP 3K Stockyard (Refrig) 9707 REF A88116 | Power Processor 3K and 3K High Speed Stockyards | 2 | CDRH | 2021-10-14 |  |  | Beckman Coulter Inc. | Missing splash guard which may expose users to hazardous materials. .
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189114 | Z-0005-2022 | Alinity m Resp-4-Plex AMP Kit (Emergency Use Authorization, US Assay) | Alinity m Resp-4-Plex AMP Kit Assay | 1 | CDRH | 2021-10-14 |  |  | Abbott Molecular, Inc. | There is a potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189116 | Z-0007-2022 | Alinity m SARS-CoV-2 Amp Kit (Non-US Assay) | Alinity m SARS-CoV-2 Amp Kit (Non-US Assay) | 1 | CDRH | 2021-10-14 |  |  | Abbott Molecular, Inc. | There is a potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189278 | Z-0116-2022 | Elekta Medical Linear Accelerator (EMLA), configured as follows: a. Elekta Precise Digital Accelerator; b. Elekta Harmony; c. Elekta Synergy; d. Elekta Infinity; e. Versa HD | Elekta Medical Linear Accelerator (EMLA), | 2 | CDRH | 2021-10-14 |  |  | Elekta, Inc. | There is a risk that the Diode D1 in the Modulator can fail in a destructive manner when removed from the machine, and that the protective tube surrounding the diode may fail to contain the ejected matter.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189113 | Z-0004-2022 | Alinity m SARS-CoV-2 Amp Kit (Emergency Use Authorization, US Assay) | Alinity m SARS-CoV-2 Amp Kit Assay | 1 | CDRH | 2021-10-14 |  |  | Abbott Molecular, Inc. | There is a potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189498 | Z-0115-2022 | Elekta Medical Linear Accelerator (EMLA): Elekta Harmony Pro, REF 1555143; Elekta Infinity, REF 1552091, and Elekta Versa HD, REF 1552093. | Elekta Harmony Pro, Elekta Infinity, and Elekta Versa HD | 2 | CDRH | 2021-10-14 |  |  | ELEKTA SOLUTIONS AB | There is a risk that the Diode D1 in the Modulator can fail in a destructive manner when removed from the machine, and that the protective tube surrounding the diode may fail to contain the ejected matter.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188906 | Z-0106-2022 | Ion Flexison Biopsy Needle, 19G, Extension Range : 0-3 cm, Ref: 490104-09 and 49004-10, 5-pack, Sterile EO, Rx only, UDI:00886874115978 | Ion Endoluminal System; Flexision Biopsy Needle | 2 | CDRH | 2021-10-13 15:18 |  |  | Intuitive Surgical, Inc. | Potential for the sheath tip on the biopsy needle to separate from the sheath shaft.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189095 | Z-0105-2022 | Thoraguard Chest Tube Kit, 20 Fr | Thoraguard Chest Tube Kit | 2 | CDRH | 2021-10-13 14:38 |  |  | Centese Inc. | An adhesive joint in the SmartValve component of the device can become compromised resulting in a leak.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189222 | Z-0096-2022 | smiths medical medex CBSS 5 ml, dreifach proximaler Entnahmeport 60 cm, Model number DPSHC0082 | smiths medical medex CBSS 5 ml, dreifach proximaler Entnahmeport | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189220 | Z-0094-2022 | smiths medical medex CBSS 2-fach Druckmesset mit Option, Model Number DPSHC0076 | smiths medical medex CBSS 2-fach Druckmesset mit Option | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189227 | Z-0101-2022 | smiths medical medex CBSS-5ML, Triple set, (172/15cm), Model Number DPSHC0092 | smiths medical medex CBSS-5ML, Triple set | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189228 | Z-0102-2022 | smiths medical medex CBSS-5ML, Single Set, (172/15cm), Model Number DPSHC0093 | smiths medical medex CBSS-5ML, Single Set | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189221 | Z-0095-2022 | smiths medical medex CBSS 5 ml, Antishunt proximaler Entnahmeport 60 cm, Model Number DPSHC0081 | smiths medical medex CBSS 5 ml, Antishunt proximaler Entnahmeport | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189226 | Z-0100-2022 | smiths medical medex CBSS-5ML Antishunt set, (172/15cm), Model Number DPSHC0091 | smiths medical medex CBSS-5ML Antishunt set | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189224 | Z-0098-2022 | smiths medical medex CBSS 5 ml, zweifach proximaler Entnahmeport 60 cm, Model Number DPSHC0084 | smiths medical medex CBSS 5 ml, zweifach proximaler Entnahmeport | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189217 | Z-0092-2022 | medex LogiCal 5mL Closed Blood Sampling System, Model Numbers: a) MX961E181P1 b) MX961E211P1 c) MX962E211P1 d) MX96AE181P1 e) MX96AE211P1 | medex LogiCal 5mL Closed Blood Sampling System | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189229 | Z-0103-2022 | smiths medical medex CBSS-5ML, Double set, (172+15cm), Model Number DPSHC0094 | smiths medical medex CBSS-5ML, Double set | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189225 | Z-0099-2022 | smiths medical medex CBSS-5ML Zweifach set, Model Number DPSHC0090 | smiths medical medex CBSS-5ML Zweifach set | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189219 | Z-0093-2022 | smiths medical medex HemoDraw Plus Closed Blood Sampling Set with LogiCal Pressure Monitoring System, Model Numbers: a) MX961H182P1 b) MX961H212P1 c) MX962H182P1 d) MX962H212P1 | smiths medical medex HemoDraw Plus Closed Blood Sampling Set with LogiCal Pressure Monitoring System | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189223 | Z-0097-2022 | smiths medical medex CBSS 5 ml, einfach proximaler Entnahmeport 60 cm, Model Number DPSHC0083 | smiths medical medex CBSS 5 ml, einfach proximaler Entnahmeport | 2 | CDRH | 2021-10-13 14:14 |  |  | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189232 | Z-0113-2022 | Flower Drill Bit Kit 2.7mm- The Drill Bit Kit, 2.7mm (DBK 027) is intended to be used with solid 2.7mm screws and can be used with any the Flower Small and Medium Implant Set Flower Ankle Plating Set , Flower Rear Foot/Calcaneus Plating Set or Flower Upper Extremity Plating Set and uses the medium locking mechanism. Catalog Number: DBK 027 | Flower Drill Bit Kit 2.7mm | 2 | CDRH | 2021-10-13 |  |  | Flower Orthopedics Corporation | Drill bits packaged in kits are missing the AO connect feature and may result in delay of procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189404 | Z-0090-2022 | EcoGel 100 - RED Medical Supplies Green Ultrasound Gel, Model No. 250UG-GLU | EcoGel 100 - RED Medical Supplies Green Ultrasound Gel | 1 | CDRH | 2021-10-13 |  |  | Red Medical Supplies | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021,FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189149 | Z-0068-2022 | ECO 50TS Gel Tens Ultrasound 1in X 5L cube 1 Dispenser Btl | ECO | 1 | CDRH | 2021-10-13 |  |  | National Distribution & Contracting Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189148 | Z-0067-2022 | ECO 85CB Lotion Ultrasound 8.5 oz 24/cs ECO 40JBC Gel Ultrasound Transmission Clr 8.5 oz BTL 24/cs ECO 40JMC Gel Ultrasound Transmission Clr 5L 4/cs NDC P301008 ProAdvantage Gel Ultrasound Transmission 8.5 oz 12/cs NDC P301105 ProAdvantage Gel Ultrasound Transmission 5 Liter 4/cs NDC P301115 ProAdvantage Ultrasound Lotion 5L W/Empty Btl 4/cs ECO 30GS Ecogel 100 Ultrasound Grn 1in X 5L Cube 1 Dispenser Btl ECO 30GB Ecogel 100 Ultrasound Grn 8.5 oz Btl 24/cs | Multiple | 1 | CDRH | 2021-10-13 |  |  | National Distribution & Contracting Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189133 | Z-0108-2022 | Evita V500 Ventilator with Installed CO2 Measurement Option-Intended for the ventilation of adult,pediatric and neonatal patients. Catalog Number: 8416400 | Draeger | 2 | CDRH | 2021-10-13 |  |  | Draeger Medical, Inc. | Software 2.51 and Lower with Installed CO2 Measurement Option, may result in Restart of the Ventilator, brief cessation of ventilation can occur and loss of PEEP for approx. 8 seconds
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189402 | Z-0088-2022 | EcoGel 200 - RED Medical Supplies Blue Ultrasound Gel, Model No. 250UG-BMU | EcoGel 200 - RED Medical Supplies Blue Ultrasound Gel | 1 | CDRH | 2021-10-13 |  |  | Red Medical Supplies | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021,FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189233 | Z-0114-2022 | Flower Lag Screw Kit 2.7 mm-intended to be used with 2.7mm lag screws and can be used with any plate that as part of the Flower Small and Medium Implant Set, Flower Ankle Plating Set , Flower Rear Foot/Calcaneus Plating Set , or Flower Upper Extremity Plating Set and uses the medium locking mechanism. Catalog Number: LSK 027 | Flower Lag Screw Kit 2.7mm | 2 | CDRH | 2021-10-13 |  |  | Flower Orthopedics Corporation | Drill bits packaged in kits are missing the AO connect feature and may result in delay of procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189134 | Z-0109-2022 | Babylog VN500 Ventilator with Installed CO2 Measurement Option-intended for the ventilation of neonatal patients from 0.4 kg (0.88 lbs) up to 10 kg (22 lbs) and pediatric patients from 5 kg (11 lbs) up to 20 kg (44 lbs) bodyweight. Catalog Number: 8417400 | Draeger | 2 | CDRH | 2021-10-13 |  |  | Draeger Medical, Inc. | Software 2.51 and Lower with Installed CO2 Measurement Option, may result in Restart of the Ventilator, brief cessation of ventilation can occur and loss of PEEP for approx. 8 seconds
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189126 | Z-0104-2022 | Chesapeake Anterior Lumbar (AL) Removal Tool; Cat. no. 2008-90068 | Chesapeake Anterior Lumbar (AL) Removal Tool | 2 | CDRH | 2021-10-13 |  |  | K2M, Inc | Units may arrive with missing pins and/or in a state in disassembly. Pins dislodging from the instrument during a procedure may enter the surgical cavity, requiring intraoperative removal or revision surgery.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189403 | Z-0089-2022 | EcoGel 200 - RED Medical Supplies Clear Ultrasound Gel, Model No. 250UG-CMU | EcoGel 200 - RED Medical Supplies Clear Ultrasound Gel | 1 | CDRH | 2021-10-13 |  |  | Red Medical Supplies | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021,FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189301 | Z-0110-2022 | United Orthopedics USTAR II Knee System USTAR II Knee System Segment Part, RHS, 25 mm Length, Catalog # 2915-1025 | United Orthopedics USTAR II Knee System | 2 | CDRH | 2021-10-13 |  |  | Uoc Usa Inc | Due to product not meeting specification for concentricity and failing to meet requirements of dynamic fatigue testing.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189230 | Z-0111-2022 | Flower Drill Bit Kit 2.0 mm - intended to be used with solid 2.0mm screws and can be used with any plate that has been cleared as part of the Flower Small and Medium Implant Set and uses the small locking mechanism. mechanism. Catalog Number: DBK 020 2013002721 | Flower Drill Bit Kit 2.0mm | 2 | CDRH | 2021-10-13 |  |  | Flower Orthopedics Corporation | Drill bits packaged in kits are missing the AO connect feature and may result in delay of procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189231 | Z-0112-2022 | Flower Drill Bit Kit 2.4 mm-intended to be used with solid 2.4mm screws and can be used with any plate as part of the Flower Small and Medium Implant Set and uses the small locking mechanism. Catalog Number: DBK 024 | Flower Drill Bit Kit 2.4mm | 2 | CDRH | 2021-10-13 |  |  | Flower Orthopedics Corporation | Drill bits packaged in kits are missing the AO connect feature and may result in delay of procedure
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189289 | Z-0085-2022 | Endoskeleton" TAS Interbody System with Titan nanoLOCK" Surface Technology, Nanolock 12 Dg Interbody, Large, 12mm | Endoskeleton" TAS Interbody System with Titan nanoLOCK" Surface Technology | 2 | CDRH | 2021-10-12 |  |  | Medtronic Sofamor Danek USA Inc | Engraving on implants may not match what is listed on the pouch labels.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189290 | Z-0086-2022 | Endoskeleton" TCS Interbody System with Titan nanoLOCK" Surface Technology, NanoLock, 6 Deg, Interbody, Medium, 10mm | Endoskeleton" TCS Interbody System with Titan nanoLOCK" Surface Technology | 2 | CDRH | 2021-10-12 |  |  | Medtronic Sofamor Danek USA Inc | Engraving on implants may not match what is listed on the pouch labels.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189370 | Z-0087-2022 | FlexLab (FLX), Accelerator a3600 (ACP), Aptio Automation (AP2). (Different brand names of same product.) | FlexLab (FLX), Accelerator a3600 (ACP), Aptio Automation (AP2) | 2 | CDRH | 2021-10-12 |  |  | Inpeco S.A. | The Firmware (FW) of the Automation System Interface Module to ADVIA Centaur XP/XPT may lead to an incorrect association of test results to sample ID. Improper management of the sample may bring an incorrect patient result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189141 | Z-0079-2022 | Philips, HeartStart FRx + Defibrillator, Model # Model #s: 861304 | Philips, HeartStart FRx + Defibrillator | 2 | CDRH | 2021-10-08 | 2021-10-08 21:06 |  | Philips North America LLC | Customers were not notified of previous recalls associated with various defibrillator models.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189144 | Z-0081-2022 | Philips, HeartStart HS1 OnSite Defibrillator, Model #'s: M5066A and M5067A | Philips, HeartStart HS1 OnSite Defibrillator | 2 | CDRH | 2021-10-08 | 2021-10-08 21:06 |  | Philips North America LLC | Customers were not notified of previous recalls associated with various defibrillator models.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189143 | Z-0080-2022 | Philips, HeartStart HS1 Home Defibrillator, Model # Model number: M5068A | Philips, HeartStart HS1 Home Defibrillator | 2 | CDRH | 2021-10-08 | 2021-10-08 21:06 |  | Philips North America LLC | Customers were not notified of previous recalls associated with various defibrillator models.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189083 | Z-0083-2022 | cobas EGFR Mutation Test v2-Somatic gene mutation detection system Catalog Number :07248563190 | Roche | 2 | CDRH | 2021-10-08 |  |  | Roche Molecular Systems, Inc. | False Mutation Detected results for the exon 20 insertion (Ex20Ins) mutation when using the cobas EGFR Mutation Test v2, may lead to inappropriate treatment
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189140 | Z-0078-2022 | Philips, HeartStart FR2 + Defibrillator, Model # M3860A and M3861A | Philips, HeartStart FR2 + Defibrillator | 2 | CDRH | 2021-10-08 | 2021-10-08 21:06 |  | Philips North America LLC | Customers were not notified of previous recalls associated with various defibrillator models.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189192 | Z-0084-2022 | MYLA Versions: 4.7, 4.7.1, 4.8, 4.8.1 and 4.8.2, part of the VIRTUO System | MYLA software v. 4.7, 4.7.1, 4.8, 4.8.1, 4.8.2 | 2 | CDRH | 2021-10-08 |  |  | BioMerieux SA | Under certain conditions, there is a risk for a false negative result.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189396 | Z-0077-2022 | Heparin, GORE VIABAHN VBX, BALLOON EXPANDABLE ENDOPROTHESIS, REF BXAL085902E, 8 MM X 59 MM 8 Fr, LARGE BALLOON EXPANDLBE, Nominal Stent Diameter 8 mm, Max Post-Dilated Stent Diameter 16 mm, UDI: | GORE VIABAHN VBX, BALLOON EXPANDABLE ENDOPROTHESIS | 2 | CDRH | 2021-10-08 |  |  | W L Gore & Associates, Inc. | Incorrect labeling, specifically, expandable endoprosthesis labeled as 8 mm x 59mm x 135 cm, however may contain a 9mm x 29 mm x 135 cm device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189173 | Z-0082-2022 | BioPlex 2200 ANA Screen Control Set, REF: 663-1131 incompatible with ANA Screen Reagent Packs/ANA Kits, REF: 665-1150; and ANA Screen Calibrator Sets, REF: 663-1101 | BioPlex 2200 ANA Screen Control Set | 2 | CDRH | 2021-10-08 |  |  | Bio-Rad Laboratories, Inc. | A new lot of ANA screen control and any subsequent control lots are incompatible with older lot numbers of ANA calibrators and ANA reagent packs.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189395 | Z-0074-2022 | ATTUNE Revision Tibial Augment Size 7/8 5mm- intended for knee joint replacement Part Code:152307001 | ATTUNE Revision Tibial Augment Size 7/8 5mm | 2 | CDRH | 2021-10-07 |  |  | DePuy Orthopaedics, Inc. | Packaging may include screws that are not intended to be used with the 5mm augments, may cause delay in surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189165 | Z-0075-2022 | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP), Model Nos. 0998-00-0800-31, 0998-00-0800-32, 0998-00-0800-33, 0998-00-0800-34, 0998-00-0800-35, 0998-00-0800-45, 0998-00-0800-52, 0998-00-0800-53, 0998-00-0800-55, 0998-00-0800-65 | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) | 2 | CDRH | 2021-10-07 |  |  | Datascope Corp. | The Helium Indicator on the Cardiosave display may over represent the amount of helium in unit in the presence of Electromagnetic Interference.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189111 | Z-0070-2022 | Philips Ultrasound, Inc. Sparq, Diagnostic Ultrasound System, REF 795090, 100-240V -, 50/60 Hz, 500 VA, CE2797, Rx Only, | Philips Ultrasound, Inc. Sparq, Diagnostic Ultrasound System | 2 | CDRH | 2021-10-07 |  |  | Philips Ultrasound Inc | Battery system data issue with the ultrasound system can intermittently cause a system shutdown, regardless of actual battery state or application of alternating current power.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189103 | Z-0069-2022 | CardioQuip, Modular Cooler-Heater 1000(i), Refrigeration Module, Model: MCH-10RMS, UDI: (01)00860000846127 used with heat exchange devices. | Modular Cooler-Heater 1000(i), Refrigeration Module (MCH-10RMS) | 2 | CDRH | 2021-10-07 |  |  | CardioQuip, LLC | There is the potential bacterial including Nontuberculous mycobacterial (NTM) to grow in the water system of heater-cooler devices which could result in contaminated water potentially aerosolized into the operating room during surgery which could result in patient infection.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189112 | Z-0071-2022 | Philips Ultrasound, Inc. Sparq, Diagnostic Ultrasound System, REF 795116, 100-240V -, 50/60 Hz, 500 VA, CE2797, Rx Only, | Philips Ultrasound, Inc. Sparq, Diagnostic Ultrasound System | 2 | CDRH | 2021-10-07 |  |  | Philips Ultrasound Inc | Battery system data issue with the ultrasound system can intermittently cause a system shutdown, regardless of actual battery state or application of alternating current power.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189139 | Z-0072-2022 | ARCH Leg Positioning System | ARCH Leg Positioning System | 2 | CDRH | 2021-10-07 |  |  | Innovative Orthopedic Technologies, LLC | Due to the holes not being bored to the specific diameter, the locking pins are unable to engage under a load.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189166 | Z-0076-2022 | Cardiosave Rescue Intra-Aortic Balloon Pump (IABP), Model Nos. 0998-00-0800-75, 0998-00-0800-83, 0998-00-0800-85 | Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | 2 | CDRH | 2021-10-07 |  |  | Datascope Corp. | The Helium Indicator on the Cardiosave display may over represent the amount of helium in unit in the presence of Electromagnetic Interference.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189276 | Z-0073-2022 | AngioDynamics NanoKnife Disposable Single Electrode Probes, Irreversible Electroporation (IRE) Device, 15 cm IRE Single Electrode RFID Activation, PG. Catalog No. 20400103 | AngioDynamics NanoKnife Disposable Single Electrode Probes | 2 | CDRH | 2021-10-07 |  |  | Angiodynamics, Inc. | Programming issue affected RFID function of a single lot and did not allow the NanoKnife probes to be recognized by the NanoKnife generator.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189239 | Z-0043-2022 | ZOLL, REF 8700-0793-40 (SL-2593AE) SOLEX 7, Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), HEPARIN COATED, Sterile EO, Rx Only, UDI: (00)849111075312 | SOLEX 7, Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189243 | Z-0046-2022 | ZOLL, REF 8700-0787-40 (IC-3893AE), ICY Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), STERILE EO, Rx Only, UDI: (01)00849111075220 | ICY Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189104 | Z-0035-2022 | CardioQuip, Modular Cooler-Heater, Model: MCH-1000(i), CE 2797, UDI: (01)00860000846103 | CardioQuip, Modular Cooler-Heater, | 2 | CDRH | 2021-10-06 | 2021-10-06 21:08 |  | CardioQuip, LLC | Labeling does not include guidance for proper water-quality maintenance and device inspection for heater cooler devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189441 | Z-0060-2022 | Zoll, REF 8700-000866-01, ICY IVTM Disposable Pack and Start UP Kit (6 ft.) | ICY IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189237 | Z-0041-2022 | ZOLL, REF: 8700-0781-01 (CL-2295), COOL LINE, Intravascular Heat Exchange Catheter Kit Applause Custom Luer Heparin, STERILE EO, Rx only, UDI: (01)00849111075190 | COOL LINE INTRAVASCULAR HEAT EXCHANGE CATHETER KIT APPLAUSE CUSTOM LUE HEPARIN | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189235 | Z-0039-2022 | Zoll REF: 8700-0782-01 (IC-3893), ICY Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens) , Femoral vein Insertion Only, Sterile EO, Rx Only, UDI: (01)00849111075237 | ICY Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189105 | Z-0036-2022 | CardioQuip, Modular Cooler-Heater 1000(i), Standard Unit, Model: MCH-10RMS, CE 2797, UDI: (01) 00860000846127 | CardioQuip, Modular Cooler-Heater, MCH-10RMS | 2 | CDRH | 2021-10-06 | 2021-10-06 21:08 |  | CardioQuip, LLC | Labeling does not include guidance for proper water-quality maintenance and device inspection for heater cooler devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189247 | Z-0047-2022 | Zoll REF: 8700-0782-40 (IC-3893AE), ICY Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens),HEPARIN COATED, Sterile EO, Rx Only, UDI: (01)00849111075251 | ICY Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189253 | Z-0051-2022 | Zoll REF: 8700-0783-14 (IC-4593CO), QUATTRO Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), HEPARIN COATED, Sterile EO, Rx Only, UDI: (01)00849111075534 | QUATTRO Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189251 | Z-0049-2022 | Zoll REF: 8700-0783-40 (IC-4593AE), QUATTRO Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), HEPARIN COATED, Sterile EO, Rx Only, UDI: (01)00849111075275 | QUATTRO Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189410 | Z-0053-2022 | Zoll, REF: 8700-0783-03 (IC-4593AE) QUATTRO Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens), Sterile EO, Rx Only, UDI: (01)00849111075305 | QUATTRO Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189254 | Z-0052-2022 | Zoll REF: 8700-0793-14 (SL-2593CO), SOLEX 7 Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), Heparin COATED, Sterile EO, Rx Only, UDI: (01)00849111075541 | Solex 7 Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189242 | Z-0045-2022 | ZOLL, REF 8700-0781-40 (CL-2295AE), COOL LINE Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens) HEPARIN COATED, STERILE EO, Rx Only, UDI: (01)00849111075213 | COOL LINE Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189412 | Z-0055-2022 | Zoll, REF 8700-0782-03 (IC-3893AE), ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens), Sterile EO, Rx Only, UDI: (01)00849111050180 | ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189434 | Z-0056-2022 | Zoll, REF 8700-0695-40 (IC-3893AE), ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens), Sterile EO, Rx Only, UDI: (01)00849111075077 | ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189442 | Z-0061-2022 | Zoll, REF 8700-000867-01, Quattro IVTM Disposable Pack and Start UP Kit (6 ft.) | Quattro IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189446 | Z-0065-2022 | Zoll, REF 8700-000871-40, Quattro (Heparin) IVTM Disposable Pack and Start UP Kit (6 ft.), EU, New Luer | Quattro (Heparin) IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189236 | Z-0040-2022 | Zoll, REF: 8700-0791-01 (IC-4593), Premium Access Kit QUATTRO, Intravascular Heat exchange Catheter Kit (Heparin), with central venous infusion capabilities (3 lumens), Femoral Vein Insertion Only, Sterile EO, Rx only, UDI: (01)008491110 7550 3 | QUATTRO, Intravascular Heat Exchange | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189110 | Z-0029-2022 | Action De Gala, FuSion Conductive Gel - Product Usage: Intended for use in acoustic coupling of ultrasound transducer to tissue. | Ecogel 200 Ultrasound gel (Action De Gala, FuSion Conductive Gel) | 1 | CDRH | 2021-10-06 |  |  | Aesthetics Systems Usa Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189252 | Z-0050-2022 | Zoll REF: 8700-0782-14 (IC-3893CO), ICY Intravascular Heat Exchange Catheter Kit with central venous infusion capabilities (3 lumens), HEPARIN COATED, Sterile EO, Rx Only, UDI: (01)00849111075527 | ICY Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189447 | Z-0066-2022 | Zoll, REF 8700-000872-40, SOLEX 7(Heparin) IVTM Disposable Pack and Start UP Kit (6 ft.), EU, New Luer | Quattro (Heparin) IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189400 | Z-0034-2022 | Flower E-Kit, Advanced-Bone Fixation Fasterner-Fixation of Ref: EWK 201 | Coventus Flower | 2 | CDRH | 2021-10-06 |  |  | Flower Orthopedics Corporation | Product kit missing CDG 200 (Cannulated Countersink for 3.0-4.5mm Screws) from EWK 201 (Flower E-Kit, Advanced)
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189277 | Z-0033-2022 | Alinity c CRP Vario Cardiac High Sensitivity Calibrator Kit (ref 07P5605) - Product Usage: use in the calibration of the Alinity c CRP Vario assay, Cardiac CRP high sensitivity method [Cardiac CRP], on the Alinity c analyzer. | Alinity c CRP Vario Cardiac High Sensitivity Calibrator Kit | 3 | CDRH | 2021-10-06 |  |  | Sentinel CH SpA | Calibrator vials were incorrectly labeled. Use of incorrectly labeled vials will cause the instrument to generate a warning that calibrators are incorrectly loaded and calibration will not be completed successfully.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189435 | Z-0057-2022 | Zoll, REF 8700-0657-40 (IC-3893AE), ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens), Heparin, Sterile EO, Rx Only, UDI: (01)00849111075107 | ICY Intravascular Heat Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189444 | Z-0063-2022 | Zoll, REF 8700-000869-40, Cool Line (Heparin) IVTM Disposable Pack and Start UP Kit (6 ft.), EU, New Luer | Cool Line (Heparin) IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189443 | Z-0062-2022 | Zoll, REF 8700-000868-01, Solex 7 IVTM Disposable Pack and Start UP Kit (6 ft.) | Solex 7 IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189436 | Z-0058-2022 | Zoll, REF 8700-000766-01, Quattro IVTM Premium Disposable Pack and Start UP Kit (6 ft.) | Quattro IVTM Premium Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189106 | Z-0037-2022 | CardioQuip, Modular Cooler-Heater 1000(m),Compact Unit, Model:MCH-1000(m),CE 2797, UDI: (01)00860000846110 | CardioQuip, Modular Cooler-Heater, Model: MCH-1000m | 2 | CDRH | 2021-10-06 | 2021-10-06 21:08 |  | CardioQuip, LLC | Labeling does not include guidance for proper water-quality maintenance and device inspection for heater cooler devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189107 | Z-0038-2022 | CardioQuip, Modular Cooler-Heater 1000(m),Thermoelectric Cooling Lid, Model: MCH-11TEC,CE 2797, UDI: (01) 00860000846134 | CardioQuip, Modular Cooler-Heater 1000(m),Thermoelectric Cooling Lid, Model: MCH-11TEC | 2 | CDRH | 2021-10-06 | 2021-10-06 21:08 |  | CardioQuip, LLC | Labeling does not include guidance for proper water-quality maintenance and device inspection for heater cooler devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189250 | Z-0048-2022 | Zoll REF: 8700-0783-01 (IC-4593), QUATTRO Intravascular Heat Exchange Catheter Kit Custom Luer Applause Heparin, Sterile EO, Rx Only, UDI: (01)00849111075282 | QUATTRO Intravascular Heat Exchange Catheter Kit Custom Luer Applause Heparin | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189238 | Z-0042-2022 | ZOLL, REF 8700-0793-01 (SL-2593) Premium Access Kit SOLEX 7, Intravascular Heat Exchange catheter Kit, with central venous infusion capabilities (3 lumens), Use Jugular or Subclavian Vein Approach Only, 20 cm, Sterile EO, Rx Only, UDI: | SOLEX 7, Intravascular Heat Exchange catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189411 | Z-0054-2022 | Zoll, REF 8700-0781-03 (CL-2295AE), COOLINE Intravascular Heat Intravascular Heat Exchange Catheter Kit, with central venous infusion capabilities (3 lumens), Sterile EO, Rx Only, UDI: (01)00849111050173 | QUATTRO Intravascular Heat Exchange Catheter Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189440 | Z-0059-2022 | Zoll, REF 8700-000865-01, Cool Line IVTM Disposable Pack and Start UP Kit (6 ft.) | Quattro IVTM Premium Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189445 | Z-0064-2022 | Zoll, REF 8700-000870-40, ICY (Heparin) IVTM Disposable Pack and Start UP Kit (6 ft.), EU, New Luer | ICY (Heparin) IVTM Disposable Pack and Start Up Kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189240 | Z-0044-2022 | ZOLL, REF 8700-0781-14 (CL-2295CO), COOL LINE Intravascular Heat Exchange Catheter kit with central venous infusion capabilities (3 lumens), HEPARIN COATED, STERILE EO, Rx Only, UDI: (01)00849111075510 | COOL LINE Intravascular Heat Exchange Catheter kit | 2 | CDRH | 2021-10-06 |  |  | ZOLL Circulation, Inc. | Labeling update concerning potential leaks from the catheter or the start-up kit (SUK) tubing
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188991 | Z-0032-2022 | Access hsTnI High Sensitivity Troponin I, REF B52699 Access Immunoassay System REAGENT PACK | Access hsTnI (High Sensitivity Troponin I) | 2 | CDRH | 2021-10-05 |  |  | Beckman Coulter Inc. | Through customer feedback, testing a sample with cardiac troponin I concentration that are >270,000 pg/mL may cause carryover into the hsTnl reagent pack. The carryover is caused by sample-to-pack carryover predominantly into the particle well of the reagent pack by way of the probe. The particle well then becomes contaminated by the high sample from the probe. The extent of the carryover is dependent on sample concentration and frequency of high concentration samples.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189249 | Z-0031-2022 | Nox T3s Recorder - Product Usage: intended for ambulatory recording of physiological signals during sleep. | Nox T3S Recorder | 2 | CDRH | 2021-10-05 |  |  | NOX MEDICAL | Error in the device firmware results in the inaccurate detection of device position.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189109 | Z-0010-2022 | Ultra-Myossage Lotion 1-gallon bottle, REF: 4262; Myossage Lotion 1-gallon bottle, REF: 4210, located in the following DJO, LLC / DJO FRANCE kits: | Ultra-Myossage | 1 | CDRH | 2021-10-05 |  |  | DJO, LLC | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188910 | Z-0030-2022 | Nutrifill Scleral, Hybrid, and Gas Permeable (GP) Lens Insertion Solution, Physiologic Formula, Preservative Free, 35 Single-use 10ml ampoules (360ml), UDI:00850012123002 | Nutrifill Scleral, Hybrid, and Gas Permeable (GP) Lens Insertion Solution | 2 | CDRH | 2021-10-05 |  |  | Contamac Solutions, Inc. | Contact lens solution may contain foreign material
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189108 | Z-0009-2022 | GEL CONDUCTOR, Description/REF: 5 LITER W/DISP/4238, 24-8.5OZ BOTTLES/4248, 5 LITER/4266, distributed by DJO, LLC / DJO FRANCE in the following kits, Part/Description: 2073/ INTELECT RPW 120V, 2073-BR/ INTELECT RPW 120V BRAZIL, 2074/ INTELECT RPW 230 V, 2082/ INTELECT RPW 100 V, 2760/ INTELECT LEGEND XT COMBOUS STD, 2791/ INTELECT LEGEND XT COMBO US STD W/CART, 2782-10/ INTELECT TRANSPORT US STD, 2782-2/ INTELECT TRANSPORT US STD, 2843/ INTELECT VET 4CH COMBO PKG, 2845K/ KIT INTELECT VET 4CH COMBO PKG, 2863/ SENIOR SOLUTIONS 2CH COMBO PKG US STD 5CM, 2863-10/ SENIOR SOLUTIONS 2CH COMBO PKG US STD 10CM APPL, 2863-2/ SENIOR SOLUTIONS 2CH COMBO PKG US STD 2CM APPL, 2864/ SENIOR SOLUTIONS 4CH COMBO PKG US STD, 2761/ VECTRA GENISYS 2CH COMBO PKG US STD 5CM APPL, 2792/ VECTRA GENISYS 2CH COMBO US STD W/CART, 2761-10/ VECTRA GENISYS 2CH COMBO PKG US STD 10CM APPL, 2761-2/ VECTRA GENISYS 2CH COMBO PKG US STD 2CM APPL, 2762CC/ INT ADV 2CH COLOR COMBO INT'L 5CM APPL, 2762CC-10/ INT ADV 2CH COLOR COMBO W/O EM INT'L STD 10CM APPL, 2762CC-2/ INT ADV 2CH COLOR COMBO W/O EM INT'L STD 2CM APPL, 2788/ LEGEND XT 4CH COMBO PKG US STD 5CM APPL, 2795/ INTELECT XT 4CH COMBO US STD 5CM APPL W/CART, 2788-10/ LEGEND XT 4CH COMBO PKG US STD 10CM APPL, 2795-10/ INTELECT XT4CH COMBO US STD 10CM APPL, 2788-2/ LEGEND XT 4CH COMBO PKG US STD 2CM APPL, 2795-2/ INTELECT XT 4CH COMBO US STD 2CM APPL, 2789/ VECTRA GENISYS 4CH COMBO PKG US 5CM APPL, 2796/ VECTRA GENISYS 4CH COMBO US STD W/CART, 2796K/ KIT VECTRA GENISYS 4CH COMBO P, 2789-10/ VECTRA GENISYS 4CH COMBO PKG US STD 10CM APPL, 2796-10/ VECTRA GENISYS 4CH COMBO US US STD 10CM APPL, 2789-2/ VECTRA GENISYS 4CH COMBO PKG US STD 2CM APPL, 2796-2/ VECTRA GENISYS 4CH COMBO US STD 2CM APPL, 2738/ INTELECT TRANSPORT COMBO PKG US STD, 8250KP/ KIT TRITON DTS PROMO, 6001/ INTELECT NEO, 2738-10/ INTELECT TRANSPORT COMBO PKG US STD, 2738-2/ INTELECT TRANSPORT COMBO PKG US STD, 2772MC/ INT ADVANCED COMBO MONOCHROMEINT'L STD, 2776/ INTELECT MOBILE INT'L STD 5CM APPL, 27800/ ULTRASOUND ON CART, 2776-BR/ INTELECT MOBILE INTL STD 5CM APPL BRAZIL, 2776-10/ INT MOBIL INT'L STD 10CM APPL, 2776-2/ INT MOBILE INT'L STD 2CM APPL, 2778/ INTELECT MOBILE COMBO INT'L STD 5CM APPL, 2778-10/ INTELECT MOBILE COMBO INT'L STD 10CM APPL, 2778-2/ INTELECT MOBILE COMBO INT'L STD 2CM APPL, 2752CC/ INT ADV 2CH COLOR COMBO INT'L STD, 2756/ INTELECT VET 2CH COMBO PKG INT'L STD, 2844K/ KIT INTELECT VET 2CH COMBO PKG, 2782/ INTELECT TRANSPORT US STD, 2895K/ KIT CART ADAPT BL W/CT/SPEC US, 2872/ VECTRA GEN 2C COMBO W/O EMG PK US STD, 2876/ VECTR GEN 2C COMBO W/O EMG US STD W/CART, 2874/ VECTRA GEN 4C COMBO W/O EMG PK US STD, 2878/ VECTR GEN 4C COMBO W/O EM US STD W/CART, 7550/ INT LEGEND COMBO 4C US STD, INT001/ INTELECT LEGEND ULTRASOUND US STD, INTC B/ INT CB PKGD US STD, 39-0701/ DIRECT SUPPLY PANACEA VECTRA 2-CHNL ULTRASOUND/ESTIM COMBO, 39-0702/ DIRECT SUPPLY PANACEA VECTRA 4-CHNL ULTRASOUND/ESTIM COMBO, 13-4248/ GEL CONDUCTOR 24-8.5OZ BOTTLES, 15-1200/ INTELECT MOBILE 2 ULTRASOUND INTL SET EU PLUG, 15-1201/ INTELECT MOBILE 2 ULTRASOUND INTL SET ALL PLUGS, 15-1204/ INTELECT MOBILE 2 COMBO INTL SET EU PLUG, 15-1205/ INTELECT MOBILE 2 COMBO INTL SET ALL PLUGS, 2173KIT/ INTELECT RPW 2 120V INTERNATIONAL KIT US PLUG, 2174KIT/ INTELECT RPW 2 230V KIT EU PLUG, 2175KIT/ INTELECT RPW 2 100V KIT JPN PLUG, 2176KIT/ INTELECT RPW 2 120V USA KIT, 12-2756-1/ INTELECT VET 2CH COMBO PKG INTL STD, 2756K/ KIT INTELECT VET 2CH COMBO PKGUS STD, 70002/ NEO MODULE ULTRASOUND | Gel Conductor | 1 | CDRH | 2021-10-05 |  |  | DJO, LLC | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189260 | Z-0021-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK Logic All-polyethylene Tibial Components, also labeled as the following: a. OPTETRAK Logic All-polyethylene CR Tibial Components; b. OPTETRAK Logic All-polyethylene CRC Tibial Components; c. OPTETRAK Logic All-polyethylene PSC Tibial Components; | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189020 | Z-0014-2022 | Scan Abutment NP, Cat. No. 30-70102 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment NP | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189019 | Z-0013-2022 | Scan Abutment SP, Cat. No. 30-70101 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment SP | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189268 | Z-0027-2022 | Vantage Total Ankle System including Vantage Mobile-Bearing Polyethylene Liner Component, not sold in the USA. | Vantage Total Ankle System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189024 | Z-0018-2022 | Scan Abutment WP, Cat. No. 30-70106 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment WP | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189015 | Z-0019-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK All-polyethylene Tibial Components, labeled as the following: a. OPTETRAK All-polyethylene CC Tibial Components; b. OPTETRAK All-polyethylene CR Tibial Components; c. OPTETRAK All-polyethylene CR Tibial Sloped Components; d. OPTETRAK All-polyethylene PS Tibial Components; 510(k)s: K011976, K082022, K932690, K932776, K933494, K933610, K954208 | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189006 | Z-0028-2022 | TECNIS Eyhance IOLs are permanent intraocular implants intended for implantation in the capsular bag. The TECNIS Simplicity Delivery System is designed as a single-use disposable preloaded modular cartridge that functions as both the primary packaging and as a sterile, disposable insertion system for delivering the IOL into the eye during cataract surgery. | TECNIS EyhanceTM Toric II IOLs with TECNIS Simplicity" Delivery System/Model Number DIU225U050) | 2 | CDRH | 2021-10-04 |  |  | Johnson & Johnson Surgical Vision Inc | Nonconforming product was distributed in error. Product was nonconforming due endotoxin levels higher than both the Firm's internal endotoxin specification limit and the FDA recommended endotoxin specification limit. Product was released in error based on the results of additional testing performed during the investigation into release testing nonconformance.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189267 | Z-0026-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK Logic RBK Tibial Components, not sold in the USA. Also labeled as OPTETRAK Logic RBK PS Tibial Components | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189264 | Z-0024-2022 | Vantage Total Ankle System including Vantage Fixed-Bearing Polyethylene Liner Component | Vantage Total Ankle System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189263 | Z-0023-2022 | TRULIANT Knee System including TRULIANT Tibial Inserts, also labeled as the following: a. TRULIANT CR Tibial Inserts; b. TRULIANT CRC Tibial Inserts; c. TRULIANT PS Tibial Inserts; d. TRULIANT PSC Tibial Inserts | TRULIANT Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189266 | Z-0025-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK All-polyethylene Tibial Components, not sold in the USA. Also labeled as OPTETRAK All-polyethylene CR Tibial Components | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189259 | Z-0020-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK HI-FLEX Polyethylene Tibial Components, also labeled as OPTETRAK HI-FLEX PS Polyethylene Tibial Components | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189021 | Z-0015-2022 | Scan Abutment MU, Cat. No. 30-70103 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment MU | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189023 | Z-0017-2022 | Scan Abutment PCA, Cat. No. 30-70105 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment PCA | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189261 | Z-0022-2022 | OPTETRAK Comprehensive Knee System including OPTETRAK Logic Modular Tibial Components, also labeled as OPTETRAK Logic Modular PS Tibial Components | OPTETRAK Comprehensive Knee System | 2 | CDRH | 2021-10-04 |  |  | Exactech, Inc. | Inserts were packaged in vacuum bags that lacked an additional oxygen barrier layer.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189022 | Z-0016-2022 | Scan Abutment SU, Cat. No. 30-70104 - Product Usage: used to take an impression with an intra oral scanner or optical scanner to provide a digital registration of an implant or abutment platform. The Paltop Advanced Dental Solutions Implant System is indicated for use in surgical and restorative applications for placement in the bone of the upper or lower jaw to provide support for prosthetic devices, such as artificial teeth, in order to restore the patient s chewing function. | Scan Abutment SU | 2 | CDRH | 2021-10-04 |  |  | Paltop Advanced Dental Solutions Ltd | Some devices are outside tolerances and showed a maximal deviation of 2 degrees in the orientation between the peek head and titanium base, and an incongruence between the size engraved on the peek has been identified. Deviations may delay the final restoration.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189064 | Z-0012-2022 | Centricity Universal Viewer 7.0, Model 5829628-0XX. Radiological image processing system | Centricity Universal Viewer | 2 | CDRH | 2021-10-01 | 2021-10-02 21:06 |  | GE Healthcare, LLC | A software anomaly exists in the Centricity Universal Viewer study management feature in which study changes are not propagated to either the Centricity Enterprise Archive (EA) or another Vendor Neutral Archive (VNA).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189291 | Z-0003-2022 | Medtronic Bio-Medicus Insertion Kits, Model number 96553, Pediatric Insertion Kit, Sterile; and Model number 96551, Venous Insertion Kit, Sterile. | Bio-Medicus" Insertion Kits | 2 | CDRH | 2021-10-01 |  |  | Medtronic Perfusion Systems | There was a label mix-up between two model numbers resulting in the incorrect guidewire being contained in the kits for both model numbers.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188854 | Z-0011-2022 | ELI, Burdick and McKesson brand 280 Resting Electrocardiographs (ELI 280, MLBUR280 and BUR280), Model Numbers: ELI280-DDB-ADAAX, ELI280-BDB-ACAAX, ELI280-CAA-AAFAD, ELI280-BDB-BDFAX, ELI280-BCB-AAAAX, ELI280-DDB-ACAAX, ELI280-DBA-BAFAX, ELI280-DCB-AAAAX, ELI280-DBA-AAFBD, ELI280-AAA-AAEBX, ELI280-BBA-ADFBD, ELI280-DDB-ACFAX, MLBUR280-81X, ELI280-CAA-ACEBX, ELI280-AAA-AAHBX, ELI280-DDB-BDFAX, ELI280-CAA-AAFBT, ELI280-DCD-ADFAD, ELI280-DDB-AACBX, ELI280-JXX-BDFAX, ELI280-CAA-ADCBX, ELI280-BCB-AACBD, ELI280-LDX-ADFBX, ELI280-LDX-ADCBX, ELI280-BDD-ADFAD, ELI280-BDB-AAFBD, ELI280-BDB-ACCAD, ELI280-BDB-ADFAD, ELI280-LDX-ADFBD, ELI280-DDB-AAFBD, ELI280-DCB-ACFAD, ELI280-AAA-BAFAF, MLBUR280-W1X, ELI280-DCB-AAABX, ELI280-AAA-ACAAX, ELI280-ADA-ABFBX, ELI280-BCB-AAFBD, ELI280-BCB-AAFBX, ELI280-CEA-ADFBX, ELI280-CDA-ADABX, ELI280-BDB-AAABX, ELI280-DDB-ACFBD, ELI280-DDB-BCFAX, ELI280-ADA-ADCAX, ELI280-BDB-AACBX, ELI280-BCB-ACAAX, BUR280-81X, ELI280-DCB-AAFBG, ELI280-BDB-AAFBT, ELI280-DCB-AAFAD, ELI280-CAA-ADFBX, ELI280-BDB-AAFBG, ELI280-BDB-AAFBX, ELI280-DCA-ACAAX, ELI280-CAB-ACFBX, ELI280-ADA-ACFAX, ELI280-CAA-AAFAT, ELI280-CAA-ACFBD, ELI280-CDA-ADCBX, ELI280-AAA-ADCBX, ELI280-CAA-AAFBD, ELI280-DCB-BAFBT, ELI280-BBA-AAAAX, ELI280-BBA-AAFAD, ELI280-DBA-AAFAD, ELI280-DCB-AAFBT, ELI280-DBA-ADFAX, ELI280-BBA-AAFBD, MLBUR280-C1X, ELI280-BBA-ADFAX, ELI280-DDD-ADFAD, ELI280-DBA-AAFAX, MLBUR280-W1D, ELI280-BCB-AACBX, ELI280-DCB-AACBD, ELI280-LDX-ADFBG, ELI280-DCB-AACBX, ELI280-DBA-ADFBD, ELI280-DDB-AAFBG, ELI280-DBA-AAABD, ELI280-DCB-AAFBX, ELI280-AAA-AAFBT, MLBUR280-C1D, ELI280-AAB-ADAAX, ELI280-DDB-AAFBT, ELI280-CAA-AACBX, ELI280-CEB-ACFBX, ELI280-CAA-ABFAX, ELI280-DDB-AAFBX, ELI280-DBA-AAAAX, ELI280-BCB-AAFBG, ELI280-BFA-ADCBX, ELI280-LDX-ADABX, MLBUR280-81D, ELI280-BBA-ADFAD, ELI280-CAA-ADHAX, ELI280-CEB-ACFBD, ELI280-AAB-ADCAD, ELI280-BDB-AACBD, ELI280-LDX-ADCBD, ELI280-AAB-ACCBX, ELI280-AAA-ABFBX, ELI280-DAB-ADCAD, ELI280-AFB-ABCBX, ELI280-DCB-AACAX, ELI280-DEB-ACFBD, ELI280-ADA-ACAAX, ELI280-AAB-ADFAD, ELI280-DCB-AAFBD, ELI280-CAA-AAAAX, ELI280-BDB-ACCAX, ELI280-CAA-AAFBX, ELI280-DDB-AAABX, ELI280-DCB-ACAAX, ELI280-CAA-ADFBD, ELI280-CBB-ACCBX, ELI280-DDB-AACBD, ELI280-DDB-ACCAX, ELI280-A, ELI280-DFA-ADCBX, ELI280-BCA-AAAAX, ELI280-DDB-AAAAX, ELI280-E, ELI280-C, ELI280-AAA-AAFBD, ELI280-DFC-ADFAD, BUR280-C1X, ELI280-D, ELI280-BAA-ACCBD, ELI280-AAA-AAAAX, BUR280-W1X, ELI280-B, ELI280-BCB-AAABX, ELI280-BCB-AACAX, BUR280-W1D, ELI280-F, ELI280-BCB-AAFAD, ELI280-BCB-BAFAX, BUR280-81D | ELI, Burdick and McKesson brand 280 Resting Electrocardiographs | 2 | CDRH | 2021-10-01 |  |  | Welch Allyn Inc Mortara | The devices malfunction under specific operator workflows.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188744 | Z-0002-2022 | REF TIN3015, 15G x 79 mm Bone Marrow Aspiration/Intraosseous Infusion Needle, For Single Use Only, Sterile EO, UDI: 10885403043796 | Illinois Bone Marrow Aspiration/Intraosseous Infusion Needle | 2 | CDRH | 2021-10-01 |  |  | Bard Peripheral Vascular Inc | Potential for needle protector to become fully dislodge exposing infusion needle which could result in needle stick.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188990 | Z-2587-2021 | GENOSYL DS (Delivery System) console, for use as a vasodilator. | GENOSYL DS | 1 | CDRH | 2021-09-30 08:17 |  |  | Vero Biotech, LLC | Manufacturing defect involving reversal of the inlet and outlet gas lines, resulting in delivery of nitrogen dioxide (NO2) instead of nitric oxide (NO).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188580 | Z-2595-2021 | Flow sensors kept as stand-alone user replaceable spare parts for the GEHC anethesia machines: 2089610-001 FLOW SENSOR, LEGACY VAR ORF BCG (blue, cleanable) 2089610-001-S FLOW SENSOR, LEGACY VAR ORF BCG, SERVICE (blue, cleanable) 2087640-001 FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG (gray, autoclavable) 2087640-001-S FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG, SERVICE (gray, autoclavable) 2096513-001-S FLOW SENSOR ASSEMBLY 5697309 R-FMI34109-FLOW SENSOR, LEGACY VAR ORF BCG 5697310 R-FMI34109-FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG | Aestiva MRI | 2 | CDRH | 2021-09-30 |  |  | GE Healthcare, LLC | The flow sensors built prior to June 2021 could have damaged tubes with small punctures or cuts that could cause leaks resulting in incorrect anesthesia machine tidal volumes potentially leading to over-delivery of volume to the patient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189060 | Z-2612-2021 | Corin Unity Knee Patella, Size 3 / 8.5mm, Centered Dome 3 PEG-UHMWPE, REF 112.018.06, Sterile, a component of the Unity Total Knee System. | Unity Total Knee System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188982 | Z-2598-2021 | FlexoStretch Model Number: 3156057 - The combination of this products allows the user to easily adjust/control the center of gravity during lift of patient. | FlexoStretch | 2 | CDRH | 2021-09-30 |  |  | Hill-Rom, Inc. | Stretch Leveller has the potential to lose the function of maintaining a patient in the horizontal position when lifting and in motion. When this function is lost the patient can slide backwards and potentially may fall out of the Oct lift sheet back onto the surface it is being lifted from.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189047 | Z-2607-2021 | Corin BIOLOX Delta Mod Head, diameter 28mm Medium 0mm, 12/14 Taper, REF 104.2805, Sterile, a component of the Trinity Acetabular System. | Trinity Acetabular System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189052 | Z-2609-2021 | Corin MetaFix Hip Stem, Size 1, Short Neck, 135 degrees, Collared Hip Stem, REF 579.0201, Sterile. | Corin MetaFix Hip Stem | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188989 | Z-2614-2021 | DuoDERM Control Gel Formula Border Dressings 15x15cm (1x5pk) | DuoDERM | 2 | CDRH | 2021-09-30 |  |  | ConvaTec, Inc | Primary packaging labeling may be incorrect.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188983 | Z-2599-2021 | Stretch Leveller Model Number: 3156200 - The combination of this products allows the user to easily adjust/control the center of gravity during lift of patient. | Stretch Leveller | 2 | CDRH | 2021-09-30 |  |  | Hill-Rom, Inc. | Stretch Leveller has the potential to lose the function of maintaining a patient in the horizontal position when lifting and in motion. When this function is lost the patient can slide backwards and potentially may fall out of the Oct lift sheet back onto the surface it is being lifted from.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189204 | Z-2616-2021 | Cardiosave Rescue Intra-Aortic Balloon Pump (IABP), Model Nos. 0998-00-0800-75, 0998-00-0800-83, 0998-00-0800-85 | Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) | 2 | CDRH | 2021-09-30 |  |  | Datascope Corp. | The Cardiosave IABP may unexpectedly shut down when the device is running on AC power, only one battery is installed in the IABP, and the battery is physically removed while the battery is being charged. This occurs during a very specific set of conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188984 | Z-2600-2021 | LikoStretch 1900 Model Number: 3156051 - The combination of this products allows the user to easily adjust/control the center of gravity during lift of patient. | Hill Rom | 2 | CDRH | 2021-09-30 |  |  | Hill-Rom, Inc. | Stretch Leveller has the potential to lose the function of maintaining a patient in the horizontal position when lifting and in motion. When this function is lost the patient can slide backwards and potentially may fall out of the Oct lift sheet back onto the surface it is being lifted from.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189203 | Z-2615-2021 | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP), Model Nos. 0998-00-0800-31, 0998-00-0800-32, 0998-00-0800-33, 0998-00-0800-34, 0998-00-0800-35, 0998-00-0800-45, 0998-00-0800-52, 0998-00-0800-53, 0998-00-0800-55, 0998-00-0800-65 | Cardiosave Hybrid Intra-Aortic Balloon Pump (IABP) | 2 | CDRH | 2021-09-30 |  |  | Datascope Corp. | The Cardiosave IABP may unexpectedly shut down when the device is running on AC power, only one battery is installed in the IABP, and the battery is physically removed while the battery is being charged. This occurs during a very specific set of conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188892 | Z-2602-2021 | DxA 5000 (DxA Automation System, DxA 5000 fit), "For in Vitro Diagnostic Use", Reference Numbers: B87352 with B57634, UDI: 15099590707118 | DxA 5000, Automation System | 2 | CDRH | 2021-09-30 |  |  | Beckman Coulter Inc. | There is a potential that sample carriers (w or w/o samples) within automated sampling system may transport samples to the outer lane which does not contain a physical barrier to restrain samples. Samples are dropped when they reach the end of the outer lane which could result in Delay in providing patient with appropriate medical treatment. New sample needs to be drawn. Substantial discharge of biohazards material - Infection with a human transmissible blood-borne or animal pathogen which could result in Infection with a human transmissible blood-borne pathogen (e.g. infectious HBV, HCV, HIV).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189065 | Z-2617-2021 | Xcela PICC with PASV Valve Technology/Kit 5 F-55cm; (Maximum Flow Rate 3.5 mL/Sec)-indicated for short or long-term peripheral access to the central venous system for intravenous therapy, including but not limited to, the administration of fluids, medications, and nutrients; the sampling of blood; and for power injection of contrast media. Catalog Number: 60M701995 | Navilyst | 2 | CDRH | 2021-09-30 |  |  | Angiodynamics Inc. (Navilyst Medical Inc.) | Mislabeled: The product is labeled as 5F (55cm; Maximum Flow Rate 3.5 mL/Sec) and the product is actually 4F (55cm; Maximum Flow Rate 3.5 mL/Sec).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189041 | Z-2603-2021 | Corin TriFit TS, Size 8, Standard Tapered Stem, REF 694.0008, Sterile. | Corin TriFit TS | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189055 | Z-2610-2021 | Corin Unity Knee Augments, Size 5, 6 & 7 / 10mm Thick, Femoral Posterior Augment, REF 1122370, Sterile, a component of the Unity Total Knee System. | Unity Total Knee System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189058 | Z-2611-2021 | Corin Unity Knee CS Insert, Size 5 / 9.0mm, Fixed Condylar Stabilised - UHMWPE, REF 1125182, Sterile, and Corin Unity Knee CS Insert, Size 8 / 10.0mm, Fixed Condylar Stabilised - UHMWPE, REF 1125243, Sterile, both are components of the Unity Total Knee System. | Unity Total Knee System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189049 | Z-2608-2021 | Corin Trinity Dual Mobility, Size 4, Dual Mobility CoCr Liner, REF 321.04.540, Sterile, and Corin Dual Mobility, Size 4 42 insert/28 head, REF 325.04.042, Sterile; both are components of the Trinity Dual Mobility System. | Trinity Dual Mobility System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189043 | Z-2605-2021 | Corin Trinity-i, diameter 50mm Taper Size 3, Non-Occluded Ti Plasma Shell, REF 320.03.350, Sterile, a component of the Trinity Acetabular System. | Trinity Acetabular System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189042 | Z-2604-2021 | Corin Trinity Screw, diameter 6.5mm 35mm Length, Cancellous Bone Screw, REF 321.035, Sterile; and Corin Trinity Screw, diameter 6.5mm 50mm Length, Cancellous Bone Screw, REF 321.050, Sterile, both are components of the Trinity Acetabular System. | Trinity Acetabular System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189067 | Z-2601-2021 | LikoStretch Mod 600 IC Model Number: 3156065 LikoStretch Mod 600 IC, Wide Model Number: 3156065B (obsolete since 12JUN2020) - The combination of this products allows the user to easily adjust/control the center of gravity during lift of patient. | LikoStretch Mod 600 IC | 2 | CDRH | 2021-09-30 |  |  | Hill-Rom, Inc. | Stretch Leveller has the potential to lose the function of maintaining a patient in the horizontal position when lifting and in motion. When this function is lost the patient can slide backwards and potentially may fall out of the Oct lift sheet back onto the surface it is being lifted from.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188981 | Z-2597-2021 | OctoStretch with Stretch Leveller - The combination of this products allows the user to easily adjust/control the center of gravity during lift of patient. Model Number: 3156056 | OctoStretch with Stretch Leveller | 2 | CDRH | 2021-09-30 |  |  | Hill-Rom, Inc. | Stretch Leveller has the potential to lose the function of maintaining a patient in the horizontal position when lifting and in motion. When this function is lost the patient can slide backwards and potentially may fall out of the Oct lift sheet back onto the surface it is being lifted from.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189045 | Z-2606-2021 | Corin Trinity Liner ECIMA, diameter 40mm Neutral Offset, Taper Size 4, REF 322.04.640, Sterile; and Corin Trinity Liner ECIMA, diameter 36mm Neutral 4mm EPW, Taper Size 4 (Elevated Wall), REF 322.04.936, Sterile; both are components of the Trinity Acetabular System. | Trinity Acetabular System | 2 | CDRH | 2021-09-30 |  |  | Corin Ltd | The packaging system on the devices have potential physical and water damage resulting in damage to the devices, loss of sterility, or contamination of the devices.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188592 | Z-2596-2021 | Flow sensors that may be installed in the following GEHC anesthesia machines: (1) Aestiva MRI, GTIN 0080682102339; (2) Aestiva 7900; (3) Aestiva 7100; (4) Aespire System Carestation, S/5 Protiva Config., Aespire 100 Compact Wall Mount (Aespire 7100); (5) S/5 Aespire 7900; (6) Aespire View; (7) Avance; (8) Avance CS2, GTIN 00840682102292; (9) Aisys; (10) Aisys CS2, GTIN 00840682102322; (11) Amingo; (12) 9100 NXT Config, Aelite NXT; (13) Protiva 7100; | Aestiva 7900 | 2 | CDRH | 2021-09-30 |  |  | GE Healthcare, LLC | The flow sensors built prior to June 2021 could have damaged tubes with small punctures or cuts that could cause leaks resulting in incorrect anesthesia machine tidal volumes potentially leading to over-delivery of volume to the patient.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189197 | Z-2499-2021 | DeRoyal Procedure Packs containing Smiths Medical NORMOFLO Irrigation Warming Set | DeRoyal | 1 | CDRH | 2021-09-29 15:53 |  |  | Deroyal Industries, Inc. Lafollette | DeRoyal manufactured procedure packs containing a Smiths Medical NORMOFLO Irrigation Warming System which is recalled due to the potential for aluminum ions leaching into warming fluids.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189073 | Z-2590-2021 | AMX Navigate Mobile X-Ray System | AMX Navigate Mobile X-Ray System | 2 | CDRH | 2021-09-29 13:53 | 2021-09-29 21:07 |  | GE Healthcare, LLC | An unexpected column motion while attempting to park the column on the AMX Navigate mobile X-ray system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189088 | Z-2577-2021 | Imperative Care, ZOOM 71 Reperfusion Catheter, REF: ICRC071137, TRX Tip, 137 cm, 6F (.083 in/2.11 mm), .071 in/1.8 mm), STERILE.EO, Rx Only, UDI: (01) 00812212030191. For neurovascular procedures. | ZOOM 71 Reperfusion Catheter | 1 | CDRH | 2021-09-29 13:48 | 2021-09-29 21:07 |  | IMPERATIVE CARE INC | There is a potential for distal end of catheters to fracture and become detached.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189091 | Z-2582-2021 | Norco Ultrasound Gel, Item/Description: NC70479/Norco Ultrasound Gel, 250 mL; NC70479/North Coast Ultrasound Gel, 250 mL; NC70479-5/Norco Ultrasound Gel, 5 liters; NC70479-5/North Coast Ultrasound Gel, 5 liters; NC70479-5C/Norco Ultrasound Gel, 5 liters (4); NC70479-5C/North Coast Ultrasound Gel, 5 liters (4); NC70479C/Norco Ultrasound Gel, 250 mL (24); NC70479C/North Coast Ultrasound Gel, 250 mL (24) | Norco Ultrasound Gel | 1 | CDRH | 2021-09-29 07:42 |  |  | North Coast Medical Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189092 | Z-2583-2021 | EcoLotion Transmission Lotion, used in ultrasound procedures. Item/Description: NC70481/Eco Lotion 5 liters, NC70482/Eco Lotion with Aloe 5 liters, NC70482C/Eco Lotion with Aloe 5 liters (4); Myossage Lotion, 1 Gallon, Item: NC31208 | Ecolotion | 1 | CDRH | 2021-09-29 07:42 |  |  | North Coast Medical Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189079 | Z-2589-2021 | Atrium ADVANTA VXT Vascular Graft, Part No. 22059, UDI 00650862220599 - Product Usage: intended for use in arterial vascular reconstruction, segmental bypass, and for arteriovenous vascular access. | Atrium ADVANTA VXT Vascular Graft | 2 | CDRH | 2021-09-29 |  |  | Atrium Medical Corporation | The product package may potentially contain a mismatched graft that does not contain the radial support ring (helix) identified in the labeling.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189071 | Z-2591-2021 | MOSAIQ Oncology Information System | MOSAIQ Oncology Information System | 2 | CDRH | 2021-09-29 |  |  | Elekta, Inc. | A drug strength in MOSAIQ can be changed during the ordering process and approved, but the saved drug strength reverts to the original value.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188689 | Z-2592-2021 | Atrium Pneumostat Chest Drain Valve, Part Number 16100 | Atrium Pneumostat Chest Drain Valve | 2 | CDRH | 2021-09-29 |  |  | Atrium Medical Corporation | The firm is revising the Indication for Use, Contraindications, Warnings, and Precautions, after the device was involved in the treatment of a preterm infant who later died.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188896 | Z-2594-2021 | uEXPLORER PET/CT System - Product Usage: intended to be operated by qualified healthcare professionals to assist in the detection, localization, diagnosis, staging, restaging, treatment planning and treatment response evaluation for diseases and disorders in, but not limit to, oncology, cardiology and neurology. | uEXPLORER PET/CT System | 2 | CDRH | 2021-09-29 |  |  | Shanghai United Imaging Healthcare Co., Ltd. | The EXAM acquisition workflow could occasionally freeze after canceling the PET scan.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188895 | Z-2593-2021 | uMI 550 System - Product Usage: intended to be operated by qualified healthcare professionals to assist in the detection, localization, diagnosis, staging, restaging, treatment planning and treatment response evaluation for diseases and disorders in, but not limit to, oncology, cardiology and neurology. | uMI 550 System | 2 | CDRH | 2021-09-29 |  |  | Shanghai United Imaging Healthcare Co., Ltd. | The EXAM acquisition workflow could occasionally freeze after canceling the PET scan.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188736 | Z-2588-2021 | Arjo Citadel beds originally assembled with the IndiGo modules during manufcturing, model number CX812A3J3AMAD0; and IndiGo modules (retrofit kits) that can be installed on Citadel (model no. INDIDAA) or Enterprise (model no. INDIJAA) beds manufactured without IndiGo modules assembled at the time of manufacturing and subsequently retrofitted on the beds. | Arjo | 2 | CDRH | 2021-09-29 |  |  | ARJOHUNTLEIGH POLSKA Sp. z.o.o. | The IndiGo power cord might wear during use which could lead to its damage, resulting in the cord being cut, burnt or short circuit with limited emission of sparks.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189199 | Z-2584-2021 | Operating room table - Product Usage: intended for use during diagnostic examinations, radiologic procedures or surgical procedures to support and position a patient. | ORT200 | 2 | CDRH | 2021-09-28 |  |  | Deerfield Imaging, Inc. | ORT200/300 rotational lock may fail to function such that it cannot change states between locked and unlocked, due to a mechanical hardware issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189201 | Z-2585-2021 | Operating room table - Product Usage: intended for use during diagnostic examinations, radiologic procedures or surgical procedures to support and position a patient. | ORT300 | 2 | CDRH | 2021-09-28 |  |  | Deerfield Imaging, Inc. | ORT200/300 rotational lock may fail to function such that it cannot change states between locked and unlocked, due to a mechanical hardware issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189202 | Z-2586-2021 | Operating room table - Product Usage: intended for use during diagnostic examinations, radiologic procedures or surgical procedures to support and position a patient. | ORT300 | 2 | CDRH | 2021-09-28 |  |  | Deerfield Imaging, Inc. | ORT200/300 rotational lock may fail to function such that it cannot change states between locked and unlocked, due to a mechanical hardware issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189077 | Z-2581-2021 | AtomLab 500Plus Dose CalibratorSoftware 2.0.00 through Revision 2.0.08, inclusive Model: 086-332 | Biodex | 2 | CDRH | 2021-09-27 14:55 |  |  | Biodex Medical Systems, Inc. | When deleting a previously entered custom isotope, the software deletes the isotope, but not the associated dial setting. Possible other stored custom isotopes may then have incorrect dial settings and may lead to an incorrect dose measurement and/or misadministration
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189076 | Z-2580-2021 | Atomlab 500 Dose Calibrator Software 2.0.00 through Revision 2.0.08, inclusive Model: 086-330. | Biodex | 2 | CDRH | 2021-09-27 14:55 |  |  | Biodex Medical Systems, Inc. | When deleting a previously entered custom isotope, the software deletes the isotope, but not the associated dial setting. Possible other stored custom isotopes may then have incorrect dial settings and may lead to an incorrect dose measurement and/or misadministration
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189008 | Z-2576-2021 | Vysis CLL FISH Probe Kit with the following components: Vysis LSI p53 SpectrumOrange/ATM, SpectrumGreen and LSI D13S319, SpectrumOrange/ 13q34, SpectrumAqua/CEP 12 SpectrumGreen Probes | Vysis CLL FISH Probe Kit | 2 | CDRH | 2021-09-27 11:26 |  |  | Abbott Molecular, Inc. | Potential for Vysis CLL FISH Probe Kits not detecting 13q deletions in known positive patient samples.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188398 | Z-2578-2021 | PAS-PORT PROXIMAL ANASTOMOSIS SYSTEM Catalog Number: FC700SU | Aesculap | 2 | CDRH | 2021-09-27 |  |  | Aesculap Implant Systems LLC | Malfuncton-Implant did not deploy successfully may necessitate the need for an additional medical intervention for patients which would result in a potential delay in surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188399 | Z-2579-2021 | PAS-PORT PROXIMAL ANASTOMOSIS DEVICE US Catalog Number: FG-000001-13 | Aesculap | 2 | CDRH | 2021-09-27 |  |  | Aesculap Implant Systems LLC | Malfuncton-Implant did not deploy successfully may necessitate the need for an additional medical intervention for patients which would result in a potential delay in surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189074 | Z-2518-2021 | Omnisound Gel, 250 ml bottle and 5 L container; and EcoGel 200, 250 ml bottle; Manufactured by Eco-Med. For use in ultrasound procedures. | Omnisound gel and EcoGel | 1 | CDRH | 2021-09-24 21:52 |  |  | Accelerated Care Plus Corporation | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188996 | Z-2565-2021 | BF-3C40 OES Bronchofiberscope, Model No. BF-3C40 | BF-3C40 OES Bronchofiberscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189000 | Z-2569-2021 | BF-XP160F EVIS EXERA Bronchovideoscope, Model No. BF-XP160F | BF-XP160F EVIS EXERA Bronchovideoscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189001 | Z-2570-2021 | BF-XT160 EVIS EXERA Bronchovideoscope, Model No. BF-XT160 | BF-XT160 EVIS EXERA Bronchovideoscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188999 | Z-2568-2021 | BF-P60 OES Bronchofiberscope, Model No. BF-P60 | BF-P60 OES Bronchofiberscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188997 | Z-2566-2021 | BF-XP60 OES Bronchofiberscope, Model No. BF-XP60 | BF-XP60 OES Bronchofiberscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188998 | Z-2567-2021 | BF-1T60 OES Bronchofiberscope, Model No. BF-1T60 | BF-1T60 OES Bronchofiberscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189002 | Z-2571-2021 | BF-P180 EVIS EXERA II Bronchovideoscope, Model No. BF-P180 | BF-P180 EVIS EXERA II Bronchovideoscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188995 | Z-2564-2021 | BF-N20 OES Bronchofiberscope, Model No. BF-N20 | BF-N20 OES Bronchofiberscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189004 | Z-2573-2021 | BF-Q180-AC EVIS EXERA II Bronchovideoscope, Model No. BF-Q180-AC | BF-Q180-AC EVIS EXERA II Bronchovideoscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189003 | Z-2572-2021 | BF-1T180 EVIS EXERA II Bronchovideoscope, Model No. BF-1T180 | BF-1T180 EVIS EXERA II Bronchovideoscope | 2 | CDRH | 2021-09-24 13:18 |  |  | Olympus Corporation of the Americas | The firm is issuing validated, new reprocessing instructions according to the FDA 2015 guidance, "Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling."
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188273 | Z-2543-2021 | The Instructions for Use (IFU) for: Carotid Patch and Gelsoft Patch vascular patches. FSN says Tapered Carotid Patch 920875t | VASCUTEK/TERUMO | 2 | CDRH | 2021-09-24 11:46 |  |  | Vascutek, Ltd. | The devices contained the IFU for distribution to ROW (rest of world) consignees rather than U.S. consignees.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188272 | Z-2542-2021 | The Instructions for Use (IFU) for: Gelsoft Plus Bifurcate, Gelsoft Plus Ax-Fem, GelsoftPlus ERS Ax-Fem, Gelweave Straight, Gelweave Trifurcate, Gelweave Ant-Flo, and Gelweave Valsalva vascular grafts: (1) Vascutek Terumo Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis, Bifurcate, 14 x 8mm, Bifurcate 45cm, REF 631408, sterile. (2) Vascutek Terumo Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis, Bifurcate, 16 x 8mm, Bifurcate 45cm, REF 631608, sterile. (3) Vascutek Terumo Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis, Bifurcate, 18 x 9mm, Bifurcate 45 cm, REF 631809, sterile. (4) Vascutek Terumo Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis, Bifurcate, 20 x 10mm, Bifurcate 45cm, REF 632010, sterile. (5) Vascutex Terumo Gelsoft Gelatin Impregnated Knitted Vascular Prosthesis, Bifurcate, 22 x 11mm, Bifurcate 45cm, REF 632211, sterile | VASCUTEK/TERUMO | 2 | CDRH | 2021-09-24 11:46 |  |  | Vascutek, Ltd. | The devices contained the IFU for distribution to ROW (rest of world) consignees rather than U.S. consignees.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188868 | Z-2485-2021 | VACUETTE TUBE 6 ml K2E K2EDTA 13x100 pink cap-pink ring, non-ridged | VACUETTE¿ TUBE 6 ml K2E K2EDTA 13x100 pink cap-pink ring, non-ridged | 2 | CDRH | 2021-09-24 10:17 |  |  | Greiner Bio-One North America, Inc. | Blood collection tubes may experience a clotting issue.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188614 | Z-2561-2021 | Covidien HysteroLux Fluid Management System Control Unit, REF: 72205000 | HysteroLux" Fluid Management System | 2 | CDRH | 2021-09-24 |  |  | WOM World of Medicine AG | When the display of inflow volume to the uterus reaches 32450 ml, the fluid management system control unit display will freeze because the internal software calculation threshold is reached. The outflow measurement will continue, and so the deficit accumulated up to this point will start counting backwards to 0 ml. There is a risk of distention fluid reaching the soft tissue circulatory system.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188926 | Z-2544-2021 | 15 drops/ml, Priming Volume: 27ml, Length: 129 in. Single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluid Material ID: 354212 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188929 | Z-2545-2021 | Outlook Pump Sets w/3 CARESITE Needle-Free Injection Sites: single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Material ID: 352413 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188974 | Z-2575-2021 | NAMIC Stopcocks: (1) NAMIC STOPCOCK, 200 psi (14 bar), REF 70015002, UPN H965700150021, Three-way Stopcock with Rotating Collar, Rx only, Sterile; (2) NAMIC STOPCOCK, 200 psi (14 bar), REF 70015003, UPN H965700150031, Three-Way Stopcock with Fixed Male Luer Lock, Rx only, Sterile; (3) NAMIC STOPCOCK, 200 psi (14 bar), REF 70015012, UPN H965700150121, Four-way Stopcock with Rotating Collar, Rx only, Sterile; (4) NAMIC STOPCOCK, 200 psi (14 bar), REF 70015013, UPN H965700150131, Four-way Stopcock with Fixed Male Luer Lock, Rx only, Sterile; (5) NAMIC STOPCOCK, 200 psi (14 bar), REF 70015015, UPN H965700150151, One-way Stopcock with Collar, Rx only, Sterile; (6) NAMIC STOPCOCK, 400 psi (28 bar), REF 70035002, UPN H965700350021, Three-way Stopcock with Rotating Collar, Rx only, Sterile; (7) NAMIC STOPCOCK, 400 psi (28 bar), REF 70035003, UPN H965700350031, Three-way Stopcock with Fixed Male Luer Lock, Rx only, Sterile; (8) NAMIC STOPCOCK, 400 psi (28 bar), REF 70035008, UPN H965700350081, Three-way Stopcock with Rotating Adaptor, Rx only, Sterile; (9) NAMIC STOPCOCK, 400 psi (28 bar), REF 70035009, UPN H965700350091, Three-way Stopcock with Rotating Adaptor, Rx only, Sterile; (10) NAMIC STOPCOCK, 1050 psi (72 bar), REF 70055003, UPN H965700550031, Three-way Stopcock with Fixed Male Luer Lock, Rx only, Sterile; (11) NAMIC STOPCOCK, 1050 psi (72 bar), REF 70055008, UPN H965700550081, Three-way Stopcock with Rotating Adapter, Rx only, Sterile; (12) NAMIC STOPCOCK, 1050 psi (72 bar), REF 70055009, UPN H965700550091, Three-way Stopcock with Rotating Adaptor, Rx only, Sterile; and (13) NAMIC STOPCOCK, 1050 psi (72 bar), REF 70055017, UPN H965700550171, One-way Stopcock with Fixed Male Luer Lock, Rx only, Sterile. | NAMIC Stopcocks | 2 | CDRH | 2021-09-24 |  |  | Medline Industries Inc | Potential for sterile barrier breach. Microscopic pinholes within the sterile pouch could potentially lead to non-sterile conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188932 | Z-2547-2021 | 60 drops/ml, Priming volume: 25 ml, Length: 120 in.in-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 354215 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188938 | Z-2552-2021 | OUTLOOK PUMP IV SET W/UNIVERSAL SPIKE-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 474004 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188942 | Z-2556-2021 | OUTLOOK PUMP SET, 3 CARESITE LADS, 115 IN-single use, for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 490361 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188935 | Z-2549-2021 | Outlook Pump Measured Volume Solution Set/ w/ Auto Shut-Off. Single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 375118 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188931 | Z-2546-2021 | Outlook¿ Pump 15 Drop 0.2 Micron Filtered IV Set with 2 CARESITE¿ Y-Sites: single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 354214 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188943 | Z-2557-2021 | OUTLOOK PUMP SET 3 CARESITE LADS 133 IN.-single use, for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 490564 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188945 | Z-2559-2021 | Outlook Pump Primary Administration Set w/ micro Tubing for Epidural Administration-single use, for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: V7453 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189270 | Z-2563-2021 | Elekta Monaco - Product Usage: used to make treatment plans for patients with prescriptions for external beam radiation therapy. | Monaco RTP System | 2 | CDRH | 2021-09-24 |  |  | Elekta Inc | Contour changes can be saved on an unintended image set. In addition, these contour edits do not cause the dose to be frozen on the plans associated with that image set.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188940 | Z-2554-2021 | OUTLOOK PUMP IV SET, 134 IN.-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 477009 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188941 | Z-2555-2021 | PUMP SET, 15 DR, 3 CARESITE, 127 in.single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 480239 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188944 | Z-2558-2021 | Outlook Pump Add-On Set-single use, for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: V7400 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188774 | Z-2562-2021 | LYMOL Medical Elite X Class Rigid Bronchoscope System, Part No. BD2410, BB2701, BB2702, BB2401, BB2402, BA2700, BT2000, BT2101, BT2103, BT2105, BT2106, BT2210, BT2201, BT2203, BT2205, BR2200, BP2200, BI2200, BR2000, BP2000, BI2999, BR2103, BP2103, BI2103, BR2101, BP2101, BI2101, BI2001, BDL2420 | LYMOL Medical Elite X Class Rigid Bronchoscope System | 2 | CDRH | 2021-09-24 |  |  | Lymol Medical Corporation | Reprocessing instructions were revised November 15, 2020, to reflect new processes and equipment used at consignee user facilities. New instructions were not sent to all consignees of the devices at the time of the revision.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188890 | Z-2574-2021 | NAMIC Manifolds. (1) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 64038301, UPN H965640383011, 3 Valve Manifold, Off Star Handles, Ports on Right, Rx only, Sterile; (2) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 64038303, UPN H965640383031, 3 Valve Manifold, On Star Handles, Ports on Right, Rx only, Sterile; (3) NAMIC Angiographic MANIFOLD High Pressure, REF 70037200, UPN H965700372001, 2 Valve Manifold, Off Handles, Ports on Right, Rx only, Sterile; (4) NAMIC Angiographic MANIFOLD, High Pressure, REF 70037202, UPN H965700372021, 2 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile. (5) NAMIC Angiographic MANIFOLD, High Pressure, REF 70037301, UPN H965700373011, 3 Valve Manifold, Off Handles, Ports on Right, Rx only, Sterile; (6) NAMIC Angiographic MANIFOLD, High Pressure, REF 70037303, UPN H965700373031, 3 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile; (7) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70038200, UPN H965700382001, 2 Valve Manifold, Off Handles, Ports on Right, Rx only, Sterile; (8) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70038202, UPN H965700382021, 2 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile; (9) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70038301, UPN H965700383011, 3 Valve Manifold, Off Handles, Ports on Right, Rx only, Sterile; (10) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70038303, UPN H96570038301, 3 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile; (11) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70039202, UPN H965700392021, 2 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile; (12) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70039301, UPN H965700393011, 3 Valve Manifold, Off Handles, Ports on Right, Rx only, Sterile; and (13) NAMIC Angiographic MANIFOLD, Medium Pressure, REF 70039303, 3 Valve Manifold, On Handles, Ports on Right, Rx only, Sterile. | NAMIC Manifolds | 2 | CDRH | 2021-09-24 |  |  | Medline Industries Inc | Potential for sterile barrier breach. Microscopic pinholes within the sterile pouch could potentially lead to non-sterile conditions.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188934 | Z-2548-2021 | 10 drops/ml, Priming volume: 46ml, Length: 130 in.-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 354216 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188936 | Z-2550-2021 | ET82HP OUTLOOK PUMP EXTENSION SET-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 470024 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188939 | Z-2553-2021 | Y-Type Blood, Outlook Pump Set, Caresite-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 477005 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188937 | Z-2551-2021 | V1500HP OUTLOOK PUMP SET 118-single-use for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: 470025 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188946 | Z-2560-2021 | Outlook Pump Blood Administration Set with SAFELINE Injection Site-single use, for use with the Outlook Safety Infusion System and/or Horizon NXT Pump for infusion of parenteral fluids Ref: NF3120 | Outlook¿ Pump Sets | 2 | CDRH | 2021-09-24 |  |  | B. Braun Medical, Inc. | Leakage of the pump set within the cassette portion of the pump sets, potential for delays in administration of medications, under-delivery or incomplete dosing and bloodstream infections.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189075 | Z-2520-2021 | 10 ga ARS Decompression Needle, Part Number ZZ-0298 | ARS Decompression Needle | 2 | CDRH | 2021-09-23 12:58 |  |  | North American Rescue LLC. | The device includes the incorrect Instructions for Use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188839 | Z-2536-2021 | Catalog A902916, SureStep" Foley Tray, Lubricath", Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188841 | Z-2537-2021 | Catalog A943216, SureStep" Foley Tray, Lubricath", Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188824 | Z-2522-2021 | Catalog A300314A, SureStep" Foley Tray, Bardex" I.C., Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188835 | Z-2533-2021 | Catalog A897514, SureStep" Foley Tray, Silicone, Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188836 | Z-2534-2021 | Catalog A899916, SureStep" Foley Tray, Lubricath", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188848 | Z-2540-2021 | Catalog A303416A, SureStep" Foley Tray, Lubri-Sil" I.C., Temperature Sensing Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188827 | Z-2525-2021 | Catalog A303316A, SureStep" Foley Tray, Bardex" I.C., Complete Care", Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188851 | Z-2541-2021 | Catalog A942216, SureStep" Foley Tray, Bard LubricathTM, Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188719 | Z-2519-2021 | DuoDERM CGF Dressing 10x10CM (1x5PK) (ICC 187660) DuoDERM Extra Thin Dressing 15X15CM (1X10PK) (ICC 187957) DuoDERM Extra Thin Dressing 10X10CM (1X10PK) (ICC 187955) | DuoDERM | 2 | CDRH | 2021-09-23 |  |  | ConvaTec, Inc | There is a potential for open seals which can compromise sterility.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188823 | Z-2521-2021 | Catalog A119216M, SureStep" Foley Tray, Lubri-Sil", Urine Meter, Temperature Sensing Dual, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188837 | Z-2535-2021 | Catalog A902616, SureStep" Foley Tray, Bardex" I.C., Lubricath", Urine Meter | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188844 | Z-2538-2021 | Catalog A319516AM, SureStep" Foley Tray, Lubri-Sil" I.C., Temperature Sensing Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188825 | Z-2523-2021 | Catalog A300316A, SureStep" Foley Tray, Bardex" I.C., Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188828 | Z-2526-2021 | Catalog A303318A, SureStep" Foley Tray, Bardex" I.C., Complete Care", Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188829 | Z-2527-2021 | Catalog A319416AM, SureStep" Foley Tray, Bardex" I.C., Complete Care", Urine Meter, Temperature Sensing, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188830 | Z-2528-2021 | Catalog A399400A, SureStep" Foley Tray, Add-a-Foley, Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188831 | Z-2529-2021 | Catalog A800061, SureStep" Foley Tray, Bard LubricathTM Foley Catheter Tray | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188845 | Z-2539-2021 | Catalog A902416, SureStep" Foley Tray, Silicone, Urine Meter, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188833 | Z-2531-2021 | Catalog A800361, SureStep" Foley Tray, Bard LubricathTM Foley Catheter Tray | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188834 | Z-2532-2021 | Catalog A800365, SureStep" Foley Tray, Bard LubricathTM Foley Catheter Tray | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188826 | Z-2524-2021 | Catalog A300318A, SureStep" Foley Tray, Bardex" I.C., Complete Care", Drainage Bag, StatLock" Stabilization Device | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188832 | Z-2530-2021 | Catalog A800360, SureStep" Foley Tray, Bard LubricathTM Foley Catheter Tray | SureStep Foley Tray System | 2 | CDRH | 2021-09-23 | 2021-09-23 21:07 |  | C.R. Bard Inc | There is potential for packaging defects that may impact the sterile barrier.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189005 | Z-2507-2021 | Greiner Bio-One VACUETTE¿ TUBE 2 ml 9NC Coagulation sodium citrate 3.2% 13x75 blue cap-white ring, sandwich tube, non-ridged | VACUETTE¿ TUBE 2 ml 9NC Coagulation sodium citrate 13x75 mm blue cap-white ring, sandwich tube, non- | 2 | CDRH | 2021-09-22 12:38 |  |  | Greiner Bio-One North America, Inc. | Complaint of tubes clotting due to variation of anticoagulant and/or tubes have low vacuum.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188977 | Z-2505-2021 | DLP Left Heart Vent Catheters, 18 FR. Sterile, Rx Only |  | 2 | CDRH | 2021-09-22 08:04 |  |  | Medtronic Perfusion Systems | Potential for a wire protrusion through the left heart vent catheter tip
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188975 | Z-2504-2021 | DLP Left Heart Vent Catheters, 16 FR. Sterile, Rx Only |  | 2 | CDRH | 2021-09-22 08:04 |  |  | Medtronic Perfusion Systems | Potential for a wire protrusion through the left heart vent catheter tip
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188396 | Z-2516-2021 | Univation X System Tray Sysem knee implant devices as follows: Tray Number/Set Name/Set Number NM1090R Navigation Set ST0560; NM1092R + NM1095R + NM1096R (Tray insert to NM1095R) univation XF ST0540; NM1093R univation XF OPT ST0544 | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188870 | Z-2506-2021 | CoLink Bone Graft Harvester, 8mm, Sterile, Disposable, REF G05 S1008 - Product Usage: is a general instrument intended for assisting surgeons in the manipulation of bone during surgical procedures. | CoLink Bone Graft Harvester | 2 | CDRH | 2021-09-22 |  |  | In2bones USA, LLC | Complaints were received of CoLink Bone Graft Harvesters breaking at the outer tube weld site.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188394 | Z-2514-2021 | Univation X System knee implant devices as follows: Catalog/Product: NO156Z AS UNIVATION XF TIBIA CEMENTED T1 RM; NO157Z AS UNIVATION XF TIBIA CEMENTED T2 RM; NO158Z AS UNIVATION XF TIBIA CEMENTED T3 RM; NO159Z AS UNIVATION XF TIBIA CEMENTED T4 RM; NO160Z AS UNIVATION XF TIBIA CEMENTED T5 RM; NO161Z AS UNIVATION XF TIBIA CEMENTED T6 RM; NO162Z AS UNIVATION XF TIBIA CEMENTED T1 LM; NO163Z AS UNIVATION XF TIBIA CEMENTED T2 LM; NO164Z AS UNIVATION XF TIBIA CEMENTED T3 LM; NO165Z AS UNIVATION XF TIBIA CEMENTED T4 LM; NO166Z AS UNIVATION XF TIBIA CEMENTED T5 LM; NO167Z AS UNIVATION XF TIBIA CEMENTED T6 LM; | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189032 | Z-2480-2021 | EcoGel 100 Imaging Gel, EcoGel 200 MultiPurpose Gel (Blue), EcoGel 200 MultiPurpose Gel (CLEAR), and EcoLotion Transmission Lotion 1) 100, 200 gel Intended for use in acoustic coupling of ultrasound transducer to tissue 2) Lotion Ecolotion transmission and massage lotion is intended to be used as an accessory to ultrasonic transducers and ultrasound monitors. Ecolotion is intended to be used externally only as a coupling medium for use in ultrasound procedures. | Eco-Med | 1 | CDRH | 2021-09-22 |  |  | AMTEC SALES Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189089 | Z-2508-2021 | Access SARS CoV-2 IgG (1st IS) Assay: Access Immunoassay Systems REAGENT PACK, REF C74339 Access SARS-CoV-2 IgG (1st IS) Calibrators, REF C74340 Access SARS-CoV-2 (1st IS) Quality Control (QC), REF C74341 | ACCESS SARS CoV-2 IgG (1st IS) Assay | 2 | CDRH | 2021-09-22 |  |  | Beckman Coulter Inc. | Firm's submission for Emergency Use Authorization (EUA) contained insufficient evidence to be granted EUA Authorization to distribute
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188389 | Z-2509-2021 | Univation X System knee implant devices as follows: Catalog/Product: NL470 UNIVATION F MENISCAL COMP.T1 RM/LM 7MM; NL471 UNIVATION F MENISCAL COMP.T2 RM/LM 7MM; NL472 UNIVATION F MENISCAL COMP.T3 RM/LM 7MM NL473 UNIVATION F MENISCAL COMP.T4 RM/LM 7MM NL474 UNIVATION F MENISCAL COMP.T5 RM/LM 7MM NL475 UNIVATION F MENISCAL COMP.T6 RM/LM 7MM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188909 | Z-2502-2021 | Centricity PACS-IW with Universal Viewer - Product Usage: Is a device that displays medical images (including mammograms) and data from various imaging sources. | Centricity PACS-IW with Universal Viewer | 2 | CDRH | 2021-09-22 |  |  | GE Healthcare, LLC | Image acquisition failures and synchronization failure with the Centricity Enterprise Archive
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188390 | Z-2510-2021 | Univation X System knee implant devices as follows: Catalog/Product: NL476 UNIVATION F MENISCAL COMP.T1 RM/LM 8MM; NL477 UNIVATION F MENISCAL COMP.T2 RM/LM 8MM; NL478 UNIVATION F MENISCAL COMP.T3 RM/LM 8MM; NL479 UNIVATION F MENISCAL COMP.T4 RM/LM 8MM; NL480 UNIVATION F MENISCAL COMP.T5 RM/LM 8MM; NL481 UNIVATION F MENISCAL COMP.T6 RM/LM 8MM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188392 | Z-2512-2021 | Univation X System knee implant devices as follows: Catalog/Product: NL488 UNIVATION F MENISCAL COMP.T1 RM/LM 10MM; NL489 UNIVATION F MENISCAL COMP.T2 RM/LM 10MM; NL490 UNIVATION F MENISCAL COMP.T3 RM/LM 10MM; NL491 UNIVATION F MENISCAL COMP.T4 RM/LM 10MM; NL492 UNIVATION F MENISCAL COMP.T5 RM/LM 10MM; NL493 UNIVATION F MENISCAL COMP.T6 RM/LM 10MM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188393 | Z-2513-2021 | Univation X System knee implant devices as follows: Catalog/Product: NL494 UNIVATION F MENISCAL COMP.T1 RM/LM 11MM; NL495 UNIVATION F MENISCAL COMP.T2 RM/LM 11MM; NL496 UNIVATION F MENISCAL COMP.T3 RM/LM 11MM NL497 UNIVATION F MENISCAL COMP.T4 RM/LM 11MM NL498 UNIVATION F MENISCAL COMP.T5 RM/LM 11MM NL499 UNIVATION F MENISCAL COMP.T6 RM/LM 11MM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188581 | Z-2517-2021 | Sight OLO-E1/U1 Software and Operators Manual Software v2.56.1 Catalog #: 1 OLO-U1 Catalog #: 2 OLO-E1 revision 1.9, revision 1.10 | Sight OLO | 2 | CDRH | 2021-09-22 |  |  | SIGHT DIAGNOSTICS LTD | Discrepancy in labeling between the adult CBC reference ranges presented in the OLO Operator's Manual and the OLO software
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188391 | Z-2511-2021 | Univation X System knee implant devices as follows: Catalog/Product: NL482 UNIVATION F MENISCAL COMP.T1 RM/LM 9MM; NL483 UNIVATION F MENISCAL COMP.T2 RM/LM 9MM; NL484 UNIVATION F MENISCAL COMP.T3 RM/LM 9MM; NL485 UNIVATION F MENISCAL COMP.T4 RM/LM 9MM; NL486 UNIVATION F MENISCAL COMP.T5 RM/LM 9MM; NL487 UNIVATION F MENISCAL COMP.T6 RM/LM 9MM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188395 | Z-2515-2021 | Univation X System knee implant devices as follows: Catalog/Product: NO180Z AS UNIVATION XF FEMUR CEMENTED F1 RM; NO181Z AS UNIVATION XF FEMUR CEMENTED F2 RM; NO182Z AS UNIVATION XF FEMUR CEMENTED F3 RM; NO183Z AS UNIVATION XF FEMUR CEMENTED F4 RM; NO184Z AS UNIVATION XF FEMUR CEMENTED F5 RM; NO185Z AS UNIVATION XF FEMUR CEMENTED F1 LM; NO186Z AS UNIVATION XF FEMUR CEMENTED F2 LM; NO187Z AS UNIVATION XF FEMUR CEMENTED F3 LM; NO188Z AS UNIVATION XF FEMUR CEMENTED F4 LM; NO189Z AS UNIVATION XF FEMUR CEMENTED F5 LM | Aesculap | 2 | CDRH | 2021-09-22 |  |  | Aesculap Implant Systems LLC | Malfunction- loosening of the implant resulting in a potential revision surgery
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188919 | Z-2503-2021 | Centricity Universal Viewer and Viewer 6.0 - Product Usage: Is a device that displays medical images (including mammograms) and data from various imaging sources. | Centricity Universal Viewer, Centricity Universal Viewer 6.0 | 2 | CDRH | 2021-09-22 |  |  | GE Healthcare, LLC | Image acquisition failures and synchronization failure with the Centricity Enterprise Archive
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188893 | Z-2501-2021 | Scorpion Portal Vein Access Set - Product Usage: used to gain access to the hepatic vein and guide a sharp puncture tool (0.040 Stylet or 17ga Needle) through the parenchyma. | Scorpion Portal Vein Access Set | 2 | CDRH | 2021-09-21 |  |  | Argon Medical Devices, Inc | As a result of design changes, sheaths have exhibited cracking/breaking at the tips.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188993 | Z-2497-2021 | EVIS EXERA BF-XT160 Bronchofibervideoscope, Model No. BF-XT160 | EVIS EXERA BF-XT160 Bronchofibervideoscope | 2 | CDRH | 2021-09-21 |  |  | Olympus Corporation of the Americas | A gluing step was not performed during the manufacturing of the affected device's insertion tube. There is a potential for the insertion tube to develop a leak, which poses an infection control risk.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189029 | Z-2500-2021 | Computed tomography x-ray systems with software syngo.CT VA20A_SP4a, VA20A_SP5, VA30A_SP2, VA30A_SP2a, VA30A_SP3,VA30A_FP2 in Somatom systems: SOMATOM go.Up - Model 11061620 SOMATOM go.Up - Model 11061628 SOMATOM go.All - Model 11061630 SOMATOM go.Top - Model 11061640 SOMATOM X.cite - Model 11330001 Computed tomography systems intended to generate and process cross-sectional images of patients by computer reconstruction of x-ray transmission data. | Siemens | 2 | CDRH | 2021-09-21 |  |  | Siemens Medical Solutions USA, Inc | Software versions may result in sporadic problems causing scanning workflow interruptions and unexpected user notifications. Sporadic software errors may also occur during interventional workflows, resulting in delay in diagnosis or scan aborts with the necessity for patient rescan may occur
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189072 | Z-2498-2021 | cobas SARS-CoV-2 & Influenza A/B Test for use on the cobas Liat System, Catalog Number 09211101190 | cobas SARS-CoV-2 & Influenza A/B Test | 2 | CDRH | 2021-09-21 |  |  | Roche Molecular Systems, Inc. | Customers have reported an increased number of false positive SARS-CoV-2 results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188623 | Z-2491-2021 | Mammary Nodules Localization | Automark, Mark, Markslim, Repomark | 2 | CDRH | 2021-09-20 13:13 |  |  | M.D.L. S.r.l. | Sterility assurance may be compromised.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188625 | Z-2493-2021 | General Manual surgical instrument | Soft Tissue Biopsy | 2 | CDRH | 2021-09-20 13:13 |  |  | M.D.L. S.r.l. | Sterility assurance may be compromised.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188624 | Z-2492-2021 | Soft Tissue Biopsy Needles | Soft Tissue Biopsy | 2 | CDRH | 2021-09-20 13:13 |  |  | M.D.L. S.r.l. | Sterility assurance may be compromised.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188621 | Z-2489-2021 | Bone Marrow Transplantation biopsy needle | Harvest, Illy, Jamblu, OsteoJ, Pen-Bone, Pickup, Pressto | 2 | CDRH | 2021-09-20 13:13 |  |  | M.D.L. S.r.l. | Sterility assurance may be compromised.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188622 | Z-2490-2021 | Aspirated Cyto-Histological Biopsy needle | Cytocut | 2 | CDRH | 2021-09-20 13:13 |  |  | M.D.L. S.r.l. | Sterility assurance may be compromised.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189040 | Z-2494-2021 | Ad-Tech Lightweight CABRIO Cable - Product Usage: Designed for the purpose of connecting these electrodes to third party monitors/stimulators. | Lightweight CABRIO Cable | 2 | CDRH | 2021-09-20 |  |  | Ad-Tech Medical Instrument Corporation | Incorrect version of labels were used.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189046 | Z-2496-2021 | Ad-Tech TECH ATTACH Cable- Product Usage: Designed for the purpose of connecting these electrodes to third party monitors/stimulators. | TECH-ATTACH Connection System Cable | 2 | CDRH | 2021-09-20 |  |  | Ad-Tech Medical Instrument Corporation | Incorrect version of labels were used.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189044 | Z-2495-2021 | Ad-Tech Lightweight TECH ATTACH Cable - Product Usage: Designed for the purpose of connecting these electrodes to third party monitors/stimulators. | Lightweight CABRIO Cable | 2 | CDRH | 2021-09-20 |  |  | Ad-Tech Medical Instrument Corporation | Incorrect version of labels were used.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189026 | Z-2487-2021 | SARS-CoV-2 Process Control (Swab) - Product Usage: intended for use as an external, non-viable control material to evaluate the performance of nucleic acid tests that detect SARS-CoV-2 virus. | HELIX¦ELITE SARS-CoV-2 Process Control (Swab) | 2 | CDRH | 2021-09-17 |  |  | Microbiologics Inc | The product indicated is contaminated with a plasmid DNA that contains the target SARS-CoV-2 sequences.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188886 | Z-2483-2021 | EPIQ Diagnostic Ultrasound Systems: Software Versions 5.0, 5.0.1 and 5.0.2 Models: EPIQ 5 - 795120 EPIQ 5C - 795205 EPIQ 5G - 795204 EPIQ 7 - 795117 EPIQ 7C - 795201 EPIQ 7G - 795200 EPIQ CVxi -795232 EPIQ CVx - 795231 | Philips EPIQ Diagnostic Ultrasound Systems | 2 | CDRH | 2021-09-17 |  |  | Philips Ultrasound Inc | Due to a software defect that can intermittently cause the system to lock-up which exiting Review Mode while performing an exam.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189025 | Z-2486-2021 | SARS-CoV-2 Process Control (Pellet) - Product Usage: intended for use as an external, non-viable control material to evaluate the performance of nucleic acid tests that detect SARS-CoV-2 virus. | HELIX¦ELITE SARS-CoV-2 Process Control (Pellet) | 2 | CDRH | 2021-09-17 |  |  | Microbiologics Inc | The SARS-CoV-2 Process Control (Pellet) contains RNA transcripts that include diagnostically relevant SARS-CoV-2 targets (CDC and WHO consensus sequences).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188925 | Z-2484-2021 | Affiniti Diagnostic Ultrasound Systems: Software Versions 5.0, 5.0.1 and 5.0.2 Models: Affiniti 30 795121 Affiniti 30 795218 Affiniti 50 795208 Affiniti 50 795118 Affiniti 70 795210 Affiniti 70 795119 | Philips Afiniti Diagnostic Ultrasound Systems | 2 | CDRH | 2021-09-17 |  |  | Philips Ultrasound Inc | Due to a software defect that can intermittently cause the system to lock-up which exiting Review Mode while performing an exam.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188985 | Z-2488-2021 | Phantom Nail System, Drill, 4.6 x 300mm, Cannulated, 3/16" Sq. Connection p28 P99-110-4630 | Phantom Hindfoot TTC/TC Nail System | 2 | CDRH | 2021-09-17 |  |  | Paragon 28, Inc. | A thin wall condition between the inner cannulation and the flutes of the drill which potentially could cause a fracture through the wall of the drill to the cannulation beneath.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188773 | Z-2469-2021 | Ultrasound Gel labeled under the following brand names: MediChoice Ultrasound Gel, Athena Liquasonic Ultrasound Gel, Eco-Med EcoGel Ultrasound Gel, MAC Medical Supply Ultrasound Gel, Eco-Med EcoGel 200 Blue High Velocity Ultrasound Gel, Eco-Med EcoGel 200 Multi-Purpose Ultrasound Gel, and Eco-Med EcoGel 200 Multi-Purpose CLEAR Ultrasound Gel | Multipe (see code info) | 1 | CDRH | 2021-09-17 |  |  | Mac Medical Supply Co Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188956 | Z-2478-2021 | Medtronic HVAD Pump Implant Kits | Medtronic HVAD" Pump Implant Kits | 2 | CDRH | 2021-09-16 |  |  | Heartware, Inc. | There is more force required to slide the driveline cover (boot cover) away from the connector after completion of a Driveline Strain Relief Repair than over an unrepaired strain relief. This may lead to difficulty in pulling back the driveline cover when attempting to access the Driveline connector during a controller exchange. Additionally, sliding the driveline cover over a strain relief repair may compromise the repair.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188768 | Z-2482-2021 | Braun ThermoScan PRO 6000 Ear Thermometer-indicated for the intermittent measurement of human body temperature for patients having ages ranging from normal weight (full term) newborn to geriatric adults in a professional use environment. Model: 6000 | Braun | 2 | CDRH | 2021-09-16 |  |  | Welch Allyn Inc | If the device is exposed to fluid ingress and is used before the cleaning fluid has completely dried, there is a risk of the device overheating potentially causing a burn to the user or the ear canal of the patient
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188707 | Z-2481-2021 | BOND Ready-To-Use Primary Antibody CDX2 (EP25), REF PA0375 | BOND | 2 | CDRH | 2021-09-16 |  |  | Leica Microsystems, Inc. | Product may not perform as specified in IFU.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188655 | Z-2479-2021 | Maxi Sky 1000 (Arjo) and V10 (BHM) Bariatric Ceiling Rail - Mounted Lift with measuring function | Maxi Sky 1000 (Arjo) and V10 (BHM) | 2 | CDRH | 2021-09-16 |  |  | Arjohuntleigh Magog, Inc. | The spreader bar can disconnect from the scale bottom attachment.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188532 | Z-2431-2021 | Cordis SUPER TORQUE MB 5F PIG, Angiographic Catheter, REF 532-598B | SUPER TORQUE MB 5F PIG | 1 | CDRH | 2021-09-15 |  |  | Cordis Corporation | Cordis has identified that the SUPER TORQUE MB Angiographic Catheter (with Marker Bands) are being used in a manner that may allow entrapment of the catheter between endovascular devices and the vessel wall, which can lead to marker band movement or dislodgement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188822 | Z-2477-2021 | Lyra Direct SARS-CoV-2 Assay Emergency Use Authorization- Ref: M124. The Lyra Direct SARS-CoV-2 Assay is a real-time RT-PCR assay intended for the qualitative detection of nucleic acid from SARS-CoV-2 in nasal (NS), nasopharyngeal (NP), or oropharyngeal (OP) direct swab specimens from individuals suspected of COVID-19 by their healthcare provider | Quidel | 2 | CDRH | 2021-09-15 |  |  | Quidel Corporation | Revised Instructions for Use: Potential for specimens from patients with high SARS-CoV-2 viral loads to be falsely interpreted as negative if target amplification occurs prior to a cycle-threshold (Ct) value of 5 when using the Thermo Fisher QuantStudio 7 Pro, the Applied Biosystems 7500 Fast Dx, the Applied Biosystems 7500, the Bio-Rad CFX96 Touch", the Roche LightCycler LC 480 II/ cobas z 480, or the Qiagen Rotor-Gene¿ MDx.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188533 | Z-2432-2021 | Cordis SUPER TORQUE MB 5F PIG, Angiographic Catheter, REF 532-598C | SUPER TORQUE MB 5F PIG | 1 | CDRH | 2021-09-15 |  |  | Cordis Corporation | Cordis has identified that the SUPER TORQUE MB Angiographic Catheter (with Marker Bands) are being used in a manner that may allow entrapment of the catheter between endovascular devices and the vessel wall, which can lead to marker band movement or dislodgement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188530 | Z-2429-2021 | Cordis 5F SUPER TORQUE PIG PIGTAIL SPECIAL, Angiographic Catheter, REF SRD7040MB | 5F SUPER TORQUE PIG PIGTAIL SPECIAL | 1 | CDRH | 2021-09-15 |  |  | Cordis Corporation | Cordis has identified that the SUPER TORQUE MB Angiographic Catheter (with Marker Bands) are being used in a manner that may allow entrapment of the catheter between endovascular devices and the vessel wall, which can lead to marker band movement or dislodgement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188531 | Z-2430-2021 | Cordis SUPER TORQUE MB 5F PIG, Angiographic Catheter, REF 532-598A | SUPER TORQUE MB 5F PIG | 1 | CDRH | 2021-09-15 |  |  | Cordis Corporation | Cordis has identified that the SUPER TORQUE MB Angiographic Catheter (with Marker Bands) are being used in a manner that may allow entrapment of the catheter between endovascular devices and the vessel wall, which can lead to marker band movement or dislodgement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188529 | Z-2428-2021 | Cordis 5F UNIVERSAL FLUSH F4 SUPER TORQUE MB, SPECIAL, Angiographic Catheter, REF SRD6785MB | 5f UNIVERSAL FLUSH F4 SUPER TORQUE MB, SPECIAL | 1 | CDRH | 2021-09-15 |  |  | Cordis Corporation | Cordis has identified that the SUPER TORQUE MB Angiographic Catheter (with Marker Bands) are being used in a manner that may allow entrapment of the catheter between endovascular devices and the vessel wall, which can lead to marker band movement or dislodgement.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188994 | Z-2473-2021 | McKesson Lap Sponge, X-Ray Detectable 12"x 12", 5/pack, Sterile McKesson Part # 16-2112121 | McKesson Lap Sponge | 2 | CDRH | 2021-09-14 10:47 |  |  | Cypress Medical Products LLC | Manufacturer issued a recall due to an EO indicator color. The indicator should change from red to blue signifying sterilization. However, on one case of product, the indicators did not change
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188853 | Z-2472-2021 | Alinity ci-series System Control Module clinical chemistry and immunoassay analyzer for in vitro diagnostic use. Model number LN 03R70-01 | Alinity ci-series System Control Module (SCM) | 2 | CDRH | 2021-09-14 09:18 |  |  | Abbott Laboratories | Due to potential performance issues with software version 3.2.3 and earlier. Performance issues are: 1) There is the potential for onboard reagents past their lot expiration date or onboard stability time to be incorrectly displayed on the Reagent Status screen with a status of "OK". In this scenario, the expired reagent will remain in the carousel for sample processing. 2) Performing monthly maintenance procedure 5701 Clean ICT Drain Tip incorrectly may potentially cause issues such as damaged connector or leaking connection. 3) A sample exception with message code 150 "Unable to process test. Previous processing module error" is produced without notifying the operator of a reagent pipettor failure and may cause processing module to transition into the Pausing state. Tests initiated prior to the sample exception may continue to process without dispense of reagents, potentially leading to incorrect results. 4. Reagent short dispense: Alinity c needs a control measure to protect against the potential for short reagent dispense. Addition of a maximum depth limit for reagent cartridges on the Alinity c to prevent the potential for a short dispense is needed. 5. RSM indicators were blinking green incorrectly: The Reagent and Sample Manager (RSM) indicators will incorrectly blink green indicating that processing is completed and the rack or cartridge can be removed from the loading area when the rack or cartridge should not be accessed. 6. Erroneous ICT Calibration Curves: Erroneous ICT calibration curves might be generated for the Alinity c when there is a malfunction of ICT reference solution delivery, use of wrong/evaporated calibrator, defective ICT modules, etc.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188585 | Z-2471-2021 | ACUSON Juniper Diagnostic Ultrasound System | ACUSON Juniper ultrasound imaging system | 2 | CDRH | 2021-09-14 |  |  | Siemens Medical Solutions USA, Inc. | Due to intermittent failures of the power supply in the ultrasound system which renders it inoperable.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188963 | Z-2474-2021 | The Allura Xper series is intended for use on human patients to perform: "Vascular, cardiovascular and neurovascular imaging applications, including diagnostic, interventional and minimally invasive procedures. This includes, e.g., peripheral, cerebral, thoracic and abdominal angiography, as well as PTAs, stent placements, embolisations and thrombolysis. "Cardiac imaging applications including diagnostics, interventional and minimally invasive procedures (such as PTCA, stent placing, atherectomies), pacemaker implanatations, and electrophysiology (EP). "Non-vascular interventions such as drainages, biopsies and vertebroplasties procedures." | Philips Allura Xper FD20 | 2 | CDRH | 2021-09-14 |  |  | Philips North America Llc | Early life failure of the ADUs only occurs when the hospital mains provide 480V to the system. This current voltage may lead to saturation and overheating of the coils that protect the IGBTs (insulated-gate bipolar transistor) of the ADU. The overheating may generate a peculiar burning smell that may be noticed. In some instances, this burning smell may be noticed before the ADU fails. When the ADU fails the X-ray performance of the system is reduced to Emergency Fluoroscopy. Exposure will not be possible and image quality will be reduced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188964 | Z-2475-2021 | The Allura Xper series is intended for use on human patients to perform: "Vascular, cardiovascular and neurovascular imaging applications, including diagnostic, interventional and minimally invasive procedures. This includes, e.g., peripheral, cerebral, thoracic and abdominal angiography, as well as PTAs, stent placements, embolisations and thrombolysis." "Cardiac imaging applications including diagnostics, interventional and minimally invasive procedures (such as PTCA, stent placing, atherectomies), pacemaker implanatations, and electrophysiology (EP). Non-vascular interventions such as drainages, biopsies and vertebroplasties procedures." | Philips Allura Xper FD20/20 | 2 | CDRH | 2021-09-14 |  |  | Philips North America Llc | Early life failure of the ADUs only occurs when the hospital mains provide 480V to the system. This current voltage may lead to saturation and overheating of the coils that protect the IGBTs (insulated-gate bipolar transistor) of the ADU. The overheating may generate a peculiar burning smell that may be noticed. In some instances, this burning smell may be noticed before the ADU fails. When the ADU fails the X-ray performance of the system is reduced to Emergency Fluoroscopy. Exposure will not be possible and image quality will be reduced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188965 | Z-2476-2021 | The Azurion series (within the limits of the used Operation Room table) are intended for use to perform: "Image guidance in diagnostic, interventional and minimally invasive surgery procedures for the following clinical application areas: vascular, non-vascular, cardiovascular and neuro procedures. "Cardiac imaging applications including diagnostics, interventional and minimally invasive surgery procedures. Additionally: "The Azurion series can be used in a hybrid Operation Room. "The Azurion series contain a number of features to support a flexible and patient centric procedural workflow. | Philips Azurion 7 M20 | 2 | CDRH | 2021-09-14 |  |  | Philips North America Llc | Early life failure of the ADUs only occurs when the hospital mains provide 480V to the system. This current voltage may lead to saturation and overheating of the coils that protect the IGBTs (insulated-gate bipolar transistor) of the ADU. The overheating may generate a peculiar burning smell that may be noticed. In some instances, this burning smell may be noticed before the ADU fails. When the ADU fails the X-ray performance of the system is reduced to Emergency Fluoroscopy. Exposure will not be possible and image quality will be reduced.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188128 | Z-2455-2021 | BD SmartSite 'Y' Extension Set 2 Needle-Free Valves, REF: 20019E7D | BD SmartSite 'Y' Extension Set 2 Needle-Free Valves | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188695 | Z-2464-2021 | BD ADD-ON Burette Set V/NV Needle-Free Valve, REF: 82113E-0006 | BD ADD-ON Burette Set V/NV Needle-Free Valve | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188696 | Z-2465-2021 | BD SE Burette Set 20 Drops 2 Injection Ports Vented/NonVented, REF: 72103E-0006 | BD SE Burette Set 20 Drops 2 Injection Ports Vented/NonVented | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188724 | Z-2433-2021 | AMSORB PLUS PREFILLED G-CAN 1.0L; Carbon dioxide absorbent canister; Product codes AMAB3801 and AMAB3801GE. Used in anesthesia. | AMSORB PLUS PREFILLED G-CAN 1.0L | 1 | CDRH | 2021-09-13 |  |  | Armstrong Medical Services Limited | Some devices may be associated with high and unexpected resistance to gas flow during clinical use, impacting or preventing adequate ventilation of an anaesthetised patient, which could cause a delay in treatment. Not all defective canisters will be identified by the pre-use test.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188187 | Z-2456-2021 | BD Gravity Burette Set 60 DROP 1.2 MF 2 SmartSite VALVES BALL VALVE DRIP CHAMBER V/NV, REF: 10037031; BD Gravity Burette Set 20 DROP 1.2 Micron Filter 2 SmartSite VALVES BALL VALVE DRIP CHAMBER V/NV , REF: 10037032; BD GRAVITY BURETTE SET 60 DROP 2 Smartsite VALVES BALL DRIP CHAMBER \VENTED/NONVENTED, REF: 82104E; BD ADD-ON BURETTE SET Smartsite VALVE, REF: 82113E; BD ADD-ON BURETTE SET Smartsite VALVE VENTED/NONVENTED, REF: 82114E; BD ADD-ON BURETTE SET Smartsite VALVE BALL VALVE DRIP CHAMBER VENTED/NONVENTED, REF: 82115E | BD Gravity Burette Set 60 DROP 1.2 MF 2 SmartSite VALVES BALL VALVE DRIP CHAMBER V/NV | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188193 | Z-2459-2021 | BD Alaris Pump Infusion Buerette Set 0.2 Micron Filter Smallbore Tubing SmartSite Port (Burette) SmartSite Y-Site, REF: 10015012; BD Alaris Pump Infusion Buerette Set Ball Valve Smallbore Tubing Smartsite Port (Burette) SmartSite Y-Site, REF: 10817920; BD Alaris Pump Infusion Buerette Set SmartSite Port (Burette) 3 SmartSite Y-Sites, REF: 10821753; BD Alaris Pump Infusion Buerette Smallbore Tubing Smartsite Port (Burette) SmartSite Y-Site, REF: 11607787; BD Alaris Pump Infusion Buerette Set 0.2Micron Filter Ball Valve SmartSite Port (Burette) 2 SmartSite Y-Sites, REF: 11613191; BD Alaris Pump Infusion Buerette Set Smartsite Port (Burette) 3 SmartSite Y-Sites, REF: 2440-0600; BD Alaris Pump Infusion Buerette Set Smallbore Tubing Smartsite Port (Burette) 3 SmartSite Y-Site, REF: 2441-0007; BD Alaris Pump Infusion Buerette Set Low Sorbing Tubing Smartsite Port (Burette), REF: 2444-0007; BD Alaris Pump Infusion Buerette Set Ball Valve Smartsite Port (Burette) 2 Smartsite Y-Sites, REF: 2447-0007; BD Alaris Pump Infusion Set Back Check Valve 4-Way Stopcock 3 Smartsite Y-Sites Extension Set With SmartSite Port, REF: C24104E | BD Alaris Pump Infusion Buerette Set 0.2 Micron Filter Smallbore Tubing SmartSite Port (Burette) Sma | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188670 | Z-2453-2021 | Configura Advance Chair, Model Number: CHAIR-0-SC1-030; Advance Chair, 100mm Castors, Std Backrest, Manual; Serial Numbers: 10001-SC1-20 to 10483-SC1-20; 10001-SC1-21 to 10224-SC1-21; 10231-SC1-21 to 10456-SC1-21; 10481-SC1-21 to 10570-SC1-21. Model Number: CHAIR-0-SC2-030; Advance Chair, 100mm (4in) Castors, Std Backrest, Electric; Serial Numbers: 10001-SC2-20 to 10135-SC2-20; 10001-SC2-21 to 10240-SC2-21. Model Number: CHAIR-0-SC3-050; Advance Chair, 125/150mm (5in/6in) Castors, Standard Backrest, Manual; Serial Numbers: 10116-SC3-20 to 10145-SC3-20; 10236-SC3-20 to 10295-SC3-20; 10001-SC3-21 to 10090-SC3-21; 10271-SC3-21 to 10360-SC3-21; 10001-SC3-20 to 10115-SC3-20; 10146-SC3-20 to 10235-SC3-20. Model Number: CHAIR-0-SC4-050; Advance Chair, 125/150mm (5in/6in) Castors, Standard Backrest, Electric; Serial Numbers: 10001-SC4-20 to 10030-SC4-20. | Configura Advance Chair | 2 | CDRH | 2021-09-13 |  |  | Accora Inc | The firm has identified a potential for the backrest to become detached from the chair.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188693 | Z-2463-2021 | BD Gravity IV Set 3-port Closed Manifold Back Check Valve 3 Needle-Free Y-sites, REF: CM42500E-07; BD Gravity IV Set 3-port Closed Stopock Manifold Back Check Valve 2 Needle-Free Y-sites, REF: CS42522E-07; BD GRAVITY 10 DP V/NV CHECK VALVE 3-WAY STOPCOCK 4 Smartsite VALVES & SmartSite EXT SET, REF: 42407E-05 | BD Gravity IV Set 3-port Closed Manifold Back Check Valve 3 Needle-Free Y-sites | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188134 | Z-2452-2021 | Cellex qSARS-Cov-2 Antigen Rapid Test, 25 tests/box, submitted as EUA202955 "declined to issue" | Cellex qSARS-Cov-2 Antigen Rapid Test | 2 | CDRH | 2021-09-13 |  |  | Cellex | The kit does not have an emergency use authorization (EUA).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188196 | Z-2460-2021 | BD Extension Set SMALLBORE Tubing SmartSite VALVE, REF: 10010511; BD SmartSite Extension Set Needle-Free Valve, REF: 10010912; BD EXTENSION SET SMALLBORE TUBING DETACHABLE SmartSite VALVE, REF: 10010983; BD Gravity Burette Set 60 DP 4 VersaSafe PORTS 3 SmartSite VALVES BALL VALVE DRIP CHAMBER V/NV, REF: 10011087; BD EXTENSION SET SMALLBORE TUBING SmartSite VALVE (DETACHABLE), REF: 10011253; BD EXTENSION SET LOW SORBING TUBING 0.2 MICRON FILTER SmartSite VALVE, REF: 10013902; BD EXTENSION SET 2 SmartSite VALVES Y-CONNECTOR, REF: 10014692; BD EXTENSION SET SMALLBORE TUBING 2 DETACHABLE SmartSite VALVES, REF: 10015488; BD SmartSite NEEDLE FREE VALVE, REF: 2000E; BD Extension Set J-LOOP SMALLBORE Tubing SmartSite VALVE, REF: 20021E; BD Extension Set SMALLBORE Tubing SmartSite VALVE, REF: 20035E; BD Extension Set SMALLBORE Tubing SmartSite VALVE SLIP LUER, REF: 20037E; BD Extension Set TRI-PORT SMALLBORE Tubing 3 SmartSite VALVES, REF: 20038E; BD Extension Set SMALLBORE Tubing SmartSite VALVE, REF: 20039E; BD Extension Set SmartSite VALVE SLIP LUER, REF: 20057E; BD Extension Set SmartSite VALVE, REF: 20059E; BD extension Set Y-CONNECTOR 2 SmartSite VALVES, REF: 20061E; BD EXTENSION SET TRI-PORT SMALLBORE TUBING CHECK VALVE 3 SmartSite VALVES, REF: 20062E; BD EXTENSION SET Y-CONNECTOR 2 SmartSite VALVE, REF: 20159E; BD Extension Set TRI-PORT SMALLBORE TUBING SmartSite VALVE DETACHABLE, REF: 20338E; BD EXTENSION SET SmartSite VALVE, REF: 21059E; BD 8.5-IN PRESSURE RATED EXTENSION SET SmartSite VALVE, REF: 22003E-07; BD Texium SmartSite VALVE MICROBORE EXTENSION SET 60 IN, REF: 22005E-07T; BD EXTENSION SET SmartSite VALVE (DETACHABLE), REF: 22059E; BD EXTENSION SET SmartSite VALVE (DETACHABLE) SLIP LUER, REF: C25015E; BD SmartSite Syringe-Syringe Transfer Device, REF: MV0450 | BD Extension Set SMALLBORE Tubing SmartSite VALVE | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188652 | Z-2462-2021 | BD Gravity Burette Set 60 DP 3 SmartSite VALVES BALL VALVE DRIP CHAMBER, REF: 10012564 | BD Gravity Burette Set 60 DP 3 SmartSite VALVES BALL VALVE DRIP CHAMBER | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188699 | Z-2468-2021 | BD SmartSite Bag Access Device Needle-Free Valve, REF: 2300E; BD Bag access Device with 1 Check Valve And Needle Free Valve, REF: 2309E-0006 | BD SmartSite Bag Access Device Needle-Free Valve | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188197 | Z-2461-2021 | BD SmartSite Bag Access Device, REF: 2300E-0500; BD BAG ACCESS DEVICE CHECK VALVE SmartSite VALVE, REF: 2309E | BD SmartSite Bag Access Device | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188697 | Z-2466-2021 | BD Alaris Pump Infusion Buerette Set Smallbore Tubing Smartsite Port (Burette) 3 SmartSite Y-Sites, REF: 2441-0006 | BD Alaris Pump Infusion Buerette Set Smallbore Tubing Smartsite Port (Burette) 3 SmartSite Y-Sites | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188127 | Z-2454-2021 | BD Extension Set SMALLBORE Tubing with 2 SmartSite VALVES, REF: 20019E | BD Extension Set SMALLBORE Tubing with 2 SmartSite VALVES | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188191 | Z-2457-2021 | BD Syringe Set 20 DP MIC TBG 0.2MF CKV Anti-Siphon SmartSite VLV Sensing Disc, REF: 10798696; BD Syringe Set 20 DP MIC TBG CKV Anti-Siphon SmartSite VLV Sensing Disc, REF: 10798703; | BD Syringe Set 20 DP MIC TBG 0.2MF CKV Anti-Siphon SmartSite VLV Sensing Disc | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188698 | Z-2467-2021 | BD SmartSite, REF: 2000E CHINA; BD SmartSite Connector, REF: 2000E-04; BD SmartSite J-LOOP Extension Set One Needle Free Valve, REF: 20021E7D; BD SmartSite Extension Set One Needle Free Valve Slip Luer, REF: 20037E7D; BD SmartSite Extension Set 3 Needle Free Valves, REF: 20038E7D; BD SmartSite Extension Set One Needle Free Valve, REF: 20039E7D; BD SmartSite Extension Set 1 Needle-Free Valve, 20049E7D; BD Extension Set 1 Needle-Free Valve, REF: 20059E-0006; BD SmartSite Extension Set 1 Needle-Free Valve, REF: 20060E7D; BD SmartSite Extension Set 2 Needle Free Valves, REF: 20061E7D; BD SmartSite Extension Set Check Valve 3 Needle Free Valves, REF: 20062E7D; BD SmartSite Extension Set 2 Needle Free Valves, REF: 20159E7D; BD Alaris GP Series SmartSite Burette Set, REF: 60103E; BD Alaris GP Series Burette Set No Filter, REF: 63441E; Alaris SmartSite needle-free valve, REF: 2000E7D | BD SmartSite, REF: 2000E CHINA; BD SmartSite Connector | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188583 | Z-2470-2021 | Arjo Disposable Repositioning Sling, Model AHD001 | Arjo Disposable Repositioning Sling | 2 | CDRH | 2021-09-13 |  |  | Getinge Dominican Republic SA | Sling loop straps on device pose a trip hazard.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188192 | Z-2458-2021 | BD BURETTE GRAVITY SET 60 DROP 2 Smartsite VALVES VENTED/NONVENTED, REF: 42163E | BD BURETTE GRAVITY SET 60 DROP 2 Smartsite VALVES VENTED/NONVENTED | 2 | CDRH | 2021-09-13 |  |  | CAREFUSION | Needle-free connector valves, either standalone or attached to extension sets, may have issues including difficulty to flush, flow issues, and partial or total occlusions, which could lead to therapy delay. Follow-up activates not required if product has already been used or was safely connected to the female luer of connecting device.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188668 | Z-2448-2021 | Colonoscope Family # 1-Pentax Video Colonoscopes (EC Family) with One Instrument Channel and a Water Jet Channel Models: EC-2990Li, EC-3490Li, EC-3890Li, EC-3490TLi, EC-3490LK, EC-3890LK, EC3890LZi, EC34-i10L, EC38-i10L; Colonoscope Family # 2-Pentax Video Colonoscopes with Two Instrument Channels and a Water Jet Channel Model: EC-3890TLK | Pentax Medical Video Colonoscopes | 2 | CDRH | 2021-09-10 10:59 |  |  | Pentax of America Inc | Updated Reprocessing Instruction For Use (rIFU): Impacted PENTAX Medical Video Upper GI Scopes (EG) and Video Colonoscopes (EC) families
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188669 | Z-2449-2021 | Gastroscope Family # 1-Gastroscopes without a Water Jet Channel Models: EG-2790i, EG-1690K, EG-2490K, EG-2790K, EG27-i10; Gastroscope Family # 2- Gastroscopes with a Water Jet Channel Models: EG-2990i, EG-2990K, EG-3490K, EG29-i10; Gastroscope Family # 3-Gastroscopes with Two Instrument Channels and a Water Jet Channel Models: EG-3890TK | Pentax Medical Video Upper GI Scopes | 2 | CDRH | 2021-09-10 10:59 |  |  | Pentax of America Inc | Updated Reprocessing Instruction For Use (rIFU): Impacted PENTAX Medical Video Upper GI Scopes (EG) and Video Colonoscopes (EC) families
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189039 | Z-2447-2021 | MR MAGNETOM Systems with RT Image Suite with syngo MR VA20A, VA30A and VA31A with license SW-Key CT_RT_IMA_SUITE_SYNTHETIC_CT Material Numbers: 11060815 | Siemens Syngo.MR | 2 | CDRH | 2021-09-10 |  |  | Siemens Medical Solutions USA, Inc | Potential for data loss when using the contouring or patient marking workflow when editing a structure set after saving and reopening a study, resulting in an incomplete structure set sent to the treatment planning system (TPS).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189037 | Z-2445-2021 | syngo.via RT Image Suite with syngo.via VB30, VB40, VB50 and VB60- Intended for Radiation Therapy Treatment Material Number: 10496180 | Siemens Syngo.via RT Image Suite | 2 | CDRH | 2021-09-10 |  |  | Siemens Medical Solutions USA, Inc | Potential for data loss when using the contouring or patient marking workflow when editing a structure set after saving and reopening a study, resulting in an incomplete structure set sent to the treatment planning system (TPS).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188701 | Z-2434-2021 | Proteus 235- designed to produce and deliver a proton beam for the treatment of patients with localized tumors and other conditions susceptible to treatment by radiation | IBA ProteusPLUS and ProteusONE | 2 | CDRH | 2021-09-10 | 2021-09-08 21:07 |  | Ion Beam Applications S.A. | Users may be misled by the popup message displayed by the Proton Therapy System when the Oncology Information System appears disconnected, may cause mistreatment
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188702 | Z-2450-2021 | Potassium Hydroxide 10% (KOH 10%), Cat. No. 300145 Potassium Hydroxide 10% (KOH 10%) is intended for use in the demonstration of yeast and fungal elements in wet clinical preparations. Potassium Hydroxide digests the proteinaceous material in the specimen, leaving the fungal cell wall intact for direct microscopic observation. | KOH 10% | 2 | CDRH | 2021-09-10 |  |  | Gibson Bioscience | A lot of KOH 10% was made with Hydrogen Peroxide 3% instead of KOH 10%.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188611 | Z-2451-2021 | Verigene EP Amplification Reagent Kit Test | Verigene | 2 | CDRH | 2021-09-10 |  |  | Luminex Corporation | There is potential for false positive results.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188593 | Z-2443-2021 | The Hero Model 100 (H-100) is a system consisting of a smart pill dispenser and a medication management app. | Hero Model 100 (H-100) | 2 | CDRH | 2021-09-10 |  |  | Hero Health | Due to dispenser prompting for a missed dose that had already been dispensed
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188571 | Z-2444-2021 | cobas infinity central lab running software version 3.01.03 through 3.02.08, Catalog#: 07154003001 | cobas infinity central lab / cobas infinity core license | 2 | CDRH | 2021-09-10 |  |  | Roche Diagnostics Operations, Inc. | Under specific circumstances created by the user, the cobas e flow test results could be replaced by an automatic result sent to the lab LIS. This automatic result could be released and interpreted by your Laboratory Information System (LIS) as a false positive or negative result. It should be noted that the "correct" result would still be available on the instrument.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189038 | Z-2446-2021 | syngo.CT VA20A, VA30A and VA40A in: syngo.via RT Image Suite / Sim&Go SOMATOM go.Now SOMATOM go.Up SOMATOM go.All SOMATOM go.Top SOMATOM go.Sim SOMATOM go.Open Pro and SOMATOM X.cite Material Numbers: 11061620 11061628 11061630 11061640 11061660 11061670 | Siemens Syngo.CT | 2 | CDRH | 2021-09-10 |  |  | Siemens Medical Solutions USA, Inc | Potential for data loss when using the contouring or patient marking workflow when editing a structure set after saving and reopening a study, resulting in an incomplete structure set sent to the treatment planning system (TPS).
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188588 | Z-2441-2021 | MIVI Super 90 Guide Catheters, MIVI Mi-Axus 8F Catheters, REF MIA-9080S (80 cm length); MIA-9090S (90 cm length), and MIA-9095S (95 cm length), MIA-9080-E, MIA-9090S-E, MIA-9095S-E, Sterile. Device is used to facilitate the insertion and guidance of microcatheters into a selected blood vessel in the peripheral, coronary and neuro vascular systems. | MIVI Super 90 Guide Catheters | 2 | CDRH | 2021-09-09 17:02 |  |  | Mivi Neuroscience Inc | There is potential for nonsterility of product due to a possible defect in the pouch seal.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188894 | Z-2442-2021 | MIVI Super 90 8F Guide Catheter, 95 cm, REF MIA-9095S-IDE, For Investigational Use Only; sterile, and Super 90 8F Guide Catheter, 90 cm, REF MIA-9090S-IDE, For Investigational Use Only. Device is used to facilitate the insertion and guidance of microcatheters into a selected blood vessel in the peripheral, coronary and neuro vascular systems. | MIVI Super 90 Guide Catheters | 2 | CDRH | 2021-09-09 17:02 |  |  | Mivi Neuroscience Inc | There is potential for nonsterility of product due to a possible defect in the pouch seal.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188953 | Z-2391-2021 | LiquaSonic Ultrasound Gel 250mL, Model No. 001205 | LiquaSonic Ultrasound Gel 250mL | 1 | CDRH | 2021-09-09 13:46 |  |  | Athena Medical Products Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188954 | Z-2392-2021 | LiquaSonic Ultrasound Gel 5L, Model No. 001222 | LiquaSonic Ultrasound Gel 5L | 1 | CDRH | 2021-09-09 13:46 |  |  | Athena Medical Products Inc | Ultrasound gels and lotions manufactured by Eco-Med Pharmaceutical, Inc. may contain bacterial contamination. On August 18, 2021, FDA published a Letter to Health Care Providers to further address this issue: https://www.fda.gov/medical-devices/letters-health-care-providers/stop-using-all-eco-med-ultrasound-gels-and-lotions-due-risk-bacterial-contamination-letter-health.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188660 | Z-2440-2021 | COPAN FLOQSwabs, STERILE.EO, for drawing clinical samples. CE0123, for the following reference numbers REF:501CS01, Minitip, sterile Single Wrapped, Molded bp 80mm; REF: 20226, P/N 501CS01Q100, Minitip, Sterile, Single Wrapped, bp 60MM, BUNDLES OF 20 pcs; REF: 501CS01.Q20, P/N 1127700, Minitip Applicator Sterile Ind. wrapped., bp 60mm, bundles of 20; REF: 502CS01, Regular, Sterile single Wrapped, Molded bp 80mm; REF: 503CS01, Flexible, Sterile Single Wrapped, Molded bp 100mm; REF: 220252, P/N 503CS01.BD, Flexible, Sterile Single Wrapped Molded bp 100 mm; REF: 503CS01/503CS01.BX, Flexible, Sterile Single Wrapped, Molded bp 100mm REF: 50U004DS, hDNA Free, Regular sterile in dry tube Molded Break Point 20 mm; REF: 50U009DS02, hDNA Free, Regular, Sterile, 2pcs Packed; REF: 518CS01, Minitip, Ster. Single Wrapped, Molded bp 100mm; REF: 220251, P/N: 518CS01.BD, Minitip, sterile Single Wrapped Molded bp 100mm; REF: 519C/519C.BX, Regular Sterile, Molded break point 100 mm; REF: 220250 P/N: 519CS01.BD, Regular, Sterile Single wrapped Molded bp 100mm; REF: 525CS01, Ultra-Thin Minitip, sterile single wrapped, molded bp 80 mm; REF: 528C, Large Tip Swab, Handle with Stopper, in dry tube REF: 534CS01, Flexible Nylon, Sterile Single Wrapped; REF: 551C, Minitip Sterile in dry tube, Molded Break Point 80 mm; REF: 552C/552C.BX, Regular sterile in dry tube Molded Break Point 80 mm; REF: 553C, Flexible sterile in dry tube, molded bp 100 mm; REF: 553C.LC, Flexible Sterile in dry tube, molded bp 100 mm; REF: 5U002S, Regular, sterile, in tube, 80 mm breakpoint, with low stopper; 5U048N01.BIO, REF: 5U055S02.HG, Regular FLOQSwabs with 20mm breakpoint 100 peelpouches, 2 swabs per peelpouch; REF: 5U056S.IUO, FLOQSwats ADULT IN DRY TUBE, Investigational Use Only; REF: 5UO57S.IUO, FLOQSwabs CHILD IN DRY TUBE, Investigational Use Only | FLOQSwab | 2 | CDRH | 2021-09-09 06:34 | 2021-09-09 21:06 |  | Copan Italia | A sterility assurance level of 10-6 cannot be guaranteed due to intentional fraud and data integrity issues at the contract sterilizer, Steril Milano.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188682 | Z-2438-2021 | Solero Generator Refurbished PG, Item No. H78712740000D0, Catalog No. 12740000D | Solero Microwave Tissue Ablation (MTA) System | 2 | CDRH | 2021-09-08 14:53 |  |  | Angiodynamics, Inc. | Specific serial numbers of the Solero MTA Generator require servicing to upgrade the software to help reduce the incidence of Error 0001, which can occur during system start-up.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188681 | Z-2437-2021 | Solero Generator US PG, Item No. H78712740000US0, Catalog No. 12740000US | Solero Microwave Tissue Ablation (MTA) System | 2 | CDRH | 2021-09-08 14:53 |  |  | Angiodynamics, Inc. | Specific serial numbers of the Solero MTA Generator require servicing to upgrade the software to help reduce the incidence of Error 0001, which can occur during system start-up.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188680 | Z-2436-2021 | Solero Generator Refurbished US PG, Item No. H78712740000USD0, Catalog No. 12740000USD | Solero Microwave Tissue Ablation (MTA) System | 2 | CDRH | 2021-09-08 14:53 |  |  | Angiodynamics, Inc. | Specific serial numbers of the Solero MTA Generator require servicing to upgrade the software to help reduce the incidence of Error 0001, which can occur during system start-up.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188683 | Z-2439-2021 | Solero Generator PG, Item No. H78712740000, Catalog No. 12740000 | Solero Microwave Tissue Ablation (MTA) System | 2 | CDRH | 2021-09-08 14:53 |  |  | Angiodynamics, Inc. | Specific serial numbers of the Solero MTA Generator require servicing to upgrade the software to help reduce the incidence of Error 0001, which can occur during system start-up.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188657 | Z-2435-2021 | Nihon Kohden NKV-550 Ventilator, REF: NKV-550-U, CE 0344, UDI: (01) 00843685100012 The Nihon Kohden NKV-550 Series Ventilator System is intended to provide continuous ventilation for adult, pediatric and neonatal patients who require invasive or noninvasive respiratory support. | Nihon Kohden NKV-550 Ventilator System | 2 | CDRH | 2021-09-08 14:31 |  |  | NIHON KOHDEN ORANGEMED, INC | Potential defective gas inlet pressure regulator within ventilator which could result in gas releasing noise from the bottom of the ventilator, a High Air/O2 Inlet Pressure alarm, a High/Low O2 alarm, or could not pass device Check prior to use.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188618 | Z-2408-2021 | FaStep COVID-19 IgG/IgM Rapid Test Device Qualitative detection and differentiation of IgM and IgG antibodies to SARS-CoV-2 in human venous whole blood (sodium EDTA), serum, plasma (sodium EDTA) and fingerstick whole blood. Intended for use as an aid in identifying individuals with an adaptive immune response to SARS-CoV-2, indicating recent or prior infection. (CLIA Waived Laboratories) | FaStep COVID-19 IgG/IgM Rapid Test Device | 2 | CDRH | 2021-09-02 21:25 |  |  | Versea Diagnostics LLC | Products may have been distributed to customers other than laboratories certified under the Clinical Laboratory Improvement Amendments as stipulated in the Emergency Use Authorization.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188619 | Z-2409-2021 | Healgen COVID-19 IgG/IgM Rapid Test Cassette (Whole Blood/Serum/Plasma) Qualitative detection and differentiation of IgM and IgG antibodies against SARS-CoV-2 in human venous whole blood, plasma (Li+- heparin, K2-EDTA and sodium-citrate), and serum. Intended for use as an aid in identifying individuals with an adaptive immune response to SARS-CoV-2, indicating recent or prior infection. Emergency use of this test is limited to authorized laboratories. Authorized Laboratories: Laboratories certified under the Clinical Laboratory Improvement Amendments of 1988 (CLIA), 42 U.S.C. 263a, to perform moderate or high complexity tests | Healgen COVID-19 IgG/IgM Rapid Test Cassette | 2 | CDRH | 2021-09-02 21:25 |  |  | Versea Diagnostics LLC | Products may have been distributed to customers other than laboratories certified under the Clinical Laboratory Improvement Amendments as stipulated in the Emergency Use Authorization.
http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188617 | Z-2407-2021 | CareStart COVID-19 Antigen Rapid Diagnostic Test for the detection of SARS-CoV-2 Antigen Qualitative detection of the nucleocapsid protein antigen from SARS-CoV-2 in nasopharyngeal or anterior nasal swab specimens directly collected from individuals suspected of COVID-19 by their healthcare provider within five days of symptom onset, or from individuals without symptoms or other epidemiological reasons to suspect COVID-19 when tested twice over two or three days with at least 24 hours and no more than 48 hours between tests. (CLIA Waived Laboratories) | CareStart COVID-19 Antigen Rapid Diagnostic Test | 2 | CDRH | 2021-09-02 21:25 |  |  | Versea Diagnostics LLC | Products may have been distributed to customers other than laboratories certified under the Clinical Laboratory Improvement Amendments as stipulated in the Emergency Use Authorization.