## Sheet1
| Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725524&pc=OSI | 3009402404-2021-00019 | 2021-09-12 | Injury | JOERNS HEALTHCARE | 2021-10-31 | OSI | BARI10A BARIATRIC BED | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725525&pc=OQY | 3017425145-2021-00026 | 2021-05-30 | Injury | ALYDIA HEALTH | 2021-10-31 | OQY | JADA SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K201199 |  | 1 | Event Description: PRIOR TO JADA INSERTION, THE PATIENT WAS TREATED WITH CAROPROST (2 DOSES) AND METHERGINE (1 DOSE). CUMULATIVE QUANTITATIVE BLOOD LOSS AT TIME OF JADA INSERTION WAS 4700 ML. JADA WAS INSERTED IN THE OR FOR "SEVERE UTERINE ATONY."THE PATIENT RECEIVED MISOPROSTOL, ONE DOSE OF TXA, AND RED BLOOD CELLS (UNKNOWN NUMBER OF UNITS) WHILE ON ACTIVE JADA TREATMENT. JADA WAS REMOVED AFTER 0.87 HOURS OF USE. TOTAL BLOOD VOLUME IN THE JADA COLLECTION CANNISTER WAS REPORTED AS 1600 ML. THE DECISION WAS MADE TO REMOVE THE JADA DEVICE AND ESCALATE CARE AS BLEEDING CONTINUED DESPITE JADA TREATMENT. PATIENT'S INTERVENTIONS ESCALATED TO A UTERINE ARTERY LIGATION, COMPRESSION SUTURES, PLATELET TRANSFUSION (UNKNOWN UNITS), FRESH FROZEN PLASMA TRANSFUSION (UNKNOWN UNITS), CRYOPRECIPITATE (UNKNOWN UNITS), OVARIAN CYST REMOVAL, AND HYSTERECTOMY. Manufacturer Narrative: IN THIS CASE, A MASSIVE PPH WAS NOT CONTROLLED BY THE JADA SYSTEM AND ESCALATION OF CARE, ULTIMATELY TO A HYSTERECTOMY, WAS NECESSARY. THE JADA DEVICE WAS NOT UTILIZED UNTIL THE PATIENT HAD ALREADY LOST 4700 ML OF BLOOD BUT THAT DIC WAS NOT DIAGNOSED UNTIL AFTER THE HYSTERECTOMY WHEN THE CUMULATIVE QBL WAS REPORTED AS 8700 ML. THE INFORMATION PROVIDED SUGGESTS THAT THIS WAS A COMPLICATED AND SERIOUS PPH EVENT, AND IT IS POSSIBLE THAT THE PATIENT WAS IN DIC PRIOR TO JADA INITIATION WHICH MAY HAVE CONTRIBUTED TO THE FAILURE OF JADA TO CONTROL HER HEMORRHAGE. ALTHOUGH THE LEVEL OF BLEEDING THAT WAS PRESENT AT THE TIME JADA WAS PLACED WAS SUFFICIENT REASON ALONE TO REQUIRE THE POST-JADA MEDICAL AND SURGICAL INTERVENTIONS, OUT OF AN ABUNDANCE OF CAUTION, WE ARE REPORTING THIS CASE AS AN MDR AS WE CANNOT RULE OUT THAT THE FAILURE OF THE JADA SYSTEM TO CONTROL HER PPH CONTRIBUTED TO THE NEED FOR THE POST-JADA MEDICAL AND SURGICAL INTERVENTIONS. THIS REPORT WILL BE AMENDED IF WE OBTAIN ADDITIONAL INFORMATION REGARDING THIS EVENT. THE SUBMISSION OF THIS REPORT IS NOT AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725526&pc=NAY | 2955842-2021-11400 | 2021-09-28 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | DA VINCI | Material Separation | No Clinical Signs, Symptoms or Conditions | K140553 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725527&pc=NAY | 2955842-2021-11401 | 2021-09-30 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST;DAVINCI SI | Arcing | No Clinical Signs, Symptoms or Conditions | K050369 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725528&pc=NAY | 2955842-2021-11402 | 2021-10-05 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST;DAVINCI SI | Arcing | No Clinical Signs, Symptoms or Conditions | K050369 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725529&pc=NEY | 1717344-2021-01473 | 2021-10-06 | Death | COVIDIEN MFG DC BOULDER | 2021-10-31 | NEY | EMPRINT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K163105 |  | 1 | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVELY, THE PATIENT DEVELOPED A BLEED IN THE LIVER FEW DAYS AFTER THE PROCEDURE AND DIED. THE ABLATION WAS DONE UNDER GENERAL ANESTHESIA USING THE CAGEN1 EMPRINT DEVICE USING THE CA15L2 ANTENNA. THE ABLATION WAS A SINGLE ABLATION, WITH NO OVERLAPPING. THE PHYSICIAN DID NOT TRACK THE PATIENT'S ABLATION WHILE REMOVING THE NEEDLE. THE PHYSICIAN DID USE COMPUTED TOMOGRAPHY (CT) TO PLACE THE NEEDLE AND OBTAINED CT IMAGING THROUGHOUT THE ABLATION, PLUS OBTAINED A CT IMMEDIATELY POST ABLATION. THE PATIENT WAS TRANSFERRED TO POST ANESTHESIA CARE UNIT (PACU) FOR OVERNIGHT OBSERVATION. WHILE IN PACU THE PATIENT BECAME HYPOTENSIVE, WAS TAKEN TO RADIOLOGY AND AN EMBOLIZATION PERFORMED. THE PATIENT WAS LATER TAKEN TO CT AND WHILE IN CT CODED AND PASSED AWAY. THE CUSTOMER REPORTED TO THE ACCOUNT MANAGER THAT THIS DEATH WAS NOT RELATED TO A MALFUNCTION OF THE EMPRINT EQUIPMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725530&pc=NEY | 3004962788-2021-00097 | 2021-10-06 | Death | COVIDIEN LP - SUPERDIMENSION INC | 2021-10-31 | NEY | EMPRINT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K133821 |  | 1 | Event Description: ACCORDING TO THE REPORTER, POSTOPERATIVELY, THE PATIENT DEVELOPED A BLEED IN THE LIVER FEW DAYS AFTER THE PROCEDURE AND DIED. THE ABLATION WAS DONE UNDER GENERAL ANESTHESIA USING THE CAGEN1 EMPRINT DEVICE USING THE CA15L2 ANTENNA. THE ABLATION WAS A SINGLE ABLATION, WITH NO OVERLAPPING. THE PHYSICIAN DID NOT TRACK THE PATIENT'S ABLATION WHILE REMOVING THE NEEDLE. THE PHYSICIAN DID USE COMPUTED TOMOGRAPHY (CT) TO PLACE THE NEEDLE AND OBTAINED CT IMAGING THROUGHOUT THE ABLATION, PLUS OBTAINED A CT IMMEDIATELY POST ABLATION. THE PATIENT WAS TRANSFERRED TO POST ANESTHESIA CARE UNIT (PACU) FOR OVERNIGHT OBSERVATION. WHILE IN PACU THE PATIENT BECAME HYPOTENSIVE, WAS TAKEN TO RADIOLOGY AND AN EMBOLIZATION PERFORMED. THE PATIENT WAS LATER TAKEN TO CT AND WHILE IN CT CODED AND PASSED AWAY. THE CUSTOMER REPORTED TO THE ACCOUNT MANAGER THAT THIS DEATH WAS NOT RELATED TO A MALFUNCTION OF THE EMPRINT EQUIPMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725531&pc=NAY | 2955842-2021-11403 | 2021-09-14 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725532&pc=NAY | 2955842-2021-11404 | 2021-09-24 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725533&pc=GDW | 2955842-2021-11405 | 2021-10-05 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | GDW | SUREFORM | Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K173721 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725534&pc=DZE | 3003184527-2021-01532 | 2021-09-29 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT MC INICELL, PF 6.0, E 6.0, L11.0 | Failure to Osseointegrate | Inadequate Osseointegration | K093615 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED - IMPLANT WAS PLACED RIGHT AFTER EXPLANTATION OF AN OLDER IMPLANT - IMPLANT DIAMETER WAS TOO BIG.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725535&pc=DZE | 3003184527-2021-01533 | 2021-09-29 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 3.5, E 3.5, L12.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725536&pc=NAY | 2955842-2021-11406 | 2021-08-13 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST;DAVINCI SI | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K050369 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725537&pc=DZE | 3003184527-2021-01534 | 2021-10-01 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 5.0, E 5.0, L8.0 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, LOCAL INFECTION / SUBACUTE CHRONIC OSTEITIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725538&pc=DZE | 3003184527-2021-01535 | 2021-10-04 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, SMOKER, IMMEDIATE IMPLANTATION, PRECEDING/SIMULTANEOUS BONE AUGMENTATION, PAIN, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725539&pc=DZE | 3003184527-2021-01536 | 2021-10-05 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725540&pc=DZE | 3003184527-2021-01537 | 2021-10-06 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, SMOKER, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725541&pc=DZE | 3003184527-2021-01538 | 2021-10-08 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT RC INICELL, PF 5.0, E 5.0, L8.0 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, SMOKER, INADEQUATE BONE QUALITY/QUANTITY, PAIN, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725542&pc=DZE | 3003184527-2021-01539 | 2021-10-14 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT MC INICELL, PF 4.5, E 4.2, L8.0 | Failure to Osseointegrate | Inadequate Osseointegration | K093615 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED..
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725543&pc=DZE | 3003184527-2021-01540 | 2021-10-03 | Injury | THOMMEN MEDICAL AG | 2021-10-31 | DZE | ELEMENT MC INICELL, PF 4.5, E 4.2, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K093615 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, PERI-IMPLANTITIS, PAIN, SWELLING, INFLAMMATION, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725544&pc=OTN | 2210968-2021-09557 |  | Injury | ETHICON INC. | 2021-10-31 | OTN | TVT EXACT RETROPUBIC SYSTEM | Migration | Erosion; Incontinence; Pain; Urinary Retention | K132054 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725545&pc=DSI | 9610816-2021-10463 | 2021-10-20 | Malfunction | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2021-10-31 | DSI | INTELLIVUE MULTI MEASUREMENT SERVER X2 | Unable to Obtain Readings; No Audible Prompt/Feedback | No Clinical Signs, Symptoms or Conditions | K150310 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS NON-INVASIVE BLOOD PRESSURE (NIBP) ERROR ALONG WITH A DEFECTIVE LOUDSPEAKER. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS NO REPORT OF PATIENT OR USER HARM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725546&pc=FTL | 2210968-2021-10627 |  | Injury | ETHICON INC. | 2021-10-31 | FTL | PROLENE POLYPROPYLENE MESH UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K962530 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725547&pc=LIT | 2134265-2021-13674 | 2021-10-20 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2021-10-31 | LIT | MUSTANG | Material Rupture | No Clinical Signs, Symptoms or Conditions | K103751 |  | 1 | Event Description: IT WAS REPORTED THAT BALLOON RUPTURE OCCURRED. THE 99% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED SUPERFICIAL FEMORAL ARTERY. A 6.0 X 150, 135CM MUSTANG BALLOON CATHETER WAS ADVANCED FOR DILATATION. HOWEVER, DURING INITIAL INFLATION AT 20 ATMOSPHERES FOR 5 SECONDS, THE BALLOON RUPTURED. THE DEVICE WAS REMOVED AND THE PROCEDURE WAS COMPLETED WITH A DIFFERENT DEVICE. THERE WERE NO COMPLICATIONS NOR INJURIES REPORTED AND THE PATIENT WAS IN GOOD CONDITION. Manufacturer Narrative: INITIAL REPORTER CITY: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725549&pc=FGE | 2134265-2021-13684 | 2021-10-21 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2021-10-31 | FGE | MUSTANG | Material Rupture | No Clinical Signs, Symptoms or Conditions | K141521 |  | 1 | Event Description: IT WAS REPORTED THAT BALLOON RUPTURE OCCURRED. VASCULAR ACCESS WAS OBTAINED VIA AN IPSILATERAL RETROGRADE APPROACH. THE 90% STENOSED TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED COMMON ILIAC ARTERY. AFTER A JUPITER FC GUIDEWIRE CROSS THE LESION, A 7.0 X 20, 75CM MUSTANG BALLOON CATHETER WAS ADVANCED FOR DILATATION. HOWEVER, DURING INITIAL INFLATION AT 8 ATMOSPHERES FOR 10 SECONDS, THE BALLOON RUPTURED. THE DEVICE WAS REMOVED AND THE PROCEDURE WAS COMPLETED WITH A DIFFERENT DEVICE. THERE WERE NO PATIENT COMPLICATIONS REPORTED AND THE PATIENT WAS IN GOOD CONDITION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725550&pc=NIQ | 2134265-2021-13686 | 2021-10-19 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2021-10-31 | NIQ | PROMUS PREMIER SELECT | Failure to Advance; Material Deformation | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT STENT DAMAGE OCCURRED. VASCULAR ACCESS WAS OBTAINED VIA FEMORAL ARTERY. THE 80% STENOSED TARGET LESION WITH A BEND OF >45 AND <90 DEGREES WAS LOCATED IN THE MILDLY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY. AFTER THE LESION WAS PRE-DILATED WITH A 2.00X8.00 NC BALLOON CATHETER, A 20 X 2.75 PROMUS PREMIER SELECT DRUG-ELUTING STENT WAS ADVANCED BUT FAILED TO TRACK THE LESION DUE TO SIGNIFICANT RESISTANCE. THE LESION WAS PRE-DILATED AGAIN WITH A 2.50X8.00 NC BALLOON BUT STILL IT FAILED TO CROSS THE LESION AND THE STENT FOUND DAMAGE. THE PROCEDURE WAS COMPLETED WITH A DIFFERENT DEVICE. THERE WERE NO PATIENT COMPLICATIONS REPORTED AND THE PATIENT STATUS WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725551&pc=NIQ | 2134265-2021-13685 | 2021-10-18 | Malfunction | BOSTON SCIENTIFIC CORPORATION | 2021-10-31 | NIQ | PROMUS PREMIER SELECT | Failure to Advance; Material Deformation | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT STENT DAMAGE OCCURRED. VASCULAR ACCESS WAS OBTAINED VIA FEMORAL ARTERY. THE 80% STENOSED, 35MMX3.5MM, CONCENTRIC, DE NOVO TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND MODERATELY CALCIFIED RIGHT CORONARY ARTERY. A 38 X 3.50 PROMUS PREMIER SELECT DRUG-ELUTING STENT WAS ADVANCED BUT FAILED TO CROSS THE LESION. WHEN REMOVED BACK, THE STENT WAS FOUND DAMAGE. THE PROCEDURE WAS COMPLETED WITH A DIFFERENT DEVICE. THERE WERE NO PATIENT COMPLICATIONS NOR INJURIES REPORTED AND THE PATIENT STATUS WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725552&pc=GDW | 2647580-2021-03676 | 2021-10-03 | Malfunction | US SURGICAL PUERTO RICO | 2021-10-31 | GDW | ENDO GIA ULTRA | Failure to Fire | No Clinical Signs, Symptoms or Conditions | K111825 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC SURGERY, THE GREEN BUTTON WAS PRESSED BUT THE DEVICE WAS UNABLE TO FIRE. ANOTHER HANDLE WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725553&pc=FZP | 2647580-2021-03677 | 2021-10-05 | Malfunction | US SURGICAL PUERTO RICO | 2021-10-31 | FZP | PREMIUM SURGICLIP II | Mechanics Altered | No Clinical Signs, Symptoms or Conditions | K143644 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THERE WHERE ISSUES EXPERIENCED WITH THE LOADING OR FIRING OF THE CLIPS. WHEN USING THE DEVICE, THE CLIPS GET STUCK IN FIRING. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725554&pc=OLW | 2936999-2021-01057 |  | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | OLW | BIS | Temperature Problem | No Clinical Signs, Symptoms or Conditions | K072286 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE DEVICE SMELLS CHARRED AND HAD LIQUID DAMAGE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725555&pc=FJS | 3009211636-2021-00303 |  | Malfunction | COVIDIEN MFG SOLUTIONS S.A. | 2021-10-31 | FJS | DIALYSIS UNKNOWN | Migration or Expulsion of Device | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE LINE FELL OUT FOLLOWING INSERTION (WITHIN SEVEN TO EIGHT). IT WAS STATED THAT CUFF PREPARATION WAS STILL BEING PERFORMED, NO NEW INSERTED AND SUTURES ARE STAYING IN FOR UP TO SIX TO EIGHT WEEKS POST LINE INSERTION. THE CATHETER WAS NOT REPAIRED, THERE WAS NO LUER ADAPTER ISSUE AND A NEW LINE WAS INSERTED. THERE WAS NO REPORTED PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725556&pc=MUD | 2936999-2021-01058 | 2021-10-08 | Malfunction | AVNET INC | 2021-10-31 | MUD | ADULT SENSOR PMSENS71-A INVOS KIT | Low Readings | No Clinical Signs, Symptoms or Conditions | K182868 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE SENSOR HAD LOW RSO2 READINGS. BOTH SENSORS WERE REPLACED WITH SAME BATCH SENSORS AND ISSUE WAS RESOLVED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725557&pc=BTR | 2936999-2021-01059 | 2021-10-08 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | BTR | TAPERGUARD | Use of Device Problem | No Clinical Signs, Symptoms or Conditions | K090352 |  | 1 | Event Description: ACCORDING TO THE REPORTER, PRIOR TO USE, THE CUFF HAD AN INFLATION AND DEFLATION ISSUE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725558&pc=DYB | 3008452825-2021-00497 | 2021-10-08 | Malfunction | ST. JUDE MEDICAL | 2021-10-31 | DYB | AGILIS NXT STEERABLE INTRODUCER DUAL-REACH 71 CM SHEATH, LARGE CURL, 8.5F | Fluid Leak | No Clinical Signs, Symptoms or Conditions | K081645 |  | 1 | Event Description: DURING AN ATRIAL FIBRILLATION PROCEDURE, A FLUID AND AIR LEAK OCCURRED. AFTER MAKING PATIENT CONTACT, A FLUID AND AIR LEAK WAS NOTED AT THE HEMOSTASIS VALVE OF THE SHEATH. THE SHEATH WAS EXCHANGED AND THE PROCEDURE WAS COMPLETED WITH NO ADVERSE CONSEQUENCES TO THE PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725559&pc=QBJ | 3004753838-2021-252749 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND A LOSS OF CONNECTION WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725560&pc=DYB | 3005334138-2021-00634 | 2021-10-06 | Malfunction | ST. JUDE MEDICAL | 2021-10-31 | DYB | SWARTZ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL0, 63 CM LENGTH, 8F | Fluid Leak | No Clinical Signs, Symptoms or Conditions | K052644 |  | 1 | Event Description: DURING AN ATRIAL FIBRILLATION PROCEDURE, A BLOOD LEAK OCCURRED. AFTER PLACING THE SHEATH, BLOOD BEGAN TO LEAK FROM THE HEMOSTASIS VALVE OF THE SHEATH. THE SHEATH WAS REPLACED AND THE PROCEDURE WAS COMPLETED WITH NO ADVERSE CONSEQUENCES TO THE PATIENT. NO CATHETER WAS INSERTED INTO THE SHEATH. NO ADDITIONAL FEMORAL VEIN PUNCTURE WAS REQUIRED FOR REPLACING THE SHEATH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725561&pc=QBJ | 3004753838-2021-252750 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725562&pc=GDW | 1219930-2021-04539 | 2021-09-06 | Malfunction | COVIDIEN LP LLC NORTH HAVEN | 2021-10-31 | GDW | SIGNIA | Unraveled Material | No Clinical Signs, Symptoms or Conditions | K160176 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING A LAPAROSCOPIC PROCEDURE, BEFORE STARTING THE INTERVENTION, WHEN THE SCRUB NURSE HAS CLOSED THE SHELL A THERE HAS A SPACE BETWEEN THE TWO (2) PARTS OF THE DEVICE. THERE WAS NO STERILITY. TO RESOLVE THE ISSUE, A NEW SHELL WAS USED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725563&pc=GDW | 2647580-2021-03678 | 2021-10-11 | Malfunction | US SURGICAL PUERTO RICO | 2021-10-31 | GDW | ENDO GIA | Failure to Cycle; Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K111825 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC PROCEDURE, AFTER THE INITIAL FIRING, TISSUE BUILD UP AND WAS PUSHED TOWARDS THE TIP OF THE RELOAD AND NOT ALL STAPLES MANAGED TO BE INSERTED. THE STAPLE LINE WAS INCOMPLETE CENTRALLY, AND WAS FIXED BY OVERLAPPING ANOTHER RELOAD. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725564&pc=QBJ | 3004753838-2021-252751 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE TEST WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WITH (B)(4) BLUETOOTH DEVICE WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725565&pc=HQC | 2028159-2021-01256 | 2021-09-30 | Injury | ALCON RESEARCH, LLC - IRVINE TECHNOLOGY CENTER | 2021-10-31 | HQC | CENTURION VISION SYSTEM | Adverse Event Without Identified Device or Use Problem | Uveitis; Toxic Anterior Segment Syndrome (TASS) | K121555 |  | 1 | Event Description: A HEALTH PROFESSIONAL REPORTED THAT A PATIENT EXPERIENCED TOXIC ANTERIOR SEGMENT SYNDROME (TASS). PROCEDURE DETAILS AND PATIENT IMPACT HAVE NOT BEEN PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. ADDITIONAL INFORMATION RECEIVED INDICATED THAT AFTER A CATARACT SURGERY, THE PATIENT EXPERIENCED TASS IN THE RIGHT EYE. THE PATIENT HAD NO SURGICAL COMPLICATIONS. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. THIS IS FIRST OF THREE REPORTS FOR THIS REPORTED EVENT. THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725566&pc=GDW | 2647580-2021-03679 | 2021-10-11 | Malfunction | US SURGICAL PUERTO RICO | 2021-10-31 | GDW | ENDO GIA | Failure to Cycle; Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K111825 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC BARIATRIC PROCEDURE, AFTER THE INITIAL FIRING, TISSUE BUILD UP AND WAS PUSHED TOWARDS THE TIP OF THE RELOAD AND NOT ALL STAPLES MANAGED TO BE INSERTED. THE STAPLE LINE WAS INCOMPLETE CENTRALLY, AND WAS FIXED BY OVERLAPPING ANOTHER RELOAD. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725567&pc=QBJ | 3004753838-2021-252752 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725568&pc=BTR | 2936999-2021-01060 | 2021-10-11 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | BTR | SHILEY | Use of Device Problem | No Clinical Signs, Symptoms or Conditions | K090352 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE PATIENT HAD BEEN INTUBATED FOR FOUR DAYS THEN A CUFF INFLATION AND DEFLATION ISSUE HAPPENED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725569&pc=QBJ | 3004753838-2021-252753 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE TESTING WAS PERFORMED AND FAILED. NO DATA WAS PROVIDED FOR EVALUATION. THE ALLEGATION WAS UNDETERMINED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725570&pc=GDW | 1219930-2021-04540 | 2021-10-11 | Malfunction | COVIDIEN LP LLC NORTH HAVEN | 2021-10-31 | GDW | EEA ORVIL | Loss of or Failure to Bond | No Clinical Signs, Symptoms or Conditions | K093402 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN INTRATHORACIC ESOPHAGECTOMY, THE ORAL ANVIL COULD NOT BE CONNECTED. THE ANVIL WAS REMOVED AND ANASTOMOSIS WAS PERFORMED BY HAND INSTEAD. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725571&pc=QBJ | 3004753838-2021-252754 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725572&pc=NVZ | 2017865-2021-35515 | 2021-10-09 | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | NVZ | ASSURITY RF DR | Premature Discharge of Battery; Capturing Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED REMOTELY VIA MERLIN.NET. UPON REVIEW, THE PACEMAKER (PM) EXHIBITED POSSIBLE PREMATURE BATTERY DEPLETION AS WELL AS HIGH CAPTURE THRESHOLD. THE PM WAS EXPLANTED AND REPLACED. THE PATIENT WAS DISCHARGED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725573&pc=NVY | 2017865-2021-35516 | 2021-10-11 | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Capturing Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR GENERATOR CHANGE PROCEDURE. UPON INTERROGATION PRIOR TO PROCEDURE, IT WAS FOUND THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INCREASED CAPTURE THRESHOLD. THE RV LEAD WAS CAPPED AND REPLACED ON (B)(6) OCT 2021. PATIENT WAS DISCHARGED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725574&pc=NIK | 2017865-2021-35520 |  | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | NIK | QUARTET LEAD SMALL-S, 86 CM | Capturing Problem; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR FOLLOW-UP. UPON INTERROGATION, IT WAS FOUND THAT THE LEFT VENTRICULAR (LV) LEAD EXHIBITED INCREASED THRESHOLD. CHEST X-RAY WAS PERFORMED TO FURTHER CONFIRM THAT THE LV LEAD HAD DISLODGED. THE LV LEAD WAS EXPLANTED AND REPLACED. PATIENT CONDITION WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725575&pc=QBJ | 3004753838-2021-252755 | 2021-09-27 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING WITH (B)(6) BLUETOOTH DEVICE WAS PERFORMED AND PASSED. A REVIEW OF THE BIN FILE WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725576&pc=FFT | 9710107-2021-00453 | 2021-10-05 | Malfunction | GIVEN IMAGING LTD., YOQNEAM | 2021-10-31 | FFT | BRAVO | Positioning Failure | Airway Obstruction | K102543 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THEY HAD A CAPSULE WHICH FAILED TO ATTACH TO THE PATIENT¿S ESOPHAGUS AND DROPPED IN THE PATIENT¿S MOUTH. THE PHYSICIAN REMOVED THE CAPSULE FROM THE PATIENT'S MOUTH WITH FINGERS. A REPEAT PROCEDURE WAS DONE DURING THE SAME EVENT AND IT WAS SUCCESSFUL. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725577&pc=GEI | 1717344-2021-01474 | 2021-10-08 | Malfunction | COVIDIEN MFG DC BOULDER | 2021-10-31 | GEI | LIGASURE | Difficult to Open or Close | No Clinical Signs, Symptoms or Conditions | K170869 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, THE DEVICE'S JAW WAS DIFFICULT AND IMPOSSIBLE TO OPEN. THE JAW WAS PARTIALLY OPENED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725578&pc=NVZ | 2017865-2021-35522 | 2018-08-27 | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | NVZ | CURRENT PLUS VR, DF-4 CONNECTOR | Premature Discharge of Battery | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR FOLLOW-UP. UPON INTERROGATION, IT WAS FOUND THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) EXHIBITED PREMATURE BATTERY DEPLETION. THE ICD WAS EXPLANTED AND REPLACED. PATIENT CONDITION WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725579&pc=QBJ | 3004753838-2021-252756 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725581&pc=MXC | 2017865-2021-35527 | 2021-10-04 | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | MXC | CONFIRM | Adverse Event Without Identified Device or Use Problem | Discomfort | K163407 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC WITH DISCOMFORT AT THE SITE WHERE THE INSERTABLE CARDIAC MONITOR (ICM) WAS INSTALLED. THE ICM WAS EXPLANTED. PATIENT CONDITION WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725582&pc=JOH | 2936999-2021-01061 |  | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Use of Device Problem | Unintended Extubation | K182861 |  | 1 | Event Description: ACCORDING TO THE REPORTER, AFTER TWO DAYS OF USE, THE DEVICE'S FLANGE WAS DAMAGED AND FELL OUT OF THE PATIENT. REPLACEMENT OF THE TRACH WAS DONE. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725583&pc=QBJ | 3004753838-2021-252757 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725584&pc=JOH | 2936999-2021-01062 |  | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Use of Device Problem | Unintended Extubation | K182861 |  | 1 | Event Description: ACCORDING TO THE REPORTER, AFTER TWO DAYS OF USE, THE DEVICE'S FLANGE WAS DAMAGED AND FELL OUT OF THE PATIENT. REPLACEMENT OF THE TRACH WAS DONE. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725585&pc=QBJ | 3004753838-2021-252758 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. NO DATA WAS PROVIDED FOR EVALUATION. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725586&pc=JOH | 2936999-2021-01063 |  | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Use of Device Problem | Unintended Extubation | K182861 |  | 1 | Event Description: ACCORDING TO THE REPORTER, AFTER TWO DAYS OF USE, THE DEVICE'S FLANGE WAS DAMAGED AND FELL OUT OF THE PATIENT. REPLACEMENT OF THE TRACH WAS DONE. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725587&pc=JOH | 2936999-2021-01064 | 2021-10-07 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Labelling, Instructions for Use or Training Problem | Insufficient Information | K142296 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE DEVICE LABEL OF THE FLANGE WAS WEARING OFF. THERE WAS NO REPORTED PATIENT OUTCOME. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725588&pc=QBJ | 3004753838-2021-252759 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WAS PERFORMED AND PASSED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725589&pc=JOH | 2936999-2021-01065 | 2021-10-14 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Disconnection; Use of Device Problem | Unintended Extubation | K945513 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE FLANGE OF THE DEVICE HAD DAMAGED AND FELL OUT OF THE PATIENT. THE PATIENT HAD A REPLACEMENT OF THE TUBE. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725590&pc=QBJ | 3004753838-2021-252760 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND A LOSS OF CONNECTION WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725591&pc=JOH | 2936999-2021-01066 | 2021-10-14 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | JOH | SHILEY | Disconnection; Use of Device Problem | Unintended Extubation | K945513 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING USE, THE FLANGE OF THE DEVICE HAD DAMAGED AND FELL OUT OF THE PATIENT. THE PATIENT HAD A REPLACEMENT OF THE TUBE. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725592&pc=QBJ | 3004753838-2021-252761 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725593&pc=QBJ | 3004753838-2021-252762 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERIOR/EXTERIOR VISUAL INSPECTION WAS PERFORMED AND PASSED. RECEIVER CHARGED AND WILL BOOT WAS PERFORMED AND PASSED. RECEIVER DOWNLOAD RETRIEVED AND ATTACHED WAS PERFORMED AND PASSED. THE RECEIVER LOG WAS DOWNLOADED AND REVIEWED FINDING NO ERRORS RELATED TO THE COMPLAINT. CONFIRMATION OF THE ALLEGATION COULD NOT BE DETERMINED. A PROBABLE CAUSE COULD NOT BE DETERMINED AS THE ALLEGATION WAS NOT FOUND. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725594&pc=MKJ | 3030677-2021-15279 |  | Malfunction | PHILIPS NORTH AMERICA LLC | 2021-10-31 | MKJ | FRX DEFIBRILLATOR, FRENCH, EXCHANGE | Failure of Device to Self-Test | No Clinical Signs, Symptoms or Conditions | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE IS FAILING SELF-TEST.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725595&pc=BTR | 2936999-2021-01067 | 2021-09-15 | Malfunction | MMJ SA DE CV (USD) | 2021-10-31 | BTR | SHILEY | Connection Problem; Gas Leak | Insufficient Information | K090352 |  | 1 | Event Description: ACCORDING TO THE REPORTER, THE EVENT WAS FOUND EIGHT DAYS AFTER INTUBATION. AN AIR LEAKAGE FROM THE CUFF HAD OCCURRED. IT WAS FOUND THAT THERE WAS A TRACE OF DAMAGE TO THE 15 MM CONNECTOR (CRACKS). THERE WAS NO PATIENT OUTCOME. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725596&pc=QBJ | 3004753838-2021-252763 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE TEST WAS PERFORMED AND PASSED. NORDIC BLUETOOTH PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS NOT FOUND. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725597&pc=HQC | 2028159-2021-01258 | 2021-09-28 | Malfunction | ALCON RESEARCH, LLC - IRVINE TECHNOLOGY CENTER | 2021-10-31 | HQC | CENTURION VISION SYSTEM (ACTIVE SENTRY) | Inability to Irrigate; Suction Problem; Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | K161794 |  | 1 | Event Description: A CUSTOMER REPORTED THAT DURING A CATARACT SURGERY, THE EQUIPMENT WAS NOT PERFORMING THE ASPIRATION AND IRRIGATION PROPERLY IN PHACOEMULSIFICATION MODE AND ALSO PRESENTED WITH INTERMITTENT PHACOEMULSIFICATION POWER. THE ASPIRATION RATE AND VACUUM WERE FOUND TO BE 30 AND 500 RESPECTIVELY. NO SYSTEM MESSAGE WAS DISPLAYED. THE PROCEDURE COMPLETED IN A SLOWER MANNER. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. THIS IS THE FIRST OF TWO REPORTS FOR THIS REPORTED EVENTS. THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725598&pc=QBJ | 3004753838-2021-252764 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE TESTING WAS PERFORMED AND FAILED. DATA WAS PROVIDED FOR EVALUATION. PER (B)(4), CASES WHERE THE SENSOR SESSION LINE IS MISSING IN SHARE SUPPORT TOOL AND SHARE SUPPORT SERVICE, COMPLAINTS WILL BE CLOSED SINCE A DATA INVESTIGATION CANNOT BE COMPLETED AS DATA AMBIGUITY CANNOT BE RESOLVED FURTHER. THE ALLEGATION WAS UNDETERMINED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725599&pc=GAM | 1219930-2021-04541 | 2021-10-13 | Malfunction | COVIDIEN LP LLC NORTH HAVEN | 2021-10-31 | GAM | POLYSORB | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K963253 |  | 1 | Event Description: ACCORDING TO THE REPORTER, PRIOR TO A PROCEDURE, WHEN THE DEVICE WAS OPENED, THE NEEDLE WAS NOT ATTACHED. ANOTHER SUTURE WAS OPENED TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725600&pc=HQC | 2028159-2021-01259 | 2021-09-28 | Malfunction | ALCON RESEARCH, LLC - IRVINE TECHNOLOGY CENTER | 2021-10-31 | HQC | CENTURION VISION SYSTEM, ACCESSORY, ACTIVE SENTRY HANDPIECE | Inability to Irrigate; Suction Problem; Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | K161794 |  | 1 | Event Description: A CUSTOMER REPORTED THAT DURING A CATARACT SURGERY, THE EQUIPMENT WAS NOT PERFORMING THE ASPIRATION AND IRRIGATION PROPERLY IN PHACOEMULSIFICATION MODE AND ALSO PRESENTED WITH INTERMITTENT PHACOEMULSIFICATION POWER. THE ASPIRATION RATE AND VACUUM WERE FOUND TO BE 30 AND 500 RESPECTIVELY. NO SYSTEM MESSAGE WAS DISPLAYED. THE PROCEDURE COMPLETED IN A SLOWER MANNER. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. THIS IS SECOND OF TWO REPORTS FOR THIS REPORTED EVENT. THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725601&pc=GCJ | 2647580-2021-03680 | 2021-10-13 | Malfunction | US SURGICAL PUERTO RICO | 2021-10-31 | GCJ | PDB | Inflation Problem | No Clinical Signs, Symptoms or Conditions | K935426 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING LAPAROSCOPIC INGUINAL HERNIA REPAIR, AT VISUAL DEPLOYMENT, THERE WAS A SENSE OF DISCOMFORT IN THE SPHERICAL SHAPE WHEN THE BALLOON WAS DEPLOYED AND BEING INFLATED. A NEW PRODUCT WAS USED TO RESOLVE THE ISSUE IN ORDER TO COMPLETE THE CASE. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725602&pc=QBJ | 3004753838-2021-252765 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE TEST WAS PERFORMED AND PASSED. (B)(4) BLUETOOTH PAIRING TEST WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725603&pc=QBJ | 3004753838-2021-252766 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL/EXTERIOR VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. NO DATA WAS PROVIDED FOR EVALUATION. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725604&pc=QBJ | 3004753838-2021-252767 | 2021-10-03 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND A LOSS OF CONNECTION WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725605&pc=LFL | 1717344-2021-01475 | 2021-10-09 | Malfunction | COVIDIEN MFG DC BOULDER | 2021-10-31 | LFL | SONICISION | Failure to Fire; Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K101797 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER THE DISSECTOR TIP WAS ASSEMBLED, THE RED WARNING LIGHT WAS ON DURING THE TRIAL FIRING PROCESS. A PART OF THE DEVICE WAS DETACHED AND NOTHING FELL ON THE PATIENT'S CAVITY. AFTER REPLACING THE NEW DISSECTOR TIP, IT COULD BE USED NORMALLY. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725606&pc=QBJ | 3004753838-2021-252768 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING WITH (B)(4) BLUETOOTH DEVICE WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725607&pc=QBJ | 3004753838-2021-252769 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725608&pc=QBJ | 3004753838-2021-252770 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725609&pc=HQC | 2028159-2021-01260 | 2021-10-01 | Malfunction | ALCON RESEARCH, LLC - IRVINE TECHNOLOGY CENTER | 2021-10-31 | HQC | CENTURION VISION SYSTEM (ACTIVE SENTRY) | Decrease in Suction | No Clinical Signs, Symptoms or Conditions | K161794 |  | 1 | Event Description: A PHYSICIAN REPORTED THAT DURING A SURGERY, AN OPHTHALMIC OPERATING CONSOLE PRESENTED WITH POOR ASPIRATION IN PHACOEMULSIFICATION AND IRRIGATION/ASPIRATION MODES. THE SURGERY WAS COMPLETED WITH THE SAME SYSTEM. NO PATIENT HARM WAS REPORTED. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725610&pc=QBJ | 3004753838-2021-252771 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. DATA WAS RECEIVED BUT DOES NOT REFLECT FULL INVESTIGATION WINDOW. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725611&pc=QBJ | 3004753838-2021-252772 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725612&pc=QBJ | 3004753838-2021-252773 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725613&pc=QBJ | 3004753838-2021-252774 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING TEST WAS PERFORMED AND FAILED. NO DATA WAS PROVIDED FOR EVALUATION. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725614&pc=FDF | 9610877-2021-01259 |  | Malfunction | HOYA CORPORATION PENTAX TOKYO OFFICE | 2021-10-31 | FDF | PENTAX | No Display/Image; Display or Visual Feedback Problem | Insufficient Information | K131855 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725615&pc=MKJ | 3030677-2021-15319 |  | Malfunction | PHILIPS NORTH AMERICA LLC | 2021-10-31 | MKJ | FR3 DEFIB, ECG, JAPANESE, EXCHANGE | Failure of Device to Self-Test | No Clinical Signs, Symptoms or Conditions | P160028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE IS FAILING SELF-TEST.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725616&pc=MKJ | 3030677-2021-15320 |  | Malfunction | PHILIPS NORTH AMERICA LLC | 2021-10-31 | MKJ | FR3 DEFIB, ECG, JAPANESE, EXCHANGE | Inappropriate Audible Prompt/Feedback; Failure of Device to Self-Test | No Clinical Signs, Symptoms or Conditions | P160028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE IS FAILING SELF-TEST AND THAT THE DEVICE SPEAKERS ARE NOT FUNCTIONING CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725617&pc=QBJ | 3004753838-2021-252775 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. IT WAS DETERMINED THAT THE SIGNAL LOSS WAS RELATED TO THE MOBILE APPLICATION. IT WAS INDICATED THAT THE OPERATING SYSTEM FOR THE SMART DEVICE WAS NOT A SUPPORTED SYSTEM, WHICH IS OFF-LABEL USAGE OF THE DEVICE. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE PATIENT CLOSED THE MOBILE APP. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725618&pc=QBJ | 3004753838-2021-252776 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE MEASUREMENT WAS PERFORMED AND PASSED. PAIRING WITH NORDIC BLUETOOTH DEVICE/DOWNLOAD WAS PERFORMED AND PASSED. A REVIEW OF THE BIN FILE WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. HOWEVER, THE DEVICE OPERATED WITHIN SPECIFICATION. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725619&pc=QBJ | 3004753838-2021-252777 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725620&pc=DYB | 3007284313-2021-01637 | 2021-10-01 | Injury | W. L. GORE & ASSOCIATES, INC. | 2021-10-31 | DYB | GORE DRYSEAL FLEX INTRODUCER SHEATH | Patient-Device Incompatibility | Rupture | K160254 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725621&pc=QBJ | 3004753838-2021-252778 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE MEASUREMENT WAS PERFORMED AND PASSED. PAIRING WITH (B)(4) BLUETOOTH DEVICE/DOWNLOAD WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725622&pc=QBJ | 3004753838-2021-252779 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE TEST WAS PERFORMED AND PASSED. NORDIC BLUETOOTH PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS NOT FOUND. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725623&pc=QBJ | 3004753838-2021-252780 | 2021-09-30 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND A LOSS OF CONNECTION WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725624&pc=QBJ | 3004753838-2021-252781 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725625&pc=QBJ | 3004753838-2021-252782 | 2021-10-06 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING TEST/DOWNLOAD WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOG WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS A DEFECTIVE TRANSMITTER. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725626&pc=QBJ | 3004753838-2021-252783 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725627&pc=MIH | 2017233-2021-02504 | 2021-09-05 | Injury | W. L. GORE & ASSOCIATES, INC. | 2021-10-31 | MIH | GORE TAG CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE THROUGH THE SSB11-02 PIVOTAL CLINICAL STUDY. ON (B)(6) 2021, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF A THORACIC AORTIC DISSECTION, ZONE 2, USING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFTS WITH ACTIVE CONTROL SYSTEM. ON (B)(6) 2021, IT WAS REPORTED TYPE B AORTIC DISSECTION. REPORTEDLY RELATED TO DEVICE AND ENDOVASCULAR PROCEDURE. ON (B)(6) 2021, IT WAS RESOLVED WITHOUT SEQUELAE. TREATMENT DETAILS WERE NOT REPORTED. Manufacturer Narrative: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICE REMAINS IMPLANTED. RETURN NOT POSSIBLE. A SEVERE INJURY, ILLNESS OR IMPAIRMENT WHICH REQUIRES HOSPITALIZATION OR MEDICAL INTERVENTION. AN ADVERSE EVENT (E.G. PATIENT HARM) APPEARS TO HAVE OCCURRED, BUT THERE DOES NOT APPEAR TO HAVE BEEN A PROBLEM WITH THE DEVICE OR THE WAY IT WAS USED. THE INVESTIGATION INVOLVED THE ANALYSIS OF RELEVANT PRODUCTION RECORDS IN VIEW OF SUPPORTING THE IDENTIFICATION OF POSSIBLE CAUSES FOR THE ADVERSE EVENT. THE ACTUAL DEVICE INVOLVED IN THE ADVERSE EVENT WAS NOT RETURNED FOR TESTING DESPITE REQUESTS BY MANUFACTURER. MEDICAL DEVICE REMAINS IMPLANTED. THE DEVICE EITHER FUNCTIONED AS INTENDED OR A PROBLEM WAS NOT FOUND. THE INVESTIGATION FINDINGS DO NOT LEAD TO A CLEAR CONCLUSION ABOUT THE CAUSE OF THE REPORTED ADVERSE EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725628&pc=QBJ | 3004753838-2021-252784 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725629&pc=CCK | 3012307300-2021-10484 | 2021-10-04 | Malfunction | NULL | 2021-10-31 | CCK | ANALYZER, GAS, CARBON-DIOXIDE, GASEOUS-PHASE | Device Damaged Prior to Use | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT UPON OPENING THE BOX, 3 PACKAGES OF SODASORB WERE BUSTED OPEN. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725630&pc=NEW | 2210968-2021-10628 | 2020-01-01 | Injury | ETHICON INC. | 2021-10-31 | NEW | PDS II POLYDIOXANONE SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Necrosis | N18331 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725631&pc=JAY | 3012307300-2021-10485 | 2021-10-05 | Malfunction | SMITHS MEDICAL INTERNATIONAL, LTD. | 2021-10-31 | JAY | PORTEX | Break; Material Disintegration | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE VELCRO COMPLETELY CAME OFF ONE SIDE OF THE TRACH TAPES. NO PATIENT INVOLVEMENT
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725632&pc=QBJ | 3004753838-2021-252785 | 2021-09-29 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE TEST WAS PERFORMED AND FAILED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. HOWEVER, THE DEVICE OPERATED WITHIN SPECIFICATION. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725634&pc=QBJ | 3004753838-2021-252786 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE MEASUREMENT WAS PERFORMED AND PASSED. PAIRING WITH (B)(4) BLUETOOTH DEVICE/DOWNLOAD WAS PERFORMED AND PASSED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725635&pc=FDS | 9610877-2021-01275 |  | Malfunction | HOYA CORPORATION PENTAX TOKYO OFFICE | 2021-10-31 | FDS | PENTAX | Degraded | Insufficient Information | K131902 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725636&pc=MIH | 2017233-2021-02505 | 2021-10-01 | Injury | W. L. GORE & ASSOCIATES, INC. | 2021-10-31 | MIH | GORE TAG THORACIC ENDOPROSTHESIS | Material Puncture/Hole | Aneurysm | P040043 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725637&pc=BTR | 3012307300-2021-10487 | 2021-09-14 | Malfunction | SMITHS MEDICAL INTERNATIONAL, LTD. | 2021-10-31 | BTR | PORTEX | Leak/Splash | No Clinical Signs, Symptoms or Conditions | K931735 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT INTUBATED WITH 2 DIFFERENT ENDOTRACHEAL TUBES, THE CUFFS DID NOT HOLD PRESSURE. NO CONSEQUENCE HEALTH FOR PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725638&pc=LGW | 3006705815-2021-05355 | 2021-10-11 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE ELECTIVE REPLACEMENT INDICATOR (ERI) MESSAGE DISPLAYED FOR THE IPG. IT IS UNKNOWN IF THE IPG DEPLETED PREMATURELY. THE PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN THE IPG WAS EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725639&pc=QBJ | 3004753838-2021-252787 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE WAS PERFORMED AND PASSED. DOWNLOAD/ PAIRING TRANSMITTER WAS PERFORMED AND PASSED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725640&pc=LGW | 3006705815-2021-05358 | 2021-10-11 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE END OF SERVICE (EOS) MESSAGE DISPLAYED FOR THE IPG ON EXTERNAL DEVICES. IT IS UNKNOWN IF THE IPG DEPLETED PREMATURELY. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN THE IPG WAS EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725642&pc=LGW | 1627487-2021-18000 | 2021-10-12 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE END OF SERVICE (EOS) MESSAGE DISPLAYED FOR THE IPG ON EXTERNAL DEVICES. IT IS UNKNOWN IF THE IPG DEPLETED PREMATURELY. IN TURN, SURGICAL INTERVENTION MAY BE PENDING TO ADDRESS THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725643&pc=LGW | 3006705815-2021-05359 | 2021-10-12 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DISPLAYED THE END OF SERVICE (EOS) MESSAGE. THE PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN THE IPG WAS EXPLANTED AND REPLACED, RESTORING THERAPY. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN WEIGHT BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725644&pc=LGW | 3006705815-2021-05360 | 2021-09-16 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Migration | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER 3006705815-2021-05361. IT WAS REPORTED THE PATIENT EXPERIENCED INEFFECTIVE THERAPY DUE TO LEAD MIGRATION. IN TURN, THE PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN BOTH EXISTING LEADS WERE EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. THERAPY WAS RESTORED POST PROCEDURE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED, A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725645&pc=LGW | 3006705815-2021-05364 | 2021-10-15 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Impedance Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER 3006705815-2021-05363. IT WAS REPORTED THAT PATIENT EXPERIENCED INEFFECTIVE THERAPY DUE TO IMPEDANCES ISSUES. REPROGRAMMING WAS UNSUCCESSFUL. AS A RESULT, PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN THE LEADS WERE EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. THERAPY WAS RESTORED POSTOPERATIVELY. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725646&pc=LGW | 3006705815-2021-05361 | 2021-09-16 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Migration | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER 3006705815-2021-05360. IT WAS REPORTED THE PATIENT EXPERIENCED INEFFECTIVE THERAPY DUE TO LEAD MIGRATION. IN TURN, THE PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN BOTH EXISTING LEADS WERE EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. THERAPY WAS RESTORED POST PROCEDURE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED, A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725647&pc=LGW | 3006705815-2021-05363 | 2021-10-15 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Impedance Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER 3006705815-2021-05364. IT WAS REPORTED THAT PATIENT EXPERIENCED INEFFECTIVE THERAPY DUE TO IMPEDANCES ISSUES. REPROGRAMMING WAS UNSUCCESSFUL. AS A RESULT, PATIENT UNDERWENT SURGICAL INTERVENTION WHEREIN THE LEADS WERE EXPLANTED AND REPLACED TO ADDRESS THE ISSUE. THERAPY WAS RESTORED POSTOPERATIVELY. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725648&pc=QBJ | 3004753838-2021-252788 | 2021-10-08 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725649&pc=BTO | 3012307300-2021-10489 | 2021-09-01 | Malfunction | SMITHS MEDICAL INTERNATIONAL, LTD. | 2021-10-31 | BTO | PORTEX | Loose or Intermittent Connection | Insufficient Information | K173384 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE USE OF THE PRODUCT, THE INFLATION LINE GOT DETACHED. NO PATIENT INJURY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725650&pc=QBJ | 3004753838-2021-252789 | 2021-10-08 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. TRANSMITTER VOLTAGE TEST WAS PERFORMED AND PASSED. (B)(4) BLUETOOTH PAIRING TEST/DOWNLOAD WAS PERFORMED AND PASSED. A REVIEW OF THE BIN FILE WAS PERFORMED AND SIGNAL LOSS WAS NOT FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION COULD NOT BE DETERMINED. A PROBABLE CAUSE COULD NOT BE DETERMINED AS THE ALLEGATION WAS NOT FOUND. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725651&pc=QBJ | 3004753838-2021-252790 | 2021-10-08 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725652&pc=CAI | 3012307300-2021-10490 | 2021-10-01 | Malfunction | NULL | 2021-10-31 | CAI | PORTEX | Material Fragmentation | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PRE-USE CHECK, THE CUSHION WAS FOUND TORN. NO PATIENT INJURY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725653&pc=NPT | 2015691-2021-06092 | 2021-10-05 | Injury | EDWARDS LIFESCIENCES | 2021-10-31 | NPT | EDWARDS EXPANDABLE INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P140031 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725654&pc=NPT | 2015691-2021-06093 | 2021-10-05 | Death | EDWARDS LIFESCIENCES | 2021-10-31 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Fluid Leak; Malposition of Device | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING AN OPEN CHEST APPROACH (MINI-THORACOTOMY) PROCEDURE A 29MM SAPIEN 3 (S3) VALVE IN THE NATIVE MITRAL ANNULUS. POST DEPLOYMENT, THE S3 VALVE WAS MALPOSITIONED AND LEAKING SEVERELY. PER REPORT, THE PATIENT WAS INTENDED TO HAVE A SURGICAL MITRAL VALVE IMPLANTED, HOWEVER UPON INSPECTION OF THE NATIVE MITRAL VALVE THROUGH A MINI-THORACOTOMY APPROACH, THE PHYSICIAN DEEMED THE ANNULUS TO BE TOO CALCIFIED TO ALLOW FOR SUCH A PROCEDURE AND IT WAS DECIDED TO PROCEED WITH A SAPIEN 3 VALVE HYBRID IMPLANT. AFTER SIZING WITH TEE AND A 29MM BALLOON INFLATION, IT WAS DECIDED TO IMPLANT A 29MM SAPIEN 3 VALVE. THE PHYSICIAN HAD SEWN A FELT STRIP AROUND THE INFLOW OF THE S3 VALVE AND PLACED A FEW SUTURES IN THE NATIVE MITRAL ANNULUS TO FURTHER SECURE THE VALVE AFTER DEPLOYMENT. POST IMPLANT, THE PHYSICIAN PROCEEDED WITH CLOSING THE SEPTUM AND ATRIUM SO THE VALVE COULD BE EVALUATED. THIS TOOK A FEW HOURS, AS THE HEART MUSCLE AND TISSUE WAS FRIABLE, REQUIRING PERICARDIAL PATCHING TO ACHIEVE A SATISFACTORY RESULT. THE PATIENT WAS TAKEN OFF OF BY-PASS, AND WHEN THE HEART WAS FILLED AND COMMENCED BEATING, IT REVEALED THE S3 WAS MALPOSITIONED AND LEAKING SEVERELY. THE PHYSICIAN PERFORMED A FULL STERNOTOMY, AND PROCEEDED TO EXPLANTING THE S3, AND IMPLANTING AN EDWARDS SURGICAL VALVE. THE FCS FOLLOWED UP WITH THE ANESTHESIOLOGIST AND WAS INFORMED THAT THE PATIENT EXPIRED AROUND MIDNIGHT IN THE OPERATING ROOM DUE TO THE ACCESS SITE BLEEDING NOT BEING ABLE TO BE CONTROLLED OR MANAGED. THERE WAS NO ALLEGATION OR INDICATION THAT THE BLEEDING RESULTING IN DEATH WAS RELATED TO THE EDWARDS DEVICES. IN ADDITION, THERE WAS NO ALLEGATION OR INDICATION OF AN EDWARDS DEVICE MALFUNCTION DURING THE PROCEDURE. Manufacturer Narrative: INVESTIGATION IS STILL ONGOING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725655&pc=QBJ | 3004753838-2021-252791 | 2021-10-08 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. THE PRODUCT WAS EVALUATED. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND PASSED. VOLTAGE CHECK WAS PERFORMED AND PASSED. PAIRING WITH NORDIC BLUETOOTH DEVICE WAS PERFORMED AND PASSED. DOWNLOADED WAS PERFORMED AND PASSED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO RESTORE THE CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725656&pc=NPT | 2015691-2021-06094 | 2021-10-04 | Injury | EDWARDS LIFESCIENCES | 2021-10-31 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | No Clinical Signs, Symptoms or Conditions | P140031 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725657&pc=QBJ | 3004753838-2021-252792 | 2021-10-08 | Malfunction | DEXCOM, INC. | 2021-10-31 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Wireless Communication Problem | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SIGNAL LOSS OVER ONE HOUR OCCURRED. A REVIEW OF THE SHARE LOGS WAS PERFORMED AND SIGNAL LOSS WAS FOUND WITHIN THE INVESTIGATION WINDOW. THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS THE TRANSMITTER AND APP WERE UNABLE TO ESTABLISH A CONNECTION. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725658&pc=DYB | 3007284313-2021-01638 | 2021-10-01 | Injury | W. L. GORE & ASSOCIATES, INC. | 2021-10-31 | DYB | GORE DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Rupture | K160254 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725659&pc=BTO | 3012307300-2021-10492 | 2021-10-01 | Malfunction | NULL | 2021-10-31 | BTO | PORTEX | Leak/Splash | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE PRODUCT WAS INSERTED INTO THE PATIENT, LEAKAGE OF AIR FROM IT WAS DETECTED. THIS EVENT WAS OBSERVED FOUR TIMES CONSECUTIVELY. NO PATIENT INJURY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725660&pc=LGZ | 3012307300-2021-10494 |  | Malfunction | NULL | 2021-10-31 | LGZ | VENTILATOR, EMERGENCY | Loose or Intermittent Connection | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A SMITHS MEDICAL VENTILATOR LOOSE OXYGEN PORT. NO PATIENT INVOLVEMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725661&pc=BTO | 3012307300-2021-10495 | 2021-10-03 | Malfunction | SMITHS MEDICAL ASD, INC. | 2021-10-31 | BTO | BIVONA | Difficult to Insert | Insufficient Information | K912469 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE, THE CUSTOMER FOUND IT DIFFICULT TO INSERT THE PRODUCT INTO THE PATIENT. THE CUSTOMER ALSO NOTICED THAT THE PATIENT HAD TROUBLE DISCHARGING SPUTUM. THE CUSTOMER CHECKED THE DEVICE AND FOUND A PROTRUSION IN IT. THE DEVICE WAS EXCHANGED TO RESOLVE THE ISSUE WITH NO PATIENT INJURY. NO PATIENT INJURY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725662&pc=OOK | 3007215625-2021-01889 | 2021-10-01 | Injury | DUBLIN | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT WHO RECEIVED COOLSCULPTING TREATMENT AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725663&pc=DZE | 0001038806-2021-02093 | 2021-09-17 | Injury | BIOMET 3I | 2021-10-31 | DZE | OSSEOTITE TAPERED CERTAIN PREVAIL IMPLANT 6/5 X 15MM | Problem with Removal of Enzymatic Cleaner | No Clinical Signs, Symptoms or Conditions | K130949 |  | 1 | Event Description: IT WAS REPORTED THAT PRIMARY STABILITY COULD NOT BE ACHIEVED WITH THE IMPLANT. TOOTH POSITION # 15. Manufacturer Narrative: ZIMMER BIOMET COMPLAINT (B)(4). A SUMMARY INVESTIGATION HAS BEEN COMPLETED FOR LACK OF PRIMARY STABILITY EVENTS THAT DO NOT ALLEGE A DEFICIENCY WITH THE IMPLANT AND IDENTIFIED THAT THE REPORTED EVENT IS LIKELY DUE TO BIOLOGICAL FACTORS WHICH HAVE AN ADVERSE EFFECT ON IMPLANT STABILITY OR IS RELATED TO SURGICAL TECHNIQUE AND INSERTION TORQUE. DUE TO A WIDE RANGE OF EXTERNAL FACTORS (NON-DESIGN/ NON-MANUFACTURING RELATED), IDENTIFYING A DEFINITIVE ROOT CAUSE IS GENERALLY NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE RECEIVED WHICH INDICATES THAT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT, AN ADDITIONAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725664&pc=BTO | 3012307300-2021-10496 | 2021-09-01 | Malfunction | SMITHS MEDICAL INTERNATIONAL, LTD. | 2021-10-31 | BTO | PORTEX | Leak/Splash | Insufficient Information | K173384 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE USE OF THE PRODUCT, LEAKAGE OF AIR WAS OBSERVED. THE CUSTOMER SUSPECTED IT WAS CAUSED BY THE ANOMALY IN THE PILOT BALLOON. NO PATIENT INJURY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725665&pc=OOK | 3007215625-2021-01890 | 2021-09-29 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: AN ALLERGAN COMPANY REPRESENTATIVE REPORTED THAT A PATIENT RECEIVED COOLSCULPTING TREATMENT AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725666&pc=EOQ | 9610877-2021-01362 |  | Malfunction | HOYA CORPORATION PENTAX TOKYO OFFICE | 2021-10-31 | EOQ | PENTAX | Physical Resistance/Sticking | Insufficient Information | K131028 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725667&pc=LGW | 1627487-2021-18024 | 2021-10-11 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | QUATTRODE LEAD WIDE SPACED, 90 CM | Adverse Event Without Identified Device or Use Problem | Erosion | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 1627487-2021-18021, 1627487-2021-18022, 1627487-2021-18023. IT WAS REPORTED THE PATIENT'S SUPRA ORBITAL LEAD BECAME SUPERFICIAL AND WAS ERODING THROUGH THEIR SKIN. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 DURING WHICH THE LEADS WERE REPOSITIONED, RESOLVING THE EROSION. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725668&pc=LGW | 1627487-2021-18023 | 2021-10-11 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | QUATTRODE LEAD WIDE SPACED, 60 CM | Adverse Event Without Identified Device or Use Problem | Erosion | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 1627487-2021-18021, 1627487-2021-18022, 1627487-2021-18024. IT WAS REPORTED THE PATIENT'S SUPRA ORBITAL LEAD BECAME SUPERFICIAL AND WAS ERODING THROUGH THEIR SKIN. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 DURING WHICH THE LEADS WERE REPOSITIONED, RESOLVING THE EROSION. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725669&pc=DZE | 0002023141-2021-03077 | 2021-05-23 | Injury | ZIMMER DENTAL | 2021-10-31 | DZE | IMP,TSV,4.1MM,SBM,10 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Inflammation; Osteolysis | K072589 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANT AT TOOTH SITE # 4 (UNIVERSAL) WAS REMOVED DUE TO INFECTION AND BONE LOSS. INFLAMMATION IS REPORTED AS A CONSEQUENCE. Manufacturer Narrative: ZIMMER BIOMET COMPLAINT (B)(4). PATIENT IDENTIFIER IS (B)(6). PATIENT WEIGHT IS NOT PROVIDED / UNKNOWN. INITIAL REPORTER¿S TITLE IS NOT PROVIDED / UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725670&pc=LGW | 1627487-2021-18022 | 2021-10-11 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | QUATTRODE LEAD WIDE SPACED, 60 CM | Adverse Event Without Identified Device or Use Problem | Erosion | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 1627487-2021-18021, 1627487-2021-18023, 1627487-2021-18024. IT WAS REPORTED THE PATIENT'S SUPRA ORBITAL LEAD BECAME SUPERFICIAL AND WAS ERODING THROUGH THEIR SKIN. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 DURING WHICH THE LEADS WERE REPOSITIONED, RESOLVING THE EROSION. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725671&pc=LGW | 3006705815-2021-05383 | 2021-10-02 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | High impedance | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 3006705815-2021-05382. IT WAS REPORTED THAT THE PATIENT'S SYSTEM IS AUTOREDUCING DUE TO HIGH IMPEDANCES. TROUBLESHOOTING WAS UNABLE TO RESOLVE THE ISSUE. AS A RESULT, THE PATIENT MAY UNDERGO SURGICAL INTERVENTION TO ADDRESS THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725672&pc=LGW | 1627487-2021-18021 | 2021-10-11 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | QUATTRODE LEAD WIDE SPACED, 90 CM | Adverse Event Without Identified Device or Use Problem | Erosion | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 1627487-2021-18021, 1627487-2021-18022, 1627487-2021-18023. IT WAS REPORTED THE PATIENT'S SUPRAORBITAL LEAD BECAME SUPERFICIAL AND WAS ERODING THROUGH THEIR SKIN. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 DURING WHICH THE LEADS WERE REPOSITIONED, RESOLVING THE EROSION. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725673&pc=LGW | 3006705815-2021-05382 | 2021-10-02 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | High impedance | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 3006705815-2021-05383. IT WAS REPORTED THAT THE PATIENT'S SYSTEM IS AUTOREDUCING DUE TO HIGH IMPEDANCES. TROUBLESHOOTING WAS UNABLE TO RESOLVE THE ISSUE. AS A RESULT, THE PATIENT MAY UNDERGO SURGICAL INTERVENTION TO ADDRESS THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725674&pc=LGW | 3006705815-2021-05343 | 2021-10-09 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Fracture; Migration | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 3006705815-2021-05343. IT WAS REPORTED THE PATIENT'S SYSTEM WAS AUTOREDUCING. X-RAY REVEALED LEAD MIGRATION AND FRACTURE. AS A RESULT, THE PATIENT MAY UNDERGO SURGICAL INTERVENTION TO ADDRESS THE ISSUE. IT IS UNKNOWN WHICH LEAD HAD MIGRATED AND WAS FRACTURED, SO IT IS BEING REPORTED ON BOTH. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725675&pc=LGW | 1627487-2021-17981 | 2021-10-11 | Malfunction | ABBOTT MEDICAL | 2021-10-31 | LGW | PROCLAIM 7 ELITE IMPLANTABLE PULSE GENERATOR | Premature Elective Replacement Indicator | No Clinical Signs, Symptoms or Conditions | P010032 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT'S CONTROLLER WAS DISPLAYING THE REPLACE GENERATOR SOON MESSAGE. THE PATIENT'S APP WAS UPDATED AND THE MESSAGE WAS CLEARED, RESOLVING THE ISSUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725676&pc=LGW | 3006705815-2021-05342 | 2021-10-09 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Fracture; Migration | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 3006705815-2021-05342. IT WAS REPORTED THE PATIENT'S SYSTEM WAS AUTOREDUCING. X-RAY REVEALED LEAD MIGRATION AND FRACTURE. AS A RESULT, THE PATIENT MAY UNDERGO SURGICAL INTERVENTION TO ADDRESS THE ISSUE. IT IS UNKNOWN WHICH LEAD HAD MIGRATED AND WAS FRACTURED, SO IT IS BEING REPORTED ON BOTH. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725677&pc=LGW | 3006705815-2021-05386 | 2021-10-11 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | EON MINI IPG, 16-CHANNEL RECHARGEABLE | Adverse Event Without Identified Device or Use Problem | Implant Pain | P010032 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WAS EXPERIENCING PAIN AT THE IPG SITE DUE TO LOSING 100 POUNDS. AS A RESULT, THE PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 DURING WHICH THE IPG WAS REPOSITIONED, RESOLVING THE PAIN. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725678&pc=LKK | 3004209178-2021-16171 | 2021-07-24 | Injury | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2021-10-31 | LKK | SYNCHROMED II | Adverse Event Without Identified Device or Use Problem | Headache | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (HEALTHCARE PROVIDER, CLINICAL STUDY) REGARDING A PATIENT WHO WAS RECEIVING DILAUDID (HYDROMORPHONE) (0.5 MG AT 1.00122 MG/DAY) AND BUPIVACAINE (7.1 MG AT 14.21729 MG/DAY) VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT ON (B)(6) 2021 THE PATIENT STAYED IN PATIENT A SECOND NIGHT FOLLOWING INITIAL PUMP AND CATHETER IMPLANT SECONDARY TO A SEVERE POSITIONAL HEADACHE. THE PATIENT WAS STARTED ON FIORICET. THE ONSET OF THE HEADACHE ONE DAY POST IMPLANT ON (B)(6) 2021. ON (B)(6) 2021 THE PATIENT REPORTED PERSISTENT HEADACHES THAT WERE WORSE WITH AMBULATION AND LIGHT SENSATIONS. THE PATIENT WAS ENCOURAGED TO INCREASE CAFFEINE AND FLUID INTAKE. ON (B)(6) 2021 THE PATIENT REPORTED THE HEADACHE HAD RESOLVED. THE EVENT RESOLVED WITHOUT SEQUELAE ON (B)(6) 2021. THE CLINICAL DIAGNOSIS WAS POSITIONAL HEADACHE. THE EVENT REQUIRED IN-PATIENT HOSPITALIZATION. THE ETIOLOGY OF THE EVENT INDICATED THE RELATIONSHIP OF THE EVENT TO THE DEVICE OR THERAPY WAS NOT RELATED AND INDICATED THE RELATIONSHIP OF THE EVENT TO THE IMPLANT PROCEDURE WAS RELATED. THE EVENT WAS RELATED TO SURGERY/ANESTHESIA. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID: 8781, SERIAL#: (B)(4), IMPLANTED: (B)(6) 2021, PRODUCT TYPE: CATHETER. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 8781, SERIAL/LOT #: (B)(4), UBD: 22-JUN-2023, UDI#: (B)(4). IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725679&pc=FTR | 1645337-2021-12110 | 2021-03-01 | Injury | MENTOR TEXAS | 2021-10-31 | FTR | MENTOR MEMORYGEL BREAST IMPLANT | Material Rupture | Deformity/ Disfigurement | P030053 |  | 1 | Event Description: IT WAS REPORTED THAT A (B)(6) YEAR-OLD CAUCASIAN FEMALE WHO UNDERWENT PRIMARY BREAST AUGMENTATION WITH A 600CC MENTOR MEMORYGEL BREAST IMPLANT EXPERIENCED LEFT SIDED RUPTURE POST PROCEDURE. AS A RESULT, AN EXPLANT IS PLANNED FOR (B)(6) 2021. Manufacturer Narrative: SINCE THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS CAN BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. IF THE COMPLAINT DEVICE IS RECEIVED IN THE FUTURE, THE INVESTIGATION WILL BE REOPENED AND CONDUCTED AS APPROPRIATE. A MANUFACTURING RECORD EVALUATION (MRE) WAS PERFORMED, AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE MRE VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. REASON FOR DEVICE EXPLANT AND/OR REOPERATION: DEFLATION. FUTURE EXPLANT DATE: (B)(6) 2021. MANUFACTURER¿S REFERENCE NUMBER: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725680&pc=OOK | 3007215625-2021-01891 | 2021-10-01 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT WHO RECEIVED COOLSCULPTING TREATMENT AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA ON THE TREATED AREA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725681&pc=NAY | 2955842-2021-11407 | 2021-07-30 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST | Failure to Deliver Energy; Material Split, Cut or Torn | No Clinical Signs, Symptoms or Conditions | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725682&pc=BTT | 9611451-2021-01236 |  | Malfunction | FISHER & PAYKEL HEALTHCARE LTD | 2021-10-31 | BTT | AIRVO 2 HUMIDIFIER | No Audible Alarm | No Clinical Signs, Symptoms or Conditions | K131895 |  | 1 | Event Description: A DISTRIBUTOR IN (B)(4) REPORTED VIA A FISHER AND PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE THAT THE AUDIBLE ALARM OF A PT101 AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: (B)(4). METHOD: THE COMPLAINT AIRVO 2 HUMIDIFIER WAS NOT RECEIVED AT FISHER & PAYKEL HEALTHCARE (F&P) FOR EVALUATION. OUR INVESTIGATION IS BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: THE CUSTOMER HAS STATED THAT THE AUDIBLE ALARM OF A PT101 AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. CONCLUSION: WE ARE UNABLE TO DETERMINE THE CAUSE OF THE FAILURE MODE IN THIS INSTANCE. A CHANGE TO THE CONTROL BOARD HAS SINCE BEEN IMPLEMENTED TO IMPROVE THE SPEAKER'S PERFORMANCE. IT SHOULD BE NOTED THAT THE AIRVO 2 IS EQUIPPED WITH BOTH AUDIBLE AND VISUAL ALARMS. THE AIRVO 2 USER MANUAL STATES THAT THE "AIRVO 2 IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH-FLOW, WARMED AND HUMIDIFIED RESPIRATORY GASES" AND THAT "THE UNIT IS NOT INTENDED FOR LIFE SUPPORT." THE USER MANUAL WARNS THE USER: PRIOR TO EACH PATIENT USE, ENSURE THAT THE AUDITORY ALARM SIGNAL IS AUDIBLE BY CONDUCTING THE ALARM SYSTEM FUNCTIONALITY CHECK DESCRIBED IN THE ALARMS SECTION THE ALARM SYSTEM FUNCTIONALITY CHECK INSTRUCTS THE USER ON HOW TO CHECK THE ALARM AND STATES THAT "IF EITHER ALARM SIGNAL IS ABSENT, DO NOT USE THE UNIT. CONTACT YOUR FISHER & PAYKEL HEALTHCARE REPRESENTATIVE."
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725683&pc=OOK | 3007215625-2021-01892 | 2021-10-01 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK COOLSCULPTING | Adverse Event Without Identified Device or Use Problem | Urticaria | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT THAT A PATIENT TREATED WITH COOLSCULPTING TREATMENT DEVELOPED SWELLING OF FACE AND BODY WITH GENERALIZED LARGE ITHCY HIVES. EMERGENCY CARE WAS NEEDED. Manufacturer Narrative: ALLERGAN AESTHETICS IS DOING DUE DILIGENCE IN REQUESTING ADDITIONAL EVENT INFORMATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725684&pc=FWM | 1645337-2021-12111 | 2021-09-08 | Injury | MENTOR TEXAS | 2021-10-31 | FWM | MENTOR SMOOTH ROUND MODERATE PROFILE | Material Rupture | Deformity/ Disfigurement | P990075 |  | 1 | Event Description: IT WAS REPORTED THAT A (B)(6) YEAR-OLD FEMALE WHO UNDERWENT PRIMARY BREAST AUGMENTATION WITH A MENTOR SMOOTH ROUND MODERATE PROFILE 350CC SALINE PROSTHESIS EXPERIENCED SPONTANEOUS LEFT SIDED DEFLATION POST PROCEDURE. THE DEFLATION WAS DIAGNOSED AT THE PHYSICIAN¿S OFFICE. AS A RESULT, AN EXPLANT WAS PERFORMED ON (B(6) 2021. Manufacturer Narrative: SINCE THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS CAN BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. IF THE COMPLAINT DEVICE IS RECEIVED IN THE FUTURE, THE INVESTIGATION WILL BE REOPENED AND CONDUCTED AS APPROPRIATE. A MANUFACTURING RECORD EVALUATION (MRE) WAS PERFORMED, AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE MRE VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. REASON FOR DEVICE EXPLANT AND/OR REOPERATION: DEFLATION. MANUFACTURER¿S REFERENCE NUMBER: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725685&pc=OOK | 3007215625-2021-01893 | 2021-09-29 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT WHO RECEIVED COOLSCULPTING TREATMENT TO THE LOWER ABDOMEN AND FLANKS AREA AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725687&pc=OOK | 3007215625-2021-01894 | 2021-07-15 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | COOLADVANTAGE APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT TREATED WITH COOLSCULPTING TO THE FLANKS AND UPPER AND LOWER ABDOMEN WITH THE COOLADVANTAGE APPLICATOR WAS DIAGNOSED WITH PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725688&pc=LKK | 3004209178-2021-16172 | 2021-10-27 | Malfunction | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2021-10-31 | LKK | SYNCHROMED II | Migration or Expulsion of Device | Pain; Discomfort | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (HEALTHCARE PROVIDER, CLINICAL STUDY) REGARDING A PATIENT WHO WAS RECEIVING DRUG, UNKNOWN ( AT ) VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT ON (B)(6) 2021 THE PATIENT REPORTED PAIN AT THE POCKET SITE AND THE PUMP APPEARED TO BE MOVING. EXAMINATION REVEALED THE SITE APPEARED TO BE TENDER. IT WAS RECOMMENDED TO ICE THREE TIMES A DAY. NO ACTION WAS TAKEN. THE OUTCOME WAS NOTED AS ONGOING. THE CLINICAL DIAGNOSIS WAS PAIN AT THE POCKET SITE. THE ETIOLOGY OF THE EVENT INDICATED THE RELATIONSHIP OF THE EVENT TO THE DEVICE OR THERAPY WAS RELATED AND INDICATED THE RELATIONSHIP OF THE EVENT TO THE IMPLANT PROCEDURE WAS NOT RELATED. THE INCISIONAL SITE/DEVICE TRACT WAS PUMP POCKET. THE EVENT DATE WAS (B)(6) 2021. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725689&pc=NAY | 2955842-2021-11408 | 2021-09-22 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-31 | NAY | ENDOWRIST | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725690&pc=HCG | 3005168196-2021-02434 | 2021-10-07 | Malfunction | PENUMBRA, INC. | 2021-10-31 | HCG | PACKING COIL LP | Separation Failure | No Clinical Signs, Symptoms or Conditions | K192955 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A COIL EMBOLIZATION PROCEDURE IN THE SUPERIOR MESENTERIC ARTERY (SMA) AND INFERIOR MESENTERIC ARTERY (IMA) ENDOLEAK SAC USING PACKING COIL LPS, RUBY COIL LPS, A LP SYSTEM DETACHMENT HANDLE (HANDLE), AND A NON-PENUMBRA MICROCATHETER. DURING THE PROCEDURE, THE PHYSICIAN ADVANCED A PACKING COIL LP INTO THE TARGET LOCATION AND ATTEMPTED TO DETACH IT USING A HANDLE; HOWEVER, THE PACKING COIL LP FAILED TO DETACH. THE PHYSICIAN THEN ATTEMPTED TO MANUALLY DETACH THE PACKING COIL LP USING A HEMOSTAT, BUT THE COIL STILL WOULD NOT DETACH. THEREFORE, THE PACKING COIL LP WAS REMOVED. SUBSEQUENTLY, THE PHYSICIAN WAS UNABLE TO DETACH TWO RUBY COIL LPS IN THE TARGET LOCATION USING THE SAME HANDLE. THEREFORE, THE HANDLE WAS REMOVED. THE PHYSICIAN ATTEMPTED TO DETACH THE SAME RUBY COILS USING A NEW HANDLE, BUT THE TWO RUBY COIL LPS WOULD NOT DETACH. IT WAS ALSO REPORTED THAT THE PHYSICIAN ATTEMPTED TO MANUALLY DETACH THE TWO RUBY COIL LPS USING A HEMOSTAT; HOWEVER, THE ATTEMPT WAS UNSUCCESSFUL. THEREFORE, THE TWO RUBY COIL LPS AND HANDLE WERE REMOVED. THE PROCEDURE WAS COMPLETED USING NEW LP COILS WITH A HEMOSTAT AND THE SAME MICROCATHETER. THERE WAS NO REPORT OF AN ADVERSE EFFECT TO THE PATIENT. Manufacturer Narrative: EVALUATION OF THE PACKING COIL LP REVEALED THAT THE EMBOLIZATION COIL WAS INTACT WITH THE PUSHER ASSEMBLY, AND THE COMPLAINT WAS CONFIRMED. EVALUATION REVEALED THAT THE PET LOCK WAS SEPARATED, AND THE PULL WIRE WAS FRACTURED OFF THE PROXIMAL END OF THE PUSHER ASSEMBLY. THIS DAMAGE LIKELY OCCURRED DURING THE ATTEMPTS TO DETACH THE EMBOLIZATION COIL MANUALLY WITH A HEMOSTAT. BASED ON THE RETURNED CONDITION, THE ROOT CAUSE OF THE REPORTED COMPLAINT COULD NOT BE DETERMINED. FURTHER EVALUATION REVEALED KINKS IN THE PUSHER ASSEMBLY. THIS DAMAGE WAS INCIDENTAL TO THE REPORTED COMPLAINT AND THE ROOT CAUSE COULD NOT BE DETERMINED. PENUMBRA COILS ARE VISUALLY INSPECTED DURING IN-PROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. (B)(4). THIS REPORT IS ASSOCIATED WITH MFR REPORT NUMBERS: 3005168196-2021-02435, 3005168196-2021-02436.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725691&pc=HCG | 3005168196-2021-02435 | 2021-10-07 | Malfunction | PENUMBRA, INC. | 2021-10-31 | HCG | RUBY COIL LP | Separation Failure | No Clinical Signs, Symptoms or Conditions | K192955 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A COIL EMBOLIZATION PROCEDURE IN THE SUPERIOR MESENTERIC ARTERY (SMA) AND INFERIOR MESENTERIC ARTERY (IMA) ENDOLEAK SAC USING PACKING COIL LPS, RUBY COIL LPS, A LP SYSTEM DETACHMENT HANDLE (HANDLE), AND A NON-PENUMBRA MICROCATHETER. DURING THE PROCEDURE, THE PHYSICIAN ADVANCED A PACKING COIL LP INTO THE TARGET LOCATION AND ATTEMPTED TO DETACH IT USING A HANDLE; HOWEVER, THE PACKING COIL LP FAILED TO DETACH. THE PHYSICIAN THEN ATTEMPTED TO MANUALLY DETACH THE PACKING COIL LP USING A HEMOSTAT, BUT THE COIL STILL WOULD NOT DETACH. THEREFORE, THE PACKING COIL LP WAS REMOVED. SUBSEQUENTLY, THE PHYSICIAN WAS UNABLE TO DETACH TWO RUBY COIL LPS IN THE TARGET LOCATION USING THE SAME HANDLE. THEREFORE, THE HANDLE WAS REMOVED. THE PHYSICIAN ATTEMPTED TO DETACH THE SAME RUBY COILS USING A NEW HANDLE, BUT THE TWO RUBY COIL LPS WOULD NOT DETACH. IT WAS ALSO REPORTED THAT THE PHYSICIAN ATTEMPTED TO MANUALLY DETACH THE TWO RUBY COIL LPS USING A HEMOSTAT; HOWEVER, THE ATTEMPT WAS UNSUCCESSFUL. THEREFORE, THE TWO RUBY COIL LPS AND HANDLE WERE REMOVED. THE PROCEDURE WAS COMPLETED USING NEW LP COILS WITH A HEMOSTAT AND THE SAME MICROCATHETER. THERE WAS NO REPORT OF AN ADVERSE EFFECT TO THE PATIENT. Manufacturer Narrative: EVALUATION OF THE RUBY COIL LP REVEALED THAT THE EMBOLIZATION COIL WAS INTACT WITH THE PUSHER ASSEMBLY, AND THE COMPLAINT WAS CONFIRMED. EVALUATION REVEALED THAT THE PET LOCK WAS SEPARATED, AND THE PULL WIRE WAS FRACTURED OFF THE PROXIMAL END OF THE PUSHER ASSEMBLY. THIS DAMAGE LIKELY OCCURRED DURING THE ATTEMPTS TO DETACH THE EMBOLIZATION COIL MANUALLY WITH A HEMOSTAT. BASED ON THE RETURNED CONDITION, THE ROOT CAUSE OF THE REPORTED COMPLAINT COULD NOT BE DETERMINED. FURTHER EVALUATION REVEALED MULTIPLE KINKS IN THE PUSHER ASSEMBLY. THIS DAMAGE WAS INCIDENTAL TO THE REPORTED COMPLAINT AND THE ROOT CAUSE COULD NOT BE DETERMINED. HOWEVER, IF THESE KINKS WERE PRESENT DURING THE PROCEDURE, THEY MAY HAVE CONTRIBUTED TO FRICTION BETWEEN THE PULL WIRE AND THE INNER LUMEN OF THE PUSHER ASSEMBLY AND PREVENTED THE EMBOLIZATION COIL FROM BEING DETACHED. PENUMBRA COILS ARE VISUALLY INSPECTED DURING IN-PROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. (B)(4). THIS REPORT IS ASSOCIATED WITH MFR REPORT NUMBERS: 3005168196-2021-02434, 3005168196-2021-02436.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725692&pc=HCG | 3005168196-2021-02436 | 2021-10-07 | Malfunction | PENUMBRA, INC. | 2021-10-31 | HCG | RUBY COIL LP | Separation Failure | No Clinical Signs, Symptoms or Conditions | K192955 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A COIL EMBOLIZATION PROCEDURE IN THE SUPERIOR MESENTERIC ARTERY (SMA) AND INFERIOR MESENTERIC ARTERY (IMA) ENDOLEAK SAC USING PACKING COIL LPS, RUBY COIL LPS, A LP SYSTEM DETACHMENT HANDLE (HANDLE), AND A NON-PENUMBRA MICROCATHETER. DURING THE PROCEDURE, THE PHYSICIAN ADVANCED A PACKING COIL LP INTO THE TARGET LOCATION AND ATTEMPTED TO DETACH IT USING A HANDLE; HOWEVER, THE PACKING COIL LP FAILED TO DETACH. THE PHYSICIAN THEN ATTEMPTED TO MANUALLY DETACH THE PACKING COIL LP USING A HEMOSTAT, BUT THE COIL STILL WOULD NOT DETACH. THEREFORE, THE PACKING COIL LP WAS REMOVED. SUBSEQUENTLY, THE PHYSICIAN WAS UNABLE TO DETACH TWO RUBY COIL LPS IN THE TARGET LOCATION USING THE SAME HANDLE. THEREFORE, THE HANDLE WAS REMOVED. THE PHYSICIAN ATTEMPTED TO DETACH THE SAME RUBY COILS USING A NEW HANDLE, BUT THE TWO RUBY COIL LPS WOULD NOT DETACH. IT WAS ALSO REPORTED THAT THE PHYSICIAN ATTEMPTED TO MANUALLY DETACH THE TWO RUBY COIL LPS USING A HEMOSTAT; HOWEVER, THE ATTEMPT WAS UNSUCCESSFUL. THEREFORE, THE TWO RUBY COIL LPS AND HANDLE WERE REMOVED. THE PROCEDURE WAS COMPLETED USING NEW LP COILS WITH A HEMOSTAT AND THE SAME MICROCATHETER. THERE WAS NO REPORT OF AN ADVERSE EFFECT TO THE PATIENT. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. WITHOUT THE RETURN OF THE DEVICE, THE ROOT CAUSE OF THE PROBLEM CANNOT BE DETERMINED. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. THIS REPORT IS ASSOCIATED WITH MFR REPORT NUMBERS: 3005168196-2021-02434, 3005168196-2021-02435.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725693&pc=OOK | 3007215625-2021-01895 | 2021-10-01 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT TREATED WITH COOLSCULPTING TREATMENT TO UPPER AND LOWER ABDOMEN MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725694&pc=BTT | 9611451-2021-01237 |  | Malfunction | FISHER & PAYKEL HEALTHCARE LTD | 2021-10-31 | BTT | AIRVO 2 HUMIDIFIER | No Audible Alarm | No Clinical Signs, Symptoms or Conditions | K131895 |  | 1 | Event Description: A HEALTHCARE FACILITY IN (B)(6) REPORTED VIA A FISHER & PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE, THAT THE AUDIBLE ALARM OF A PT101 AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: (B)(4). METHOD: THE COMPLAINT PT101 AIRVO 2 HUMIDIFIER WAS NOT RECEIVED AT FISHER & PAYKEL HEALTHCARE (F&P) FOR EVALUATION. OUR INVESTIGATION IS BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: THE CUSTOMER REPORTED THAT THE AUDIBLE ALARM OF A PT101 AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. CONCLUSION: WITHOUT THE COMPLAINT DEVICE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THE FAILURE MODE IN THIS INSTANCE. IT SHOULD BE NOTED THAT THE AIRVO 2 IS EQUIPPED WITH BOTH AUDIBLE AND VISUAL ALARMS. THE AIRVO 2 USER MANUAL STATES THAT THE "AIRVO 2 IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH-FLOW, WARMED AND HUMIDIFIED RESPIRATORY GASES" AND THAT "THE UNIT IS NOT INTENDED FOR LIFE SUPPORT." THE USER MANUAL WARNS THE USER: PRIOR TO EACH PATIENT USE, ENSURE THAT THE AUDITORY ALARM SIGNAL IS AUDIBLE BY CONDUCTING THE ALARM SYSTEM FUNCTIONALITY CHECK DESCRIBED IN THE ALARMS SECTION. THE ALARM SYSTEM FUNCTIONALITY CHECK INSTRUCTS THE USER ON HOW TO CHECK THE ALARM AND STATES THAT "IF EITHER ALARM SIGNAL IS ABSENT, DO NOT USE THE UNIT. CONTACT YOUR FISHER & PAYKEL HEALTHCARE REPRESENTATIVE."
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725695&pc=DQX | 3003775027-2021-00175 | 2021-10-10 | Injury | ASAHI INTECC CO., LTD. | 2021-10-31 | DQX | GAIA THIRD | Detachment of Device or Device Component | Foreign Body In Patient | K171933 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725696&pc=LGW | 3006705815-2021-05397 | 2021-10-12 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Adverse Event Without Identified Device or Use Problem | Implant Pain | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING DISCOMFORT AT THE POCKET SITE. AS A RESULT, ON (B)(6) 2021, THE PHYSICIAN REVISED THE BATTERY LOCATION BY PLACING THE IPG SLIGHTLY LOWER IN THE POCKET. SURGICAL INTERVENTION RESOLVED THE ISSUE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. BASED ON THE INFORMATION PROVIDED A DEVICE PROBLEM WAS NOT IDENTIFIED, AS A RESULT A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725697&pc=BTT | 9611451-2021-01238 |  | Malfunction | FISHER & PAYKEL HEALTHCARE LTD | 2021-10-31 | BTT | AIRVO 2 HUMIDIFIER | No Audible Alarm | No Clinical Signs, Symptoms or Conditions | K131895 |  | 1 | Event Description: A HEALTHCARE FACILITY IN (B)(6) REPORTED VIA A FISHER & PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE, THAT THE AUDIBLE ALARM OF A PT101 AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: (B)(4). METHOD: THE COMPLAINT AIRVO 2 HUMIDIFIER WAS RECEIVED AT OUR FISHER & PAYKEL HEALTHCARE (F&P) REGIONAL OFFICE IN (B)(6) AND WAS INSPECTED BY A TRAINED F&P SERVICE TECHNICIAN. THE DEVICE WAS PERFORMANCE TESTED AND THE AUDIBLE ALARM FUNCTION WAS CHECKED. THE DEVICE WAS THEN DISPOSED OF. RESULTS: DURING TESTING IT WAS FOUND THAT THE AUDIBLE ALARM OF THE AIRVO 2 HUMIDIFIER WAS NOT FUNCTIONING. CONCLUSION: IT WAS CONFIRMED THAT THE AUDIBLE ALARM WAS NOT FUNCTIONING. WE ARE UNABLE TO DETERMINE THE CAUSE OF THE FAILURE MODE IN THIS INSTANCE. IT SHOULD BE NOTED THAT THE AIRVO 2 IS EQUIPPED WITH BOTH AUDIBLE AND VISUAL ALARMS. THE AIRVO 2 HUMIDIFIER USER MANUAL STATES THAT THE "AIRVO 2 IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH-FLOW, WARMED AND HUMIDIFIED RESPIRATORY GASES" AND THAT "THE UNIT IS NOT INTENDED FOR LIFE SUPPORT." THE USER MANUAL WARNS THE USER: PRIOR TO EACH PATIENT USE, ENSURE THAT THE AUDITORY ALARM SIGNAL IS AUDIBLE BY CONDUCTING THE ALARM SYSTEM FUNCTIONALITY CHECK DESCRIBED IN THE ALARMS SECTION. THE ALARM SYSTEM FUNCTIONALITY CHECK INSTRUCTS THE USER ON HOW TO TEST THE ALARM AND STATES THAT "IF EITHER ALARM SIGNAL IS ABSENT, DO NOT USE THE UNIT. CONTACT YOUR FISHER & PAYKEL HEALTHCARE REPRESENTATIVE."
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725699&pc=NIK | 2017865-2021-35549 | 2021-10-09 | Malfunction | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-31 | NIK | QUADRA ASSURA ICD | Capturing Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2021-35550. IT WAS REPORTED THAT THE PATIENT PRESENTED FOR ROUTINE FOLLOW UP. WHEN THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS INTERROGATED, IT WAS NOTED THAT THERE WERE INCORRECT PARAMETER VALUES. THESE VALUES WERE PROGRAMMED BACK TO THE NORMAL VALUES AFTER WHICH IT WAS NOTED THAT THE LEFT VENTRICULAR (LV) LEAD EXHIBITED HIGH CAPTURE THRESHOLDS. IT WAS UNKNOWN WHETHER THIS WAS AN ICD OR LV LEAD ISSUE. THEREFORE, FURTHER PROGRAMMING CHANGES WERE MADE TO ADDRESS THE ISSUE. THE PATIENT WAS IN STABLE CONDITION. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725701&pc=LGW | 1627487-2021-18027 | 2021-10-12 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS RECEIVING THE "REPLACE GENERATOR" MESSAGE AND HAD LOST THERAPY. A RESULT, THE PHYSICIAN EXPLANTED AND REPLACED THE PATIENT'S IPG. SURGICAL INTERVENTION RESOLVED THE ISSUE. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725702&pc=OOK | 3007215625-2021-01896 | 2021-05-15 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | COOLMINI APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: A TREATMENT PROVIDER REPORTED THAT A PATIENT TREATED WITH COOLSUCLPTING TO THE SUBMENTAL WITH THE COOLMINI APPLICATOR WAS DIAGNOSED WITH PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725703&pc=LGW | 1627487-2021-18026 | 2021-10-12 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IPG WAS INOPERABLE. AS A RESULT, THE PHYSICIAN EXPLANTED AND REPLACED THE IPG. SURGICAL INTERVENTION RESOLVED THE ISSUE. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725704&pc=LGW | 3006705815-2021-05393 | 2021-10-11 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Battery Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IPG IS AT END OF LIFE AND UNABLE TO CONNECT TO EXTERNAL DEVICES. SURGICAL INTERVENTION IS ANTICIPATED. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725705&pc=LGW | 1627487-2021-18025 | 2021-10-11 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | Wireless Communication Problem | Inadequate Pain Relief | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT, FOLLOWING A HAND SURGERY, THE PATIENT'S IPG IS INOPERABLE AND UNABLE TO CONNECT TO EXTERNAL DEVICES. SURGICAL INTERVENTION IS ANTICIPATED. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. THE DEVICE IS INCLUDED IN THE NEUROMODULATION IMPLANTABLE PULSE GENERATOR (IPG) INOPERABLE WHEN EXPOSED TO MONOPOLAR ELECTROSURGERY ADVISORY NOTICE ISSUED BY ABBOTT ON 02 JUNE 2017.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725708&pc=OOK | 3007215625-2021-01897 | 2021-10-01 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK COOLSCULPTING | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT TREATED WITH COOLSCULPTING AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725709&pc=OOK | 3007215625-2021-01898 | 2021-10-01 | Injury | ALLERGAN PLEASANTON | 2021-10-31 | OOK | UNK APPLICATOR | Adverse Event Without Identified Device or Use Problem | Hyperplasia | K160259 |  | 1 | Event Description: ALLERGAN AESTHETICS RECEIVED A REPORT OF A PATIENT TREATED WITH COOLSCULPTING TO THE ABDOMEN AND FLANKS AND MAY HAVE DEVELOPED PARADOXICAL ADIPOSE HYPERPLASIA. Manufacturer Narrative: ACCORDING TO THE COOLSCULPTING USER MANUAL, UNDER RARE ADVERSE EVENTS, PARADOXICAL ADIPOSE HYPERPLASIA IS CHARACTERIZED BY A VISIBLY ENLARGED TISSUE VOLUME WITHIN THE TREATMENT AREA, WHICH MAY DEVELOP TWO TO FIVE MONTHS AFTER TREATMENT. SURGICAL INTERVENTION MAY BE REQUIRED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF AND WHEN ADDITIONAL INFORMATION IS OBTAINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725710&pc=LGW | 3006705815-2021-05406 | 2021-10-05 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-31 | LGW | PROCLAIM 5 ELITE IMPLANTABLE PULSE GENERATOR | High impedance | Failure of Implant; Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBERS: 1627487-2021-18031, 1627487-2021-18032 IT WAS REPORTED THAT THE PATIENT'S SYSTEM COULD NOT GO INTO MRI MODE. THE PATIENT ALSO EXPERIENCED INEFFECTIVE STIMULATION AND NEEDED BETTER COVERAGE. DIAGNOSTICS REVEALED HIGH IMPEDANCES ON SEVERAL CONTACTS. AS A RESULT, SURGICAL INTERVENTION WAS UNDERTAKEN WHEREIN THE LEADS WERE EXPLANTED AND REPLACED . INTRAOPERATIVELY, THE IMPEDANCES WERE CHECKED WITH THE NEW LEADS AND THERE WERE STILL IMPEDANCE ISSUES THAT DID NOT CLEAR. AS A RESULT, THE PHYSICIAN ALSO EXPLANTED AND REPLACED THE IPG DURING THE SAME PROCEDURE. POST-OPERATIVELY, STIMULATION THERAPY WAS RESTORED AND THE SYSTEM COULD GO INTO MRI MODE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE PATIENT INFORMATION. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725711&pc=LGW | 1627487-2021-18031 | 2021-10-05 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | High impedance | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBERS: 3006705815-2021-05406, 1627487-2021-18032. IT WAS REPORTED THAT THE PATIENT'S SYSTEM COULD NOT GO INTO MRI MODE. THE PATIENT ALSO EXPERIENCED INEFFECTIVE STIMULATION AND NEEDED BETTER COVERAGE. DIAGNOSTICS REVEALED HIGH IMPEDANCES ON SEVERAL CONTACTS. AS A RESULT, SURGICAL INTERVENTION WAS UNDERTAKEN WHEREIN THE LEADS WERE EXPLANTED AND REPLACED. INTRAOPERATIVELY, THE IMPEDANCES WERE CHECKED WITH THE NEW LEADS AND THERE WERE STILL IMPEDANCE ISSUES THAT DID NOT CLEAR. AS A RESULT, THE PHYSICIAN ALSO EXPLANTED AND REPLACED THE IPG DURING THE SAME PROCEDURE. POST-OPERATIVELY, STIMULATION THERAPY WAS RESTORED AND THE SYSTEM COULD GO INTO MRI MODE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE PATIENT INFORMATION. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725712&pc=LGW | 1627487-2021-18032 | 2021-10-05 | Injury | ABBOTT MEDICAL | 2021-10-31 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | High impedance | Inadequate Pain Relief | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBERS: 3006705815-2021-05406, 1627487-2021-18031. IT WAS REPORTED THAT THE PATIENT'S SYSTEM COULD NOT GO INTO MRI MODE. THE PATIENT ALSO EXPERIENCED INEFFECTIVE STIMULATION AND NEEDED BETTER COVERAGE. DIAGNOSTICS REVEALED HIGH IMPEDANCES ON SEVERAL CONTACTS. AS A RESULT, SURGICAL INTERVENTION WAS UNDERTAKEN WHEREIN THE LEADS WERE EXPLANTED AND REPLACED . INTRAOPERATIVELY, THE IMPEDANCES WERE CHECKED WITH THE NEW LEADS AND THERE WERE STILL IMPEDANCE ISSUES THAT DID NOT CLEAR. AS A RESULT, THE PHYSICIAN ALSO EXPLANTED AND REPLACED THE IPG DURING THE SAME PROCEDURE. POST-OPERATIVELY, STIMULATION THERAPY WAS RESTORED AND THE SYSTEM COULD GO INTO MRI MODE. Manufacturer Narrative: DATE OF EVENT IS ESTIMATED. DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE PATIENT INFORMATION. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725713&pc=MGB | 3010252479-2021-00156 | 2021-09-28 | Injury | ESSENTIAL MEDICAL, INC | 2021-10-31 | MGB | MANTA 18FR | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P180025 |  | 1 | Event Description: AS REPORTED: PHYSICIAN REPORTED THAT THERE WAS A SUSPECTED COMPLICATION WITH THE 18FR MANTA THAT WAS USED ON A PATIENT. ACCESS WAS GAINED ON THE RCFA UNDER NORMAL PROTOCOL WITH ULTRASOUND AND MICRO-PUNCTURE. A PRE ANGIO WAS TAKEN SHOWING PROPER SHEATH PLACEMENT AND ADEQUATE FLOW OF CONTRAST DOWN THE VESSEL. THE DEPTH MEASUREMENT FOR THE MANTA WAS PERFORMED PER IFU. THE LARGE BORE PROCEDURE WAS THEN PERFORMED WITH AN 18FR SHEATH PLUS VALVE. ONCE THE PROCEDURE WAS COMPLETED 40 UNITS OF PROTAMINE WAS GIVEN. THE MANTA SHEATH WAS THEN ADVANCED INTO THE VESSEL, DILATER WAS REMOVED AND THE MANTA DEVICE WAS ADVANCED OVER THE WIRE PER IFU. THE MANTA WAS DEPLOYED BY PHYSICIAN ACHIEVING HEMOSTASIS. THE WIRE WAS REMOVED AND A POST ANGIO WAS TAKEN. DUE TO THE PATIENT HAVING A PROSTHETIC HIP IT WAS DIFFICULT TO PROPERLY VISUALIZE THE STAINLESS STEAL LOCK AND WHERE THE MANTA WAS DEPLOYED. THE ANGIO DID NOT SHOW ANY EXTRAVASATION, BUT FLOW WAS SLIGHTLY SLOWED. PULSES WERE GOOD, GROIN WAS SOFT AND THE PATIENT WAS SENT TO THE RECOVERY AREA. 2 TO 3 HRS LATER THE PATIENT PRESENTED WITH A COLD LEG. THEY WERE TAKEN BACK TO THE HYBRID OR WHERE A CUTDOWN WAS PERFORMED. THE MANTA WAS REMOVED, AND THE VESSEL WAS CLOSED. ACCORDING TO THE STAFF THE VASCULAR SURGEON STATED IT LOOKED LIKE THERE WAS A DISSECTION IN THE RCFA. THE PATIENT TOLERATED THE PROCEDURE WELL AND WAS DISCHARGED THE NEXT DAY. Manufacturer Narrative: DEVICE WILL NOT BE RETURNED. AN INVESTIGATION HAS BEEN OPENED TO REVIEW DEVICE HISTORY RECORD AND RISK DOCUMENTATION. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725714&pc=MGB | 3010252479-2021-00157 | 2021-10-05 | Malfunction | ESSENTIAL MEDICAL, INC | 2021-10-31 | MGB | MANTA 14F | Physical Resistance/Sticking | No Clinical Signs, Symptoms or Conditions | P180025 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725715&pc=MGB | 3010252479-2021-00158 | 2021-10-06 | Injury | ESSENTIAL MEDICAL, INC | 2021-10-31 | MGB | MANTA 18F | Activation, Positioning or SeparationProblem | Insufficient Information | P180025 |  | 1 | Event Description: AS REPORTED: PHYSICIAN ATTEMPTED TO USE MANTA TO BAILOUT A COMPETITOR DEVICE FAILURE. WHEN THE MANTA WAS UNSUCCESSFUL A COVERED STENT WAS USED. PATIENT IS REPORTED AS DOING OK. Manufacturer Narrative: DEVICE WILL NOT RETURNED. AN INVESTIGATION HAS BEEN OPENED TO REVIEW DEVICE HISTORY RECORD AND RISK DOCUMENTATION. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725200&pc=FRN | 2016493-2021-68027 | 2021-07-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725201&pc=FRN | 2016493-2021-68006 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725202&pc=FRN | 2016493-2021-67993 | 2021-10-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD LVP BEZEL POST RECALL. DURING SERVICING, REPLACED BEZEL ASSEMBLY, DIM U4 ON DISPLAY BOARD. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725203&pc=FRN | 2016493-2021-68026 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725204&pc=FRN | 2016493-2021-68022 | 2021-09-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725205&pc=DZE | 3003184527-2021-01522 | 2021-06-01 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L8.0 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, SMOKER, INADEQUATE BONE QUALITY/QUANTITY, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725206&pc=FRN | 2016493-2021-68019 | 2021-07-15 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. DURING SERVICING, THEY CHECK FOR ALL RECALLS AND IT WAS "DISCOVERED" THAT THE DEVICE HAD "CASE REAR DAMAGE AND IUI CORROSION. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725207&pc=FRN | 2016493-2021-68004 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725208&pc=FRN | 2016493-2021-68024 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725209&pc=FRN | 2016493-2021-68021 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725210&pc=DRM | 3010617000-2021-00992 | 2021-10-04 | Malfunction | ZOLL CIRCULATION | 2021-10-30 | DRM | AUTOPULSE PLATFORM RESUSCITATION | Visual Prompts will not Clear | No Clinical Signs, Symptoms or Conditions | K112998 |  | 1 | Event Description: THE AUTOPULSE PLATFORM SN (B)(4) WAS USED TO RESUSCITATE A PATIENT IN CARDIAC ARREST. THE CREW TURNED ON THE AUTOPULSE PLATFORM, PULLED UP ON THE LIFEBAND (LOT # UNKNOWN), AND PRESSED THE GREEN START/CONTINUE BUTTON TO START CHEST COMPRESSIONS. HOWEVER, THE LIFEBAND WAS NOT RETRACTED, AND THE PLATFORM DISPLAYED A FAULT CODE 42 (FORCE OVERLOAD TRIPPED) ERROR MESSAGE AFTER THE FIRST COMPRESSION FOLLOWED BY THE SAME ERROR AFTER THE FIRST COMPRESSION DURING THE SECOND ATTEMPT. PER THE REPORTER, THE LIFEBAND BAND "CAME OFF", THE END OF LINER ON THE BAND WAS FORCED FROM THE LOCKING TAPE ATTACHED TO THE HINGED BELT GUARD ON THE LIFEBAND. MANUAL CPR WAS PERFORMED DURING THE TROUBLESHOOTING. AFTER REPLACING THE LIFEBAND, THE AUTOPULSE PLATFORM WAS FUNCTIONAL. NO CONSEQUENCES OR IMPACT TO PATIENT. PLEASE SEE THE FOLLOWING RELATED MFR REPORTS: MFR 3010617000-2021-00993 FOR AUTOPULSE LIFEBAND. Manufacturer Narrative: THE REPORTED COMPLAINT OF THE AUTOPULSE PLATFORM SN (B)(4) DISPLAYED FAULT CODE "42" (FORCE OVERLOAD TRIPPED) ERROR MESSAGE WAS NOT CONFIRMED DURING THE INITIAL FUNCTIONAL TESTING BUT CONFIRMED DURING ARCHIVE DATA REVIEW. NO DEVICE MALFUNCTION WAS OBSERVED DURING THE TESTING, AND THE AUTOPULSE PLATFORM WORKED AS INTENDED. DURING VISUAL INSPECTION OF THE AUTOPULSE PLATFORM, THE FRONT ENCLOSURE WAS OBSERVED CRACKED AT THE FRONT END AREA. THIS OBSERVATION IS UNRELATED TO THE REPORTED COMPLAINT AND APPEARED TO BE THE CHARACTERISTICS OF MISHANDLING SUCH AS A DROP. THE FRONT ENCLOSURE WAS REPLACED TO ADDRESS THE OBSERVED ISSUE. ADDITIONALLY, NOTED THE ENCODER DRIVE SHAFT DOES NOT ROTATE SMOOTHLY, EXHIBITS BINDING AND RESISTANCE DUE TO A STICKY CLUTCH PLATE, UNRELATED TO THE REPORTED COMPLAINT. THE CAUSE FOR THE STICKY CLUTCH COULD BE DUE TO THE NORMAL USE OF THE DEVICE. THE IMPACT OF A STICKY CLUTCH WAS NOT SEVERE ENOUGH TO MAKE THE PLATFORM NON-FUNCTIONAL. DEBURRING WAS PERFORMED TO REMEDY THE STICKY CLUTCH PLATE. THE AUTOPULSE PLATFORM PASSED THE INITIAL FUNCTIONAL TESTING WITHOUT ANY FAULT OR ERROR. LOAD CELL CHARACTERIZATION TEST WAS PERFORMED AND CONFIRMED BOTH LOAD CELL MODULES ARE FUNCTIONING WITHIN THE SPECIFICATION. A REVIEW OF THE AUTOPULSE PLATFORM ARCHIVE WAS PERFORMED, AND DATA SHOWED FAULT CODE "42" (FORCE OVERLOAD TRIPPED) ERROR MESSAGE OCCURRED, THUS CONFIRMING THE REPORTED COMPLAINT. THE FAULT CODE 42 ERROR MESSAGE INDICATES THE MOTOR WAS ON FOR TOO LONG DURING ACTIVE OPERATION AND THE LOAD FORCE MONITORING CIRCUIT DETECTED TOO MUCH FORCE APPLIED ON THE LOAD PLATE(S) CHECK BEFORE THE INITIATE TAKE-UP WAS PERFORMED. AS THE CUSTOMER REPORTED, THE LIFEBAND WAS NOT RETRACTED, DUE TO ONE SIDE OF THE LIFEBAND BEING TWISTED AND BECOMING JAMMED IN THE ROLLER/SKIRT AREA, THIS WOULD HAVE RESULTED IN THE CUSTOMER REPORTING FAULT CODE 42 ERROR. FOLLOWING SERVICE, THE AUTOPULSE PLATFORM PASSED THE RUN-IN TEST USING THE 95% PATIENT LARGE RESUSCITATION TEST FIXTURE (LRTF) WITH GOOD KNOWN TEST BATTERIES UNTIL DISCHARGED WITHOUT ANY FAULT OR ERROR. THE AUTOPULSE PLATFORM PASSED THE FINAL TESTING WITHOUT ANY FAULT OR ERROR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725211&pc=FRN | 2016493-2021-68036 | 2021-08-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725212&pc=DZE | 3003184527-2021-01519 | 2021-07-10 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | CONTACT MC INICELL, PF 4.5, E3.5, L11.0 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, IMPLANT ACHIEVED OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725213&pc=DZE | 3003184527-2021-01515 | 2021-07-21 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, PAIN, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725214&pc=JJE | 9612296-2021-00381 | 2021-10-05 | Malfunction | BECKMAN COULTER | 2021-10-30 | JJE | DXC 700 AU CLINICAL CHEMISTRY ANALYZER | High Test Results | No Clinical Signs, Symptoms or Conditions | K161837 |  | 1 | Event Description: THE CUSTOMER REPORTED ERRONEOUS HIGH ISE (ION SELECTIVE ELECTROLYTE) PATIENT RESULTS WERE GENERATED ON THE DXC 700 AU CLINICAL CHEMISTRY ANALYZER. THE ERRONEOUS RESULTS WERE NOT RELEASED FROM THE LABORATORY. THERE WAS NO CHANGE IN PATIENT TREATMENT IN ASSOCIATION WITH THIS EVENT. Manufacturer Narrative: A BECKMAN COULTER FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND EVALUATED THE DEVICE. THE FSE REPLACED ISE INTERNAL REFERENCE SOLUTION OF THE REFERENCE ELECTRODE, ISE MIX MOTOR AND ISE MID SOLUTION WHICH RESOLVED THE ISSUE. BECKMAN COULTER INTERNAL IDENTIFIER IS CASE-(B)(4). DATE OF BIRTH: PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. SEX:PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. WEIGHT: PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED. ETHNICITY: PATIENT DEMOGRAPHIC INFORMATION WAS NOT PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725215&pc=MEA | 2016493-2021-68011 | 2021-07-08 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Failure to Align | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8120 PCA WAS AFFECTED BY SYRINGE SIZER MISALIGN RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725216&pc=QCB | 2023826-2021-04179 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.50/018, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PER UPDATED INFORMATION AN EXCHANGED WAS CARRIED OUT ON (B)(6) 2021. THE REPORTER STATED THAT THE SURGEON HAS CONFIRMED THAT THE PATIENT IS NOW DOING WELL. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725217&pc=FRN | 1314492-2021-04213 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | Device Difficult to Program or Calibrate | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP WOULD NOT GO PAST 257 MILLILITERS. THIS OCCURRED DURING AN UNSPECIFIED PROCESS STEP. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725218&pc=FRN | 2016493-2021-68005 | 2021-04-26 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING, REPLACED CORRODED IUI. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725219&pc=FRN | 2016493-2021-68031 | 2021-07-01 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725220&pc=FRN | 2016493-2021-68032 | 2021-07-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725221&pc=FRN | 1314492-2021-04214 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | False Alarm | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP ALARMED AIR IN LINE DURING TESTING AT THE BIOMED SERVICE DEPARTMENT. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED CONDITION. AIR IN LINE ALARMS WERE NOT REPRODUCED DURING FUNCTIONAL TESTING. A REVIEW OF THE EVENT HISTORY LOG REVEALED MULTIPLE 'AIR IN LINE!' ALARMS. A SERVICE HISTORY REVIEW REVEALED NO INDICATION THAT THE PARTS REPLACED DURING SERVICING CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE REPORTED CONDITION WAS VERIFIED. THE CAUSE OF THE CONDITION WAS DETERMINED TO BE AN OUT OF CALIBRATION ULTRASONIC SENSOR. THE ULTRASONIC SENSOR REQUIRES REPLACEMENT TO ADDRESS THIS ISSUE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725222&pc=DZE | 3003184527-2021-01521 | 2021-06-21 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L9.5 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, IMPLANT ACHIEVED OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, INADEQUATE BONE QUALITY/QUANTITY, BONE RESORPTION, PERI-IMPLANTITIS, INFLAMMATION, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725223&pc=FRN | 2016493-2021-68028 | 2021-10-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725224&pc=KYB | 1119421-2021-02038 |  | Malfunction | ALCON RESEARCH, LLC - HUNTINGTON | 2021-10-30 | KYB | ACRYSOF IQ SINGLEPIECE IOL WITH ULTRASERT DELIVERY SYSTEM | Device Damaged by Another Device | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930014 |  | 1 | Event Description: THIS REPORT SUMMARIZES MALFUNCTION REPORTS OF INTRAOCULAR LENS (IOL) DAMAGED BY DELIVERY SYSTEM. THE REPORTED PATIENT AGES RANGE FROM UNKNOWN TO 83 YEARS. THERE WERE 2 MALE PATIENTS, 1 MALE PATIENT AND 11 UNKNOWN GENDER. Manufacturer Narrative: TWO SAMPLES WERE RETURNED. TWELVE SAMPLES WERE NOT RETURNED. THERE WERE TEN CASES WITH METHOD CODES OF DEVICE NOT RETURNED (4114) ANALYSIS OF PRODUCTION RECORDS (3331), A RESULTS CODE OF NO FINDINGS AVAILABLE (3221), AND A CONCLUSION CODE OF CAUSE NOT ESTABLISHED (4315). THERE WERE TWO CASES WITH METHOD CODES OF TESTING OF ACTUAL/SUSPECTED DEVICE (10) ANALYSIS OF PRODUCTION RECORDS (3331), A RESULTS CODE OF OPERATIONAL PROBLEM IDENTIFIED (114), AND A CONCLUSION CODE OF CAUSE NOT ESTABLISHED (4315). THERE WAS ONE CASE WITH METHOD CODES OF ANALYSIS OF PRODUCTION RECORDS (3331) INCOMPLETE DEVICE RETURNED (4116), A RESULTS CODE OF OPERATIONAL PROBLEM IDENTIFIED (114) AND A CONCLUSION CODE OF CAUSE CANNOT BE TRACED TO DEVICE (4310) THERE WAS ONE CASE WITH A METHOD CODE OF TESTING OF ACTUAL/SUSPECTED DEVICE (10), A RESULTS CODE OF INAPPROPRIATE MATERIAL (202), AND A CONCLUSION CODE OF FAILURE TO FOLLOW INSTRUCTIONS (18). THERE WERE TWO CASES WITH A METHOD CODES OF TESTING OF ANALYSIS OF DATA PROVIDED BY USER/THIRD PARTY (4112) ANALYSIS OF PRODUCTION RECORDS (3331) DEVICE NOT RETURNED (4114), A RESULTS CODE OF NO FINDINGS AVAILABLE (3221), AND A CONCLUSION CODE OF CAUSE NOT ESTABLISHED (4315). THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725225&pc=FRN | 2016493-2021-68034 | 2021-06-29 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Corroded; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING, REPLACED REAR CASE ASSEMBLY, CRACKED DOOR ASSEMBLY WITH KEYPAD, REPLACED CRACKED BEZEL ASSEMBLY, AIL ASSEMBLY, PATIENT SIDE PRESSURE SENSOR, MEMBRANE FRAME DISCOLORED, LEFT AND RIGHT IUI CONTAMINATED. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725226&pc=FRN | 2016493-2021-68033 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY SOFTWARE RECALL AND KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725227&pc=FRN | 2016493-2021-68045 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY BEZEL POST RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725228&pc=DZE | 3003184527-2021-01523 | 2021-09-13 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L8.0 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725229&pc=FRN | 2016493-2021-68025 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725230&pc=HQL | 1119421-2021-02039 |  | Malfunction | ALCON RESEARCH, LLC - HUNTINGTON | 2021-10-30 | HQL | ACRYSOF IQ NATURAL SINGLEPIECE IOL | Scratched Material | No Clinical Signs, Symptoms or Conditions | P930014 |  | 1 | Event Description: THIS REPORT SUMMARIZES MALFUNCTION REPORTS OF SCRATCHED INTRAOCULAR LENSES (IOL). THE REPORTED PATIENT AGES RANGE FROM UNKNOWN TO 74 YEARS. THERE WERE 2 FEMALE PATIENTS, 1 MALE PATIENT AND 6 UNKNOWN GENDER. Manufacturer Narrative: ONE SAMPLE WAS RETURNED. EIGHT SAMPLES WERE NOT RETURNED. THERE WERE EIGHT CASES WITH A METHOD CODE OF DEVICE NOT RETURNED (4114), ANALYSIS OF PRODUCTION RECORDS (3331), A RESULTS CODE OF NO FINDINGS AVAILABLE (3221), AND A CONCLUSION CODE OF CAUSE NOT ESTABLISHED (4315). THERE WAS ONE CASE WITH A METHOD CODE OF TESTING OF ACTUAL/SUSPECTED DEVICE (10) ANALYSIS OF PRODUCTION RECORDS (3331), A RESULTS CODE OF OPERATIONAL PROBLEM IDENTIFIED (114) AND A CONCLUSION CODE OF CAUSE TRACED TO USER (19). THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725231&pc=FRN | 2016493-2021-68017 | 2021-05-19 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725232&pc=FRN | 1314492-2021-04212 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | Device Difficult to Program or Calibrate | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP WOULD NOT GO PAST 261.9 MILLILITERS. THIS OCCURRED DURING AN UNSPECIFIED PROCESS STEP. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725233&pc=DZE | 3003184527-2021-01524 | 2021-09-17 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT MC INICELL, PF 3.5, E 3.5, L14.0 | Failure to Osseointegrate | Inadequate Osseointegration | K093615 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WASN'T COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, IMMEDIATE IMPLANTATION, PERI-IMPLANTITIS, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725234&pc=FRN | 2016493-2021-68052 | 2021-10-08 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE WON'T TURN ON/OFF. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725235&pc=FRN | 2016493-2021-68037 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725236&pc=FRN | 2016493-2021-68038 | 2021-04-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725237&pc=DZE | 3003184527-2021-01525 | 2021-09-21 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.0, E 4.0, L11.0 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, INFECTION, SWELLING, FISTULA, INFLAMMATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725238&pc=MTA | 2023826-2021-04178 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725239&pc=FRN | 2016493-2021-68042 | 2021-05-18 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725240&pc=FRN | 2016493-2021-68056 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725241&pc=FRN | 2016493-2021-68041 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725242&pc=FRN | 2016493-2021-68043 | 2021-07-19 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD BEZEL POST RECALL. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725243&pc=FRN | 2016493-2021-68049 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725244&pc=MEA | 2016493-2021-68018 | 2021-07-12 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8120 PCA WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725245&pc=FRN | 2016493-2021-68062 | 2021-07-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack; Failure to Analyze Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYRINGE PUMP MODULE WAS AFFECTED BY PLASTIC RECALL AND FORCE SENSOR. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725246&pc=FRN | 1314492-2021-04211 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | Failure to Sense | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP KEYPAD DID NOT WORK DURING AN UNSPECIFIED PROCESS STEP. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725247&pc=FRN | 2016493-2021-68047 | 2021-07-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725248&pc=DZE | 3003184527-2021-01526 | 2021-09-14 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.5, E 4.2, L9.5 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, IMPLANT ACHIEVED OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, SMOKER, BONE RESORPTION, INFECTION, FISTULA, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725249&pc=FRN | 1314492-2021-04215 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | False Alarm | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP ALARMED DOWNSTREAM OCCLUSION DURING AN UNSPECIFIED PROCESS STEP. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED CONDITION. A DOWNSTREAM OCCLUSION ALARM WAS NOT REPRODUCED DURING FUNCTIONAL TESTING. A REVIEW OF THE EVENT HISTORY LOG REVEALED DOWNSTREAM OCCLUSION MESSAGES. A SERVICE HISTORY REVIEW WAS PERFORMED AND REVEALED THAT THE DEVICE HAS NO PREVIOUS SERVICE EVENTS; THEREFORE, SERVICING DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. THE REPORTED CONDITION WAS VERIFIED. THE CAUSE OF THE CONDITION WAS DETERMINED TO BE AN OUT OF SPECIFICATION DOWNSTREAM SENSOR. THE FORCE SENSOR REQUIRES REPLACEMENT TO ADDRESS THIS ISSUE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725250&pc=FRN | 2016493-2021-68029 | 2021-07-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725251&pc=FRN | 2016493-2021-68069 | 2021-06-01 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725252&pc=QCB | 2023826-2021-04180 | 2021-08-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/5.0/143, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS / SYMPTOMS) THE REPORTER STATED "STILL HAS EDEMA AND STITCH DUE TO WOUND LEAKAGE INTRA-OP. AC IS NOT SHALLOW ANYMORE. BUT IT WILL RECOVER SOON." CAUSE OF EVENT WAS REPORTED TO BE "PATIENT-RELATED FACTOR" AND WAS REPORTED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "PATIENT HAS NARROW SULCUS. WTW IS 100% CORRECT BUT THE SULCUS IS NARROW." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725253&pc=FRN | 2016493-2021-68051 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725254&pc=FRN | 2016493-2021-68057 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725255&pc=FRN | 2016493-2021-68050 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD LVP BEZEL / KEYPAD RECALL: BAD BEZEL OR KEYPAD ONLY DO NO REPAIRS. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725256&pc=FRN | 1314492-2021-04210 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | Protective Measures Problem | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP ALARMED SYSTEM ERROR 345 (THERMISTOR DISPARITY) DURING AN UNSPECIFIED PROCESS STEP. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED CONDITION. A SYSTEM ERROR 345 DID NOT OCCUR DURING FUNCTIONAL TESTING. A REVIEW OF THE EVENT HISTORY LOG REVEALED SYSTEM ERROR 345 MESSAGES. A SERVICE HISTORY REVIEW REVEALED NO ISSUES THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THE REPORTED CONDITION WAS VERIFIED. THE CAUSE OF THE CONDITION WAS DETERMINED TO BE A FAILING UPSTREAM THERMISTOR. THE ULTRASONIC SENSOR REQUIRES REPLACEMENT TO ADDRESS THIS ISSUE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725257&pc=FRN | 2016493-2021-68039 | 2021-04-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725258&pc=DZE | 3003184527-2021-01527 | 2021-09-20 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 3.5, E 3.5, L11.0 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WASN'T COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, SMOKER, IMMEDIATE IMPLANTATION, INFECTION, PAIN, SWELLING, INFLAMMATION, IMMEDIATE LOADING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725259&pc=FRN | 2016493-2021-68059 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725260&pc=FRN | 2016493-2021-68066 | 2021-10-05 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Device Alarm System | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLINICIAN OPENED THE DOOR DUE TO AN ¿AIR-IN-LINE¿ ALARM. IT WAS NOTED THAT THE PUMP MODULE PHYSICALLY DISCONNECTED FROM THE PCU AND FELL OFF AFTER THE DOOR WAS OPENED, WHICH CAUSED SEPARATION OF THE TUBING THAT WAS INFUSING TOTAL PARENTERAL NUTRITION (TPN). IT IS BELIEVED THAT THE FORCE AND WEIGHT OF THE PUMP MODULE FALLING ON THE FLOOR SNAPPED THE TUBING AT THE SILICONE SEGMENT. THE CLINICIAN WAS UNSURE AS TO WHAT ATTRIBUTED TO THE MALFUNCTION OF THE PUMP. PATIENT WAS UNABLE TO RECEIVE THE REMAINDER OF THE SOLUTION AS TPNS ARE ONLY PREPARED ONCE A DAY. HOWEVER, THE PATIENT RECEIVED CUSTOM FLUIDS OVERNIGHT, WHICH WAS MADE BY THE PHARMACIST STAT. THERE WAS PATIENT INVOLVEMENT BUT NO HARM. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. DEVICE WAS NOT RETURNED TO MANUFACTURING FACILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725261&pc=FRN | 2016493-2021-68060 | 2021-03-25 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725262&pc=FRN | 2016493-2021-68044 | 2021-07-20 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725263&pc=FRN | 2016493-2021-68061 | 2021-05-17 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725264&pc=DZE | 3003184527-2021-01529 | 2021-09-20 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.5, E 4.2, L8.0 | Failure to Osseointegrate | Inadequate Osseointegration |  |  | 1 | Event Description: IMPLANT DIDN'T ACHIEVE OSSEOINTEGRATION, INFECTION, SWELLING, ABSCESS, INFLAMMATION, MOBILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725265&pc=FRN | 2016493-2021-68070 | 2021-07-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Align | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYR WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725266&pc=FRN | 2016493-2021-68055 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725267&pc=FRN | 2016493-2021-68035 | 2021-08-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725268&pc=DZE | 3003184527-2021-01530 | 2021-09-22 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.5, E 4.2, L9.5 | Failure to Osseointegrate | Inadequate Osseointegration | K090154 |  | 1 | Event Description: PRIMARY STABILITY NOT ACHIEVED, NO THREAD TAP USED, SMOKER, INADEQUATE BONE QUALITY/QUANTITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725269&pc=FRN | 2016493-2021-68064 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725270&pc=FRN | 2016493-2021-68073 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL, LEFT AND RIGHT IUI CONTAMINATED. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725271&pc=FRN | 2016493-2021-68054 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Peeled/Delaminated | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725272&pc=FRN | 2016493-2021-68040 | 2021-10-08 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD FAILED TO TURN ON/OFF. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725273&pc=FRN | 2016493-2021-68046 | 2021-07-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725274&pc=FRN | 2016493-2021-68065 | 2021-07-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Device Misassembled During Manufacturing /Shipping | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICE, REPLACED DOOR LATCH, FRONT DOOR, IUI RIGHT SIDE, IUI LEFT SIDE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725275&pc=FRN | 2016493-2021-68075 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725276&pc=QCB | 2023826-2021-04182 | 2021-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | NA |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/5.0/122, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT. FOR STATUS OF THE EYE (SIGNS / SYMPTOMS) THE REPORTER STATED "STILL HAS EDEMA AND STITCH DUE TO WOUND LEAKAGE INTRA-OP. AC IS NOT SHALLOW ANYMORE. BUT IT WILL RECOVER SOON." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725277&pc=MEA | 2016493-2021-68058 | 2021-07-12 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Failure to Align | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8120 PCA WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725278&pc=FRN | 2016493-2021-68071 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Misassembled | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 PCA WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725279&pc=FRN | 2016493-2021-68078 | 2021-05-10 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack; Peeled/Delaminated | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725280&pc=HQL | 1119421-2021-02040 |  | Malfunction | ALCON RESEARCH, LLC - HUNTINGTON | 2021-10-30 | HQL | ACRYSOF IQ NATURAL SINGLEPIECE IOL | Defective Device | No Clinical Signs, Symptoms or Conditions | P930014 |  | 1 | Event Description: THIS REPORT SUMMARIZES MALFUNCTION REPORTS OF DEFECTIVE INTRAOCULAR LENSES (IOL). THE REPORTED PATIENT AGES RANGE FROM UNKNOWN TO 74 YEARS. THERE WERE 2 FEMALE PATIENTS, 1 MALE PATIENT AND 9 UNKNOWN GENDER. Manufacturer Narrative: TWELVE SAMPLES WERE NOT RETURNED. THERE WERE TWELVE CASES WITH A METHOD CODE OF DEVICE NOT RETURNED (4114) ANALYSIS OF PRODUCTION RECORDS (3331), A RESULTS CODE OF NO FINDINGS AVAILABLE (3221) AND A CONCLUSION CODE OF CAUSE NOT ESTABLISHED (4315). THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725281&pc=FRN | 2016493-2021-68085 | 2021-07-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725282&pc=FRN | 2016493-2021-68081 | 2021-08-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD PM / CALIBRATION ISSUE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725283&pc=FRN | 2016493-2021-68079 | 2021-10-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Align | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYR WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725284&pc=FRN | 2016493-2021-68048 | 2021-10-20 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725285&pc=FRN | 2016493-2021-68076 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725286&pc=FRN | 2016493-2021-68093 | 2021-07-29 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725287&pc=FRN | 2016493-2021-68072 | 2021-05-05 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Contamination; Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING IT WAS CONFIRMED THAT THE DEVICE HAD REPLACED CONTAMINATED AIL ASSY AND IUI. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725288&pc=FRN | 2016493-2021-68077 | 2021-06-22 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725289&pc=DZE | 3003184527-2021-01528 | 2021-09-22 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 5.0, E 5.0, L11.0 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, IMPLANT ACHIEVED OSSEOINTEGRATION, PRIMARY STABILITY WAS ACHIEVED, IMPLANT WAS COMPLETELY COVERED WITH BONE, NO THREAD TAP USED, ILLNESS REQUIRING STEROIDS, BIOMECHANICAL OVERLOAD, PERI-IMPLANTITIS, PAIN, INFLAMMATION, HIGH TORQUE WHEN INSERTING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725290&pc=FRN | 2016493-2021-68087 | 2021-06-25 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Naturally Worn | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE SYRINGE WAS NOT MOVING FREELY BECAUSE OF ITS STIFFNESS. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725291&pc=FRN | 2016493-2021-68097 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725292&pc=FRN | 2016493-2021-68063 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725293&pc=FRN | 2016493-2021-68086 | 2021-03-25 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725294&pc=FRN | 2016493-2021-68089 | 2021-10-14 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725295&pc=FRN | 2016493-2021-68053 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725296&pc=FRN | 2016493-2021-68082 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Circuit Failure | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725297&pc=FRN | 2016493-2021-68091 | 2021-08-10 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL, REPLACED BEZEL ASSEMBLY, AIR IN LINE, FRONT DOOR, IUI RIGHT SIDE, IUI LEFT SIDE, MEMBRANE FRAME AND DOOR COVER ISSUE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725298&pc=FRN | 2016493-2021-68104 | 2021-08-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD BROKEN DAMAGED ISSUE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725299&pc=FRN | 2016493-2021-68092 | 2021-05-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725300&pc=FET | 8010047-2021-13852 |  | Malfunction | OLYMPUS MEDICAL SYSTEMS CORP. | 2021-10-30 | FET | HD AUTOCLAVABLE CAMERA HEAD | No Display/Image | No Clinical Signs, Symptoms or Conditions | K955404 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725301&pc=FRN | 2016493-2021-68067 | 2021-04-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725302&pc=FRN | 2016493-2021-68068 | 2021-07-19 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725303&pc=FRN | 2016493-2021-68080 | 2021-08-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725304&pc=FRN | 2016493-2021-68074 | 2021-07-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725305&pc=FRN | 2016493-2021-68098 | 2021-08-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY DIM SEGMENT RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725306&pc=FRN | 2016493-2021-68109 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725307&pc=FRN | 2016493-2021-68106 | 2021-08-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD PREVENTIVE MAINTENANCE UPDATE AND REPLACED SPRING LATCH DUE TO CRACK THAT DEVELOPED. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725308&pc=FRN | 2016493-2021-68102 | 2021-03-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Circuit Failure; Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725309&pc=FET | 8010047-2021-13851 |  | Malfunction | OLYMPUS MEDICAL SYSTEMS CORP. | 2021-10-30 | FET | 4K CAMERA HEAD | No Display/Image | No Clinical Signs, Symptoms or Conditions | K172817 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725310&pc=FRN | 2016493-2021-68099 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725311&pc=FRN | 2016493-2021-68114 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725312&pc=FRN | 2016493-2021-68101 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725313&pc=FRN | 2016493-2021-68100 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725314&pc=FRN | 2016493-2021-68084 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING, IT WAS FOUND AIL SENSOR ASSY: CRACKED HOUSING. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725315&pc=FRN | 2016493-2021-68095 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725316&pc=FRN | 2016493-2021-68113 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725317&pc=FRN | 2016493-2021-68121 | 2021-08-16 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Output Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYRINGE MODULE WAS AFFECTED BY SYRINGE SPLIT NUT RECALL. DURING SERVICING, REPLACED CRACK REAR CASE AND CALIBRATED. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725318&pc=FRN | 2016493-2021-68107 | 2021-09-14 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Circuit Failure | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725319&pc=FRN | 2016493-2021-68083 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725320&pc=FRN | 2016493-2021-68096 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725321&pc=FRN | 2016493-2021-68115 | 2021-07-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Failure to Align; Naturally Worn | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYR WAS AFFECTED BY SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725322&pc=FRN | 2016493-2021-68112 | 2021-06-08 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725323&pc=FRN | 2016493-2021-68088 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725324&pc=FRN | 2016493-2021-68094 | 2021-05-24 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725325&pc=FRN | 2016493-2021-68105 | 2021-06-28 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read; Failure to Align | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYRINGE MODULE WAS AFFECTED BY MISALIGNED SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725326&pc=FRN | 2016493-2021-68120 | 2021-08-10 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725327&pc=FRN | 2016493-2021-68119 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL, R118 REAR CASE AND REPLACED MEMBRANE FRAME DISCOLORED, LEFT AND RIGHT IUI CONTAMINATED. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725328&pc=FRN | 2016493-2021-68090 | 2021-08-26 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725329&pc=FRN | 2016493-2021-68127 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725330&pc=FRN | 2016493-2021-68136 | 2021-09-16 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack; Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY DIM SEGMENT RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725331&pc=FRN | 2016493-2021-68125 | 2021-05-20 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725332&pc=FRN | 2016493-2021-68131 | 2021-08-09 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD FAILED IN PREVENTIVE MAINTENANCE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725333&pc=FRN | 2016493-2021-68141 | 2021-06-24 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725334&pc=CCK | 2016493-2021-68137 | 2021-10-13 | Malfunction | CAREFUSION SD | 2021-10-30 | CCK | ALARIS SYSTEM | Loose or Intermittent Connection | No Clinical Signs, Symptoms or Conditions | K031741 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD ALARM AND ERROR CODE MESSAGES 571.6241. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725335&pc=FRN | 2016493-2021-68122 | 2021-04-08 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY KEYPAD RECALL AND DEVICE HAD LED DISPLAY CHANNEL ERROR 211.2000 OR DIM SEGMENT ISSUE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725336&pc=FRN | 2016493-2021-68142 | 2021-07-26 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE NEEDS TUBING SENSOR. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725337&pc=FRN | 2016493-2021-68111 | 2021-07-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725338&pc=FRN | 2016493-2021-68103 | 2021-06-22 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725339&pc=FRN | 2016493-2021-68130 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725340&pc=QBJ | 3004753838-2021-252730 | 2021-10-11 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT AN ERROR MESSAGE INSTEAD OF WARMUP OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED AS SIGNAL LOSS GREATER THAN AN HOUR WAS CONFIRMED. THE SENSOR WAS INSERTED INTO THE ABDOMEN ON (B)(6) 2021. THE PROBABLE CAUSE WAS THE PATIENT CLOSING THE MOBILE APP WHICH LED TO SIGNAL LOSS. THE REPORTED EVENT OF AN ERROR MESSAGE INSTEAD OF WARMUP IS REPORTABLE BASED ON THE FINDING OF SIGNAL LOSS GREATER THAN AN HOUR. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725341&pc=FRN | 2016493-2021-68116 | 2021-09-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Electrical /Electronic Property Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725342&pc=FRN | 2016493-2021-68132 | 2021-08-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725343&pc=FRN | 2016493-2021-68148 | 2021-10-15 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725344&pc=FRN | 2016493-2021-68118 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725345&pc=FRN | 2016493-2021-68146 | 2021-10-20 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725346&pc=DTB | 2649622-2021-21652 | 2021-10-11 | Injury | MPRI | 2021-10-30 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Capturing Problem; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS ASYNCHRONOUS PACING AND ON DEVICE INTERROGATION IT WAS NOTED THAT THE RIGHT ATRIAL (RA) LEAD HAD DISLODGED INTO THE INFERIOR VENA CAVA. THE RA LEAD WAS CAPPED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: G151 CRT-D, IMPLANTED ON (B)(6) 2019. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725347&pc=FRN | 2016493-2021-68135 | 2021-08-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725348&pc=FRN | 2016493-2021-68156 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725349&pc=FRN | 2016493-2021-68124 | 2021-03-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725350&pc=FRN | 2016493-2021-68155 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725351&pc=FRN | 2016493-2021-68110 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725352&pc=FRN | 2016493-2021-68147 | 2021-10-13 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD FAILED PREVENTIVE MAINTENANCE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725353&pc=FRN | 2016493-2021-68108 | 2021-07-28 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 PCA WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725354&pc=FRN | 2016493-2021-68134 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725355&pc=FRN | 2016493-2021-68117 | 2021-08-12 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE WAS AFFECTED BY LVP BEZEL POST RECALL AND CRACKED HOUSING ON AIL SENSOR ASSEMBLY. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725356&pc=FRN | 2016493-2021-68153 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725357&pc=DTB | 2649622-2021-21650 | 2021-10-11 | Injury | MPRI | 2021-10-30 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Failure to Capture; Fracture; High impedance | Bradycardia; Fatigue | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS EXPERIENCING FATIGUE AND SEVERE BRADYCARDIA. ON DEVICE INTERROGATION IT WAS NOTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NO CAPTURE AT MAXIMUM OUTPUT AND HIGH AND UNDEFINED IMPEDANCE. THE RV LEAD WAS SUSPECTED OF A FRACTURE. THE RV LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725358&pc=FRN | 2016493-2021-68140 | 2021-07-19 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725359&pc=FRN | 2016493-2021-68123 | 2021-08-31 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICE, REPLACED REAR CASE DUE TO CRACKED LOWER LEFT CORNER, REPLACED LEFT HANDLE DUE TO LATERAL CRACKS, REPLACED RIGHT HANDLE DUE TO CRACKED REAR PORTION, REPLACED WORN OUT LEAF SPRING, REPLACED BARREL CLAMP DUE TO CRACKED INNER PORTION OF HANDLE. REPLACED CRACKED FLANGE GRIPPER RETAINER, REPLACED CONTAMINATED CARRIAGE BLOCK, REPLACED CONTAMINATED TUBE DRIVE ARM, REPLACED WORN OUT TUBE DRIVE ARM SLEEVE, REPLACED WORN OUT WIPER SEAL, REPLACED RIGHT IUI. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725360&pc=FRN | 2016493-2021-68143 | 2021-04-26 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL AND DAMAGED IUI. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725361&pc=FRN | 2016493-2021-68129 | 2021-05-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725362&pc=DTB | 2649622-2021-21653 |  | Injury | MPRI | 2021-10-30 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Fracture; High impedance | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED HIGH AND UNDEFINED IMPEDANCE. LEAD FRACTURE WAS SUSPECTED. IT WAS NOTED THAT THE PATIENT HAD PACEMAKER SYNDROME. THE RA LEAD WAS CAPPED AND REPLACED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725363&pc=FRN | 2016493-2021-68126 | 2021-04-13 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725364&pc=DTB | 2649622-2021-21651 |  | Injury | MPRI | 2021-10-30 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture; Device Dislodged or Dislocated; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT X-RAY CONFIRMED RIGHT ATRIAL (RA) LEAD DISLODGEMENT AT A DEVICE UPGRADE PROCEDURE. IT WAS NOTED THAT THE LEAD EXHIBITED HIGH THRESHOLDS WITH NO CAPTURE AT HIGH OUTPUT. THE RA LEAD WAS CAPPED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725365&pc=FRN | 2016493-2021-68138 | 2021-08-10 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725366&pc=FRN | 2016493-2021-68133 | 2021-08-05 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725367&pc=FRN | 2016493-2021-68149 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725368&pc=FRN | 2016493-2021-68139 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING IT WAS CONFIRMED THAT THE DEVICE HAD BROKEN BEZEL, DAMAGED DOOR ASSEMBLY WITH KEYPAD AND BROKEN AIR IN LINE. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725369&pc=FRN | 2016493-2021-68128 | 2021-07-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725370&pc=FRN | 2016493-2021-68154 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725371&pc=FRN | 2016493-2021-68145 | 2021-08-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725372&pc=FRN | 2016493-2021-68144 | 2021-06-08 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725373&pc=MEA | 2016493-2021-68150 | 2021-06-28 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Failure to Align | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8120 PCA WAS AFFECTED BY MISALIGNED SYRINGE SIZER RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725374&pc=FRN | 2016493-2021-68152 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725375&pc=MEA | 2016493-2021-68161 | 2021-06-22 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Contamination; Corroded; Crack; Device Markings/Labelling Problem | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8120 PCA WAS AFFECTED BY PLASTIC RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725376&pc=FRN | 2016493-2021-68151 | 2021-07-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Break; Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725377&pc=FRN | 2016493-2021-68162 | 2021-10-19 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. DURING SERVICING, REPLACED UNRESPONSIVE KEYPAD. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725378&pc=FRN | 2016493-2021-68164 | 2021-07-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725379&pc=FRN | 2016493-2021-68159 | 2021-06-30 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725380&pc=FRN | 2016493-2021-68163 | 2021-07-15 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725381&pc=FRN | 2016493-2021-68160 | 2021-05-21 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725382&pc=FRN | 2016493-2021-68165 | 2021-07-01 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725383&pc=MEA | 2016493-2021-68158 | 2021-10-27 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Device Markings/Labelling Problem | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD BROKEN/DAMAGED AND PLUNGER ASSEMBLY IS NOT MOVING SMOOTHLY. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725384&pc=FRN | 2016493-2021-68157 | 2021-10-27 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Crack; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD ALARM: ERROR CODES / MESSAGES. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725385&pc=MEA | 2016493-2021-68175 | 2021-09-14 | Malfunction | CAREFUSION SD | 2021-10-30 | MEA | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K043299 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725386&pc=FRN | 2016493-2021-68174 | 2021-07-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Corroded | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING, REPLACED BROKEN RARE CASE, BROKEN FRONT CASE, CORRODED CARRIAGE BLOCK, BENT TUBE DRIVE, BROKEN BARREL CLAMP, CORRODED LEFT AND RIGHT IUIS. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725387&pc=FRN | 2016493-2021-68172 | 2021-07-01 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725388&pc=FRN | 2016493-2021-68168 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725389&pc=FRN | 2016493-2021-68167 | 2021-06-02 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725390&pc=FRN | 2016493-2021-68178 | 2021-04-23 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Naturally Worn | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8110 SYR WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING, REPLACED CARRIAGE BLOCK DUE TO FAILED PLUNGER FORCE ACCURACY TEST. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725391&pc=FRN | 2016493-2021-68176 | 2021-10-11 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Display Difficult to Read | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725392&pc=FRN | 2016493-2021-68171 | 2021-08-05 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725393&pc=FRN | 2016493-2021-68170 | 2021-06-04 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725394&pc=FRN | 2016493-2021-68166 | 2021-08-06 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725395&pc=FRN | 2016493-2021-68169 | 2021-06-03 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725396&pc=FRN | 2016493-2021-68173 | 2021-05-17 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725397&pc=HQC | 2028159-2021-01255 | 2021-10-05 | Malfunction | ALCON RESEARCH, LLC - IRVINE TECHNOLOGY CENTER | 2021-10-30 | HQC | LAUREATE WORLD PHACO SYSTEM | Unexpected Shutdown | No Clinical Signs, Symptoms or Conditions | K021566 |  | 1 | Event Description: A PHYSICIAN REPORTED A MACHINE SWITCHED OFF ON ITS OWN. NO PATIENT HARM WAS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE EVENT OCCURRED DURING THE PHACOEMULSIFICATION (PHACO) PHASE DURING A CATARACT PROCEDURE. THE PROCEDURE WAS COMPLETED BY SHIFTING TO ANOTHER MACHINE. THERE WAS NO PATIENT IMPACT. Manufacturer Narrative: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. THE MANUFACTURER INTERNAL REFERENCE NUMBER IS: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725398&pc=FRN | 2016493-2021-68177 | 2021-09-24 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Charging Problem; Computer Operating System Problem | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A FAILURE WAS OBSERVED DURING A PLANNED PREVENTIVE MAINTENANCE OR RECALL REMEDIATION SERVICE EVENT. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725399&pc=FRN | 2016493-2021-68180 | 2021-07-01 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Circuit Failure; Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE PCA/SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725400&pc=FRN | 2016493-2021-68179 | 2021-04-22 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination; Crack | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8100 LVP WAS AFFECTED BY PLASTIC RECALL. DURING SERVICING IT WAS CONFIRMED THAT THE DEVICE HAD DAMAGED BOTH IUI. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725401&pc=FEB | 2150060-2021-00025 | 2021-10-01 | Injury | MEDIVATORS INC | 2021-10-30 | FEB | RAPICIDE PA HIGH LEVEL DISINFECTANT | Adverse Event Without Identified Device or Use Problem | Erythema | K082988 |  | 1 | Event Description: A USER FACILITY REPORTED THAT RAPICIDE PA HIGH-LEVEL DISINFECTANT PART A SPLASHED ON AN EMPLOYEE'S FOREARM. THE EMPLOYEE EXPERIENCED REDNESS AND RAISING OF THE SKIN. THE EMPLOYEE WASHED THE AFFECTED AREA. THE EMPLOYEE WAS REPORTED TO BE WEARING GLOVES, BUT IT IS UNKNOWN IF OTHER PERSONAL PROTECTIVE EQUIPMENT (PPE) WAS WORN AT THE TIME OF THE EXPOSURE. NO ADDITIONAL REPORTS OF INJURY WERE REPORTED. Manufacturer Narrative: IT IS UNKNOWN HOW THE DISINFECTANT CAME INTO CONTACT WITH THE EMPLOYEE'S SKIN. THE DURATION OF THE REPORTED SKIN REDNESS AND IRRITATION WAS NOT PROVIDED. FOLLOWING THE EVENT, THE USER CONTACTED CHEMTREC AND WAS PROVIDED THE RAPICIDE PA PART A SDS. THE EMPLOYEE SOUGHT MEDICAL ATTENTION AND A BANDAGE TO THE AFFECTED AREA WAS APPLIED. NO FURTHER MEDICAL INTERVENTION WAS REQUIRED. IT IS NOT KNOWN IF THE USER WAS WEARING PROPER PPE AT THE TIME OF THE REPORTED EVENT. THE RAPICIDE PA PART A SDS PROVIDES THE FOLLOWING WARNING LANGUAGE AND GUIDANCE, RELATED TO SKIN EXPOSURE AND PPE USAGE: "CAUSES SEVERE SKIN BURNS AND EYE DAMAGE. WASH HANDS THOROUGHLY AFTER HANDLING. WEAR PROTECTIVE GLOVES/PROTECTIVE CLOTHING/EYE PROTECTION/FACE PROTECTION...IF ON SKIN (OR HAIR): RINSE SKIN WITH WATER/SHOWER." NO ADDITIONAL REPORTS OF ADVERSE EVENT OR HARM WERE RECEIVED. MEDIVATORS WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS TO ENSURE THE PRODUCT CONTINUES TO PERFORM AS EXPECTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725402&pc=FRN | 2016493-2021-68182 | 2021-05-17 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725403&pc=CBI | 3003898360-2021-00987 | 2021-10-07 | Malfunction | TELEFLEX MEDICAL | 2021-10-30 | CBI | HUDSON ET TUBE, SHER-I-BRONCH, LS, 37 FR | Inflation Problem | No Clinical Signs, Symptoms or Conditions | K180253 |  | 1 | Event Description: IT WAS REPORTED THAT "THE DISTAL CUFF COULD ONLY BE FILLED MARGINALLY, WHICH WAS CHECKED BRONCHOSCOPICALLY. THE PILOT BALLOON DID NOT REFLECT THE INFLATION STATUS, NOR DID THE CUFF PRESSURE MEASUREMENT REFLECT THE INFLATION STATUS". NO PATIENT HARM OR INJURY REPORTED. PATIENT CONDITION REPORTED AS "FINE". Manufacturer Narrative: (B)(4). COMPLAINT VERIFICATION TESTING COULD NOT BE PERFORMED AS NO SAMPLE WAS RETURNED FOR ANALYSIS. FROM THE PICTURE ATTACHED OF 5-16037 ET TUBE, SHER-I-BRONCH, LS, 37 FR, IT IS UNABLE TO BE CONFIRMED THE FAILURE MODE REPORTED AS "INFLATION ISSUE - OTHER". HOWEVER, IT IS NECESSARY TO RECEIVE THE PHYSICAL SAMPLE TO PERFORM A PROPER INVESTIGATION, DETERMINE ROOT CAUSE & IMPLEMENT CORRECTIVE ACTIONS. PART NUMBER REPORTED IS NOT BEING MANUFACTURED CURRENTLY, HOWEVER, ANOTHER PART NUMBER FROM THE SAME FAMILY WAS USED FOR THE "VERIFICATION OF FAILURE MODE REPORTED IN THE CURRENT MANUFACTURING PROCESS" (PART # 00267-35 LOT # 3142026) AND WAS CONDUCTED AS FOLLOWS: 80 SAMPLES WERE TAKEN FROM THE CURRENT PRODUCTION; THE SAMPLES WERE VISUALLY INSPECTED AND ISSUE REPORTED "INFLATION ISSUE - OTHER" WAS NOT OBSERVED IN THE CURRENT MANUFACTURING PROCESS. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED, AND NO RELEVANT FINDINGS WERE IDENTIFIED. WITHOUT THE DEVICE TO EVALUATE THE COMPLAINT COULD NOT BE CONFIRMED AND THE PROBABLE CAUSE COULD NOT BE DETERMINED F ROM THE AVAILABLE INFORMATION. TELEFLEX WILL CONTINUE TO MONITOR AND TREND FOR REPORTS OF THIS NATURE. IF THE SAMPLE BECOMES AVAILABLE AT A LATER DATE A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725404&pc=FRN | 2016493-2021-68183 | 2021-10-13 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725405&pc=QBJ | 3004753838-2021-252731 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725406&pc=PQF | 3004753838-2021-252732 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | PQF | DEXCOM G5 MOBILE CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | P120005 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS DETERMINED TO BE A LOW TRANSMITTER BATTERY. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725407&pc=QBJ | 3004753838-2021-252733 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725408&pc=QBJ | 3004753838-2021-252734 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725409&pc=QBJ | 3004753838-2021-252735 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725410&pc=DTB | 2649622-2021-21654 | 2021-10-13 | Injury | MPRI | 2021-10-30 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Fracture; High impedance; Over-Sensing; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED HIGH AND UNDEFINED IMPEDANCE, HIGH THRESHOLDS AND OVER-SENSING. LEAD FRACTURE WAS SUSPECTED. THE RA LEAD WAS CAPPED AND REPLACED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: CONCOMITANT PRODUCTS: DDMC3D1 ICD IMPLANTED ON (B)(6) 2020. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725411&pc=FRN | 2016493-2021-68181 | 2021-10-07 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Failure to Read Input Signal | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT AN 8015 PCU WAS AFFECTED BY KEYPAD RECALL. THERE WAS NO REPORTED PATIENT INVOLVEMENT. Manufacturer Narrative: PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725412&pc=QBJ | 3004753838-2021-252736 | 2021-10-07 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT EARLY SENSOR EXPIRATION OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725413&pc=QBJ | 3004753838-2021-252737 | 2021-10-26 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Premature End-of-Life Indicator | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725414&pc=QBJ | 3004753838-2021-252738 | 2021-10-30 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Premature End-of-Life Indicator | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT EARLY SENSOR EXPIRATION OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725415&pc=QBJ | 3004753838-2021-252739 | 2021-10-28 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Premature End-of-Life Indicator | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT EARLY SENSOR EXPIRATION OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS DETERMINED TO BE POTENTIAL PATIENT MISUSE. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725416&pc=FRN | 2016493-2021-68184 | 2021-05-24 | Malfunction | CAREFUSION SD | 2021-10-30 | FRN | ALARIS SYSTEM | Contamination | No Clinical Signs, Symptoms or Conditions | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE SYR IUI BOARDS WHICH HAD FLUX RESIDUE AND/OR EXPOSED COPPER WHICH REQUIRED REWORK. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: A DEVICE HISTORY RECORD REVIEW IS PERFORMED ON EACH DEVICE REPORTED IN A MDR REPORTABLE EVENT ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED WAS OBTAINED FROM SERVICING ACTIVITIES PERFORMED ON THE DEVICE. THERE WERE NO ADDITIONAL DETAILS OBTAINABLE OR PROVIDED AT THE TIME OF SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725417&pc=DZE | 3003184527-2021-01531 | 2021-09-27 | Injury | THOMMEN MEDICAL AG | 2021-10-30 | DZE | ELEMENT RC INICELL, PF 4.5, E 4.2, L9.5 | Loss of Osseointegration | Inadequate Osseointegration | K090154 |  | 1 | Event Description: LOSS OF OSSEOINTEGRATION, BISPHOSPHONATE TREATMENT, PERI-IMPLANTITIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725418&pc=QBJ | 3004753838-2021-252740 | 2021-10-28 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Premature End-of-Life Indicator | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT EARLY SENSOR EXPIRATION OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED. THE PROBABLE CAUSE WAS DETERMINED TO BE LOW COUNTS ABERRATION VIA DATA. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725419&pc=QBJ | 3004753838-2021-252741 | 2021-10-10 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT AN ERROR MESSAGE CONNECTING AND PAIRING THE TRANSMITTER OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED AS SIGNAL LOSS GREATER THAN AN HOUR WAS CONFIRMED. THE SENSOR WAS INSERTED INTO AN UNKNOWN INSERTION SITE ON AN UNKNOWN DATE. THE PROBABLE CAUSE WAS THE PATIENT CLOSING THE MOBILE APP WHICH LED TO SIGNAL LOSS. THE REPORTED EVENT OF AN ERROR MESSAGE CONNECTING AND PAIRING THE TRANSMITTER IS REPORTABLE BASED ON THE FINDING OF SIGNAL LOSS GREATER THAN AN HOUR. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725421&pc=DTZ | 8010762-2021-00580 | 2021-10-25 | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | DTZ | HLS SET ADVANCED | Improper Chemical Reaction | No Clinical Signs, Symptoms or Conditions | K112360 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725422&pc=DTZ | 8010762-2021-00578 |  | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | DTZ | QUADROX-ID ADULT O.FILT | Improper Chemical Reaction | Low Oxygen Saturation | K150267 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THE OXYGENATION OF THE QUADROX-ID WAS INSUFFICIENT. COMPLAINT ID: (B)(4). Manufacturer Narrative: FURTHER PATIENT AND PERFUSION DATA WAS REQUESTED BUT NOT YET RECEIVED. A FOLLOW-UP EMDR WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725423&pc=DTQ | 8010762-2021-00583 | 2021-10-25 | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | DTQ | HEART LUNG MACHINE | Failure to Pump | No Clinical Signs, Symptoms or Conditions | K943803 |  | 1 | Event Description: IT WAS REPORTED THAT THE PUMPS ARE SHOWING "HEAD ERROR" . COMPLAINT ID # (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725424&pc=DTQ | 8010762-2021-00581 | 2021-10-19 | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | DTQ | HEART LUNG MACHINE | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | K943803 |  | 1 | Event Description: IT WAS REPORTED THAT THE HL20 DISPLAYED THE ERROR MESSAGE: "STOP--". THE PUMP STOPPED. NO PATIENT INVOLVEMENT. REFERENCE NUMBER: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT THE HL20 DISPLAYED THE ERROR MESSAGE: "STOP--". THE PUMP STOPPED. NO PATIENT INVOLVED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WILL BE SENT ONSITE FOR INVESTIGATION OF THE DEVICE. AS SOON AS NEW INFORMATION BECOMES AVAILABLE, A FOLLOW UP MEDWATCH WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725425&pc=DTZ | 8010762-2021-00579 | 2021-10-25 | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | DTZ | HLS SET ADVANCED | Improper Chemical Reaction | No Clinical Signs, Symptoms or Conditions | K112360 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725426&pc=KFM | 8010762-2021-00582 | 2021-10-27 | Malfunction | MAQUET CARDIOPULMONARY GMBH | 2021-10-30 | KFM | HEART LUNG MACHINE | Filtration Problem | No Clinical Signs, Symptoms or Conditions | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT THE ERROR MESSAGE "FAULTBUB" WAS DISPLAYED ON THE ROTAFLOW CONSOLE CONTINUOUSLY. NO MORE INFORMATION PROVIDED. NO INDICATION OF ACTUAL OR POTENTIAL FOR HARM OR DEATH HAS BEEN REPORTED. COMPLAINT ID:(B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725427&pc=MKJ | 1220908-2021-03749 |  | Injury | ZOLL MEDICAL CORPORATION | 2021-10-30 | MKJ | R SERIES DEFIBRILLATOR | Self-Activation or Keying; Unintended Electrical Shock | Insufficient Information | K060559/P160 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE SELF-DISCHARGED AND A CLINICIAN AND A CRNA BOTH RECEIVED AN UNINTENDED DELIVERY OF ENERGY. COMPLAINANT INDICATED THAT THERE WAS NO ADVERSE EFFECT TO THE PATIENT, THE CLINICIAN PROCEEDED TO THE EMERGENCY DEPARTMENT, WAS DEEMED OK AND IT WAS NOT KNOWN IF THERE WAS ANY ADVERSE EFFECT TO THE CRNA DUE TO THE REPORTED MALFUNCTION. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725428&pc=BTO | 3012307300-2021-10480 | 2021-09-01 | Malfunction | SMITHS MEDICAL INTERNATIONAL, LTD. | 2021-10-30 | BTO | PORTEX | Leak/Splash | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT TWELVE DAYS AFTER THE CUSTOMER STARTED TO USE THE PRODUCT, LEAKAGE OF AIR WAS OBSERVED. SO HE CHANGED THE PRODUCT TO ANOTHER NEW ONE. THE CUSTOMER THEN CHECKED THE REMOVED PRODUCT AND FOUND A TEAR IN THE INFLATION LINE, WHICH WAS SUSPECTED TO BE THE ROOT CAUSE OF THE EVENT. NO PATIENT INJURY. Manufacturer Narrative: RETURNED DEVICE WAS RECEIVED FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE THE CUSTOMER REPORTED CONDITION WAS CONFIRMED. PROBLEM SOURCE WAS TRACED TO USER INTERFACE. THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THAT THIS DEVICE MET ALL MANUFACTURING SPECIFICATION FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725429&pc=BTL | 3012307300-2021-10481 | 2021-09-19 | Malfunction | ST PAUL | 2021-10-30 | BTL | PNEUPAC | Failure to Cycle | No Clinical Signs, Symptoms or Conditions | K020899 |  | 1 | Event Description: IT WAS REPORTED THAT A SMITHS MEDICAL VENTILATOR WAS NOT STOPPING, CYCLING INTERMITTENTLY. NO PATIENT INVOLVEMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725430&pc=BTL | 3012307300-2021-10482 |  | Malfunction | ST PAUL | 2021-10-30 | BTL | PNEUPAC | Improper Flow or Infusion; Pressure Problem | Insufficient Information | K020899 |  | 1 | Event Description: IT WAS REPORTED THAT A SMITHS MEDICAL VENTILATOR HAD LOW PRESSURE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: RETURNED DEVICE WAS RECEIVED FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE THE CUSTOMER REPORTED CONDITION WAS NOT CONFIRMED. NO FAULT FOUND. THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THAT THIS DEVICE MET ALL MANUFACTURING SPECIFICATION FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725431&pc=CCK | 3012307300-2021-10483 |  | Malfunction | ST PAUL | 2021-10-30 | CCK | BCI | Device Alarm System | Insufficient Information | K970209 |  | 1 | Event Description: WHILE CALIBRATING IT GAVE AN OCCLUSION ERROR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725432&pc=ILQ | 1000136418-2021-00002 | 2021-08-09 | Malfunction | FORBES AAC | 2021-10-30 | ILQ | SOUNDPOD | Temperature Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725433&pc=FWM | 1645337-2021-12109 | 2021-01-01 | Injury | MENTOR TEXAS | 2021-10-30 | FWM | MENTOR SMOOTH ROUND MODERATE PLUS PROFILE | Material Rupture | Deformity/ Disfigurement | P990075 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE PATIENT WHO HAD A MENTOR SMOOTH ROUND MODERATE PLUS PROFILE 250CC SALINE PROSTHESIS PLACED EXPERIENCED UNILATERAL RIGHT SIDED DEFLATION POST PROCEDURE. AT THE TIME OF THIS REPORT, MENTOR HAS RECEIVED NO INFORMATION REGARDING EXPLANTATION OR AN EXPECTED EXPLANTATION DATE. Manufacturer Narrative: AT THE TIME OF THIS REPORT, MENTOR HAS RECEIVED NO INFORMATION REGARDING EXPLANTATION OR AN EXPECTED EXPLANTATION DATE. IT IS UNKNOWN AT THIS TIME IF THE DEVICE WILL BE MADE AVAILABLE FOR RETURN. AS A RESULT, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS CAN BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. IF THE COMPLAINT DEVICE IS RECEIVED IN THE FUTURE, THE INVESTIGATION WILL BE REOPENED AND CONDUCTED AS APPROPRIATE. SINCE NO LOT NUMBER WAS PROVIDED, NO MANUFACTURING RECORD EVALUATION REVIEW COULD BE PERFORMED. REASON FOR DEVICE EXPLANT AND/OR REOPERATION: N/A. MANUFACTURER¿S REFERENCE NUMBER: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725434&pc=LWP | 2017865-2021-35544 |  | Injury | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2021-10-30 | LWP | ENDURITY MRI PACEMAKER, SR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE IS SUBJECT TO THE ASSURITY AND ENDURITY PACEMAKERS HEADER ANOMALY ADVISORY ISSUED BY ABBOTT ON 15 MARCH 2021. THE PHYSICIAN CLAIMED THE DEVICE WAS EXPLANTED AND REPLACED PREVENTATIVELY. THE PATIENT WAS STABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725435&pc=GDW | 3005075853-2021-06565 | 2021-10-07 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | PROXIMATE*RL VASCULAR LIN STAP | Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON HAD TAKEN THE LINEAR STAPLER AND FOLLOWED THE FIRING SEQUENCE. HE COULD NOT GET THE STAPLER FORMATION DUE TO THE MALFORMATION OF THE STAPLES. HE TOOK ANOTHER NEW ONE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). BATCH #: UNK. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? NO. COULD YOU PLEASE CLARIFY IF WAS ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725436&pc=QBJ | 3004753838-2021-252742 | 2021-10-04 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725437&pc=QBJ | 3004753838-2021-252743 | 2021-10-01 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725438&pc=QBJ | 3004753838-2021-252744 | 2021-09-30 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725439&pc=QBJ | 3004753838-2021-252745 | 2021-10-01 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. NO PRODUCT OR DATA WAS PROVIDED FOR EVALUATION. CONFIRMATION OF THE ALLEGATION AND A PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725440&pc=QBJ | 3004753838-2021-252746 | 2021-10-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Device Output | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT TRANSMITTER FAILED ERROR OCCURRED. DATA WAS EVALUATED AND THE ALLEGATION WAS NOT CONFIRMED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725442&pc=GDW | 3005075853-2021-06566 | 2021-10-07 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | 75 MM NTLC SELECTABLE RELOAD | Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K092577 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLE HAS COME OUT AFTER THE FIRING. NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). BATCH #: UNK. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE PROVIDE MORE DETAILS FOR "STAPLE HAS COME OUT"? WHAT WAS THE SHAPE OF THE STAPLES (B-FORMATION, IRREGULAR, LEGS STRAIGHT)? WERE THE STAPLES DEPLOYED INTO THE TISSUE? WERE THE STAPLES SEEN IN THE SURGICAL FIELD? COULD YOU PLEASE CLARIFY IF THE RELOAD FALLS AFTER THE FIRING? COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? COULD YOU PLEASE CLARIFY IF WAS ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725443&pc=MNT | 2031642-2021-05422 | 2021-10-11 | Malfunction | RESPIRONICS CALIFORNIA, LLC | 2021-10-30 | MNT | RESPIRONICS | Inadequate User Interface | No Clinical Signs, Symptoms or Conditions | K102985 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT A VENTILATOR EXPERIENCED A TOUCHSCREEN ISSUE. THE DEVICE WAS REPORTED TO BE IN USE. THERE WAS NO PATIENT HARM OR DELAY NOTED. THE PHILIPS FIELD SERVICE ENGINEER (FSE) CONFIRMED AND DUPLICATED THE REPORTED ISSUE. A REPLACEMENT TOUCHSCREEN WAS ORDERED. FOLLOWING THE EVALUATION OF THE DEVICE, THE FSE REPLACED THE TOUCHSCREEN TO RESOLVE THE PROBLEM. THE UNIT WAS THEN FUNCTIONALLY TESTED AND SUCCESSFULLY PASSED SPECIFIED TESTS. NO OTHER ABNORMALITY WAS OBSERVED. Manufacturer Narrative: (B)(6).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725444&pc=QBJ | 3004753838-2021-252747 | 2021-09-05 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SENSOR STOPPED WORKING OCCURRED. THE SENSOR WAS INSERTED INTO THE ABDOMEN ON (B)(6) 2021. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED AS SIGNAL LOSS GREATER THAN AN HOUR WAS CONFIRMED. THE PROBABLE CAUSE WAS DETERMINED TO BE SIGNAL LOSS DUE TO THE TRANSMITTER AND APP WERE NOT BEING ABLE TO ESTABLISH A CONNECTION. THE REPORTED EVENT OF A SENSOR STOPPED WORKING IS REPORTABLE BASED ON THE FINDING OF SIGNAL LOSS GREATER THAN AN HOUR. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: COM-(B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725445&pc=GDW | 3005075853-2021-06567 | 2021-10-06 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | PROXIMATE*PPH PROCEDURE SET | Failure to Cut | No Clinical Signs, Symptoms or Conditions | K051301 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE FIRED DURING THE HEMORRHOIDECTOMY DID NOT FIRE PLASTIC TO PLASTIC. IT STOPPED BEFORE THE BLADE COULD TOUCH THE ANVIL. STAPLES FIRED COMPLETELY BUT THE BLADE DID NOT PROCEED TILL THE END AND STOPPED BEFORE THE END. DEVICE REMOVED AND CHECKED FOR LEAKS AND HAEMOSTASIS. SURGEON CHECKED FOR STAPLE FORMATION AND DID MULTIPLE LEAK TEST TO CHECK WHETHER THERE WERE ANY LEAKS IN THE PATIENTS CAUSED BY THE DEVICE. SURGERY WAS DELAYED 60 MINUTES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. Manufacturer Narrative: (B)(4). BATCH #: UNK. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT CONSEQUENCE AS A RESULT OF THE PROCEDURE BEING PROLONGED? COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? COULD YOU PLEASE CLARIFY IF WAS ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? ANSWER- ADDITIONAL HOSPITAL STAY AND AN ADDITIONAL PROCEDURE TO EXTRACT THE DONUTS PATIENT WAS IN MILD HYPOTENSION IN THE FIRST 24 HOURS. NEEDED ADDITIONAL SURGERY AND ADDITIONAL HOSPITAL STAY
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725446&pc=FZP | 3005075853-2021-06568 | 2021-10-06 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | FZP | LIGACLIP*ENDO ROTATING MCA | Failure to Form Staple | No Clinical Signs, Symptoms or Conditions | K864102 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PERFORMING A GASTRIC SLEEVE THE SURGEON WAS USING THE ER420 TO SUPPORT THE STAPLE LINE. HE RECOGNIZED THAT THE FIRST CLIPS WERE NOT CLOSED PROPERLY. HE REMOVED THE CLIPS OUT OF THE PATIENT AND OPENED A NEW DEVICE. WITH THE NEW DEVICE THERE WERE NO FURTHER MALFUNCTIONS OF THE CLIPS. PROCEDURE COULD BE FINISHED WITHOUT FURTHER CIRCUMSTANCES. Manufacturer Narrative: (B)(4). BATCH # UNK. THE LOT/ BATCH WAS NOT PROVIDED; THEREFORE, THE MANUFACTURING RECORDS EVALUATION COULD NOT BE PERFORMED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725447&pc=CBK | 2518422-2021-05853 | 2021-09-30 | Malfunction | RESPIRONICS, INC | 2021-10-30 | CBK | TRILOGY 100 | Failure to Power Up | No Clinical Signs, Symptoms or Conditions | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A VENTILATOR WOULD NOT POWER ON. THERE WAS NO HARM OR INJURY REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725448&pc=GDW | 3005075853-2021-06569 | 2021-10-06 | Injury | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | ENDO ILS 25MM, CURVED | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K940967 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725449&pc=GDW | 3005075853-2021-06570 | 2021-08-20 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | BLUE RELOAD FOR ECHELON 60 | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | K051002 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE LOBECTOMY SURGERY, AFTER FIRED, THE STAPLE LINE AND CUT LINE ARE BOTH COMPLETE. NOTED OOZING. TO STOP OOZING WITH COMPRESSION HEMOSTASIS. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). BATCH #: 176A67. (B)(4). INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO SURGERY FOR EVALUATION. VISUAL INSPECTION AND FUNCTIONAL TESTING WERE CONDUCTED ON THE RETURNED RELOAD. VISUAL ANALYSIS OF THE RETURNED SAMPLE REVEALED THAT THE ECR60B RELOAD (D) WAS RETURNED UNFIRED WITH 3 DOUBLE DRIVERS LOOSE, MAKING THE RELOAD NON-FUNCTIONAL. IN ADDITION, THE RELOAD PAN WAS NOTED TO BE PARTIALLY DISLODGED AT THE PROXIMAL END FROM RIGHT SIDE. NO FUNCTIONAL TEST WAS PERFORMED DUE THE CONDITION OF THE RELOAD. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS BATCH NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. AS PART OF ETHICON ENDO SURGERY¿S QUALITY PROCESS ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. NO CONCLUSION COULD BE REACHED ON WHAT MAY HAVE CAUSED THE RELOAD PAN TO DISLODGE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725450&pc=GDW | 3005075853-2021-06571 | 2021-09-27 | Injury | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | FLEX 60 ARTICULATING | Failure to Fire; Difficult to Open or Close | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K081146 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE USE OF THE DEVICE ON THE RECTUM WHICH JAMMED IN CLOSED AND CURVED POSITION WITHOUT THE POSSIBILITY TO OPENING IT, RESTORING ITS STRAIGHT POSITION OR STAPLING. THE PROCEDURE HAD TO BE CONVERT IN LAPAROTOMY. RESECTION OF THE RECTUM BETWEEN 2 LINES OF STAPLES WITH CONTOUR CUTTER. REMOVAL OF THE STILL CURVED DEVICE THROUGH THE HOLE OF THE TROCAR, AFTER REMOVING THE TROCAR. EXTRACTION OF THE RECTAL SECTION FROM THE JAWS OF THE DEVICE. THERE WAS NO STAPLING. Manufacturer Narrative: (B)(4). BATCH # UNK. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/ PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WAS THE INITIAL PROCEDURE? WHAT WAS THE INDICATIONS FOR THE PROCEDURE? DID THE ISSUE OCCUR ON FIRST OR SUBSEQUENT FIRING? WHAT WAS THE APPROXIMATE THICKNESS OF TISSUE BEING STAPLED? WHAT ATTEMPT WERE MADE TO GET THE DEVICE DE-ARTICULATED? WHAT TROUBLESHOOTING STEPS WERE FOLLOWED? WAS THE DEVICE FIRED OVER EXISTING STAPLE LINE? WAS THE CONTOUR DEVICE USED TWICE? WHAT WERE THE STEPS FOR USE OF THE EC60 DEVICE? DID THE DEVICE COMPLETE THE FIRING SEQUENCE? HOW WAS THE DEVICE ATTEMPTED TO BE OPENED? WHAT POINT IN THE USAGE DID THE DEVICE JAM? AFTER FIRING, DURING FIRING? HOW WAS THE DEVICE REMOVED FROM THE PATIENT? WHAT COLOR CARTRIDGE WAS BEING USED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725451&pc=GDW | 3005075853-2021-06572 | 2021-09-23 | Malfunction | ETHICON ENDO-SURGERY, LLC. | 2021-10-30 | GDW | PWRD 31MM CURVED CIRCULAR, 18CM SHAFT | Failure to Fire | No Clinical Signs, Symptoms or Conditions | K163523 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE STAPLER WAS CONNECTED, BUT DID NOT FIRE. SURGEON CHANGED BATTERY TO BATTERY TAKEN FROM NEW DEVICE. DEVICE COULD NOW BE FIRED. NO PATIENT CONSEQUENCE. Manufacturer Narrative: (B)(4). BATCH #: UNK. THE LOT/BATCH WAS NOT PROVIDED; THEREFORE, THE MANUFACTURING RECORDS EVALUATION COULD NOT BE PERFORMED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725452&pc=LKK | 3004209178-2021-16168 | 2021-09-01 | Malfunction | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2021-10-30 | LKK | SYNCHROMED II | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM A FOREIGN HEALTHCARE PROVIDER VIA A DISTRIBUTOR REGARDING A PATIENT RECEIVING AN UNKNOWN DOSE AND CONCENTRATION OF GABLOFEN VIA AN IMPLANTABLE PUMP FOR A THORACIC TUMOR. IT WAS REPORTED THERE WAS A TEMPORARY STOP OF THE PUMP ON (B)(6) 2021. THE SEVERITY OF THE EVENT WAS CLASSIFIED AS NON-SERIOUS. THERE WAS A "MOTOR STALL OCCURRED" ERROR DURING THE REFILL. THE PUMP PROGRAM WAS UPDATED SEVERAL TIMES BUT EVEN THOUGH THE PROGRAM WAS SUCCESS THE SAME ERROR APPEARED. AFTER ABOUT TWENTY MINUTES, WHEN TRYING AGAIN, THE ERROR MESSAGE DISAPPEARED AND THE DEVICE RETURNED TO NORMAL. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725454&pc=FRN | 1314492-2021-04216 |  | Malfunction | BAXTER HEALTHCARE CORPORATION | 2021-10-30 | FRN | SPECTRUM INFUSION PUMP | False Alarm | No Clinical Signs, Symptoms or Conditions | K133801 |  | 1 | Event Description: IT WAS REPORTED THAT A SPECTRUM PUMP ALARMED DOWNSTREAM OCCLUSION. THE EVENT HAPPENED DURING TESTING AT THE BIOMED. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED CONDITION. DURING FUNCTIONAL TESTING A DOWNSTREAM OCCLUSION ALARM DI NOT OCCUR. A REVIEW OF THE EVENT HISTORY LOG REVEALED MULTIPLE DOWNSTREAM OCCLUSION ALARM. THE REPORTED CONDITION WAS VERIFIED. THE CAUSE OF THE CONDITION WAS DETERMINED TO BE AN OUT OF CALIBRATION FORCE SENSOR. THE FORCE SENSOR REQUIRES REPLACEMENT TO ADDRESS THIS ISSUE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725455&pc=LWR | 3007113487-2021-00070 |  | Injury | ST. JUDE MEDICAL, INC. | 2021-10-30 | LWR | EPIC STENTED PORCINE HEART VALVE W/FLEXFIT SYSTEM | Backflow; Degraded | Mitral Valve Insufficiency/ Regurgitation | P040021 |  | 1 | Event Description: IT WAS REPORTED THAT AN EPIC STENTED PORCINE HEART VALVE W/FLEXFIT SYSTEM WAS IMPLANTED (B)(6) 2017. DURING A FOLLOW UP VISIT IN THE CLINIC IN 2019, MILD MITRAL TRANSVALVULAR (CENTRAL PART) REGURGITATION WAS CONFIRMED. THE PHYSICIAN HAS REPORTED THAT THE THICKNESS OF THE EPIC LEAFLET SEEMS TO HAVE BECOME THINNER SINCE IT WAS IMPLANTED AND THIS MIGHT HAVE CAUSED THE REGURGITATION. AN EXPLANT OF THE VALVE IS NOT EXPECTED DUE TO THE PATIENTS AGE. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE RESULTS/METHOD AND CONCLUSION CODES ALONG WITH INVESTIGATION RESULTS WILL BE PROVIDED IN THE FINAL REPORT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725457&pc=QBJ | 3004753838-2021-223114 | 2021-10-01 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT THE TRANSMITTER UNPAIRED ITSELF. DATA WAS PROVIDED FOR INVESTIGATION. CONFIRMATION OF THE COMPLAINT WAS UNDETERMINED VIA DATA. THE PROBABLE CAUSE WAS UNABLE TO BE DETERMINED VIA DATA. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725458&pc=FWM | 9617229-2021-54379 | 2021-10-05 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FWM | STYLE 68 SALINE FILLED BREAST IMPLANT | Fluid Leak | Failure of Implant | P990074 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED RIGHT SIDE DEFLATION. DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: DEFLATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725459&pc=FTR | 9617229-2021-56290 | 2012-05-19 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | STYLE 115 SILICONE GEL FILLED BREAST IMPLANT | Malposition of Device | Seroma; Wrinkling | P020056 |  | 1 | Event Description: PATIENT REPORTED LEFT SIDE "RIPPLING AND FEELS LIKE THEY MAY BE LEAKING", "ANXIETY", AND "FLUID". HEALTHCARE PROFESSIONAL REPORTED "RIPPLING", "INCORRECT IMPLANT POSITION", "DOCTOR DID NOT SUSPECT OR CONFIRM A RUPTURE ", "PALPABILITY OF IMPLANT RIPPLES", AND "THEY ARE CLOSE TOGETHER". DEVICE REMAINS IMPLANTED. Manufacturer Narrative: INFORMATION CONTAINED IN THIS REPORT WAS PREVIOUSLY SUBMITTED THROUGH PSR ON 04/23/2013. A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. THE EVENTS OF WRINKLING AND SEROMA ARE PHYSIOLOGICAL COMPLICATIONS AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THESE EVENTS. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: "ANXIETY" AND "FLUID IN LEFT IMPLANT."
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725460&pc=LGW | 3006705815-2021-05374 | 2021-06-28 | Injury | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2021-10-30 | LGW | PROTEGE IPG, 16-CHANNEL RECHARGEABLE | Insufficient Information | Insufficient Information | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 1627487-2021-18011, 1627487-2021-18012. IT WAS REPORTED THAT PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 WHEREIN THE ENTIRE SYSTEM WAS EXPLANTED FOR BEING NONFUNCTIONAL. Manufacturer Narrative: DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE EVENT AND PATIENT WEIGHT. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725461&pc=LGW | 1627487-2021-18012 | 2021-06-28 | Injury | ABBOTT MEDICAL | 2021-10-30 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Insufficient Information | Insufficient Information | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 3006705815-2021-05374, 1627487-2021-18011. IT WAS REPORTED THAT PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 WHEREIN THE ENTIRE SYSTEM WAS EXPLANTED FOR BEING NONFUNCTIONAL. Manufacturer Narrative: DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE EVENT AND PATIENT WEIGHT. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725462&pc=LGW | 1627487-2021-18011 | 2021-06-28 | Injury | ABBOTT MEDICAL | 2021-10-30 | LGW | OCTRODE LEAD KIT, 60CM LENGTH | Insufficient Information | Insufficient Information | P010032 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 3006705815-2021-05374, 1627487-2021-18012. IT WAS REPORTED THAT PATIENT UNDERWENT SURGICAL INTERVENTION ON (B)(6) 2021 WHEREIN THE ENTIRE SYSTEM WAS EXPLANTED FOR BEING NONFUNCTIONAL. Manufacturer Narrative: DURING PROCESSING OF THIS COMPLAINT, ATTEMPTS WERE MADE TO OBTAIN COMPLETE EVENT AND PATIENT WEIGHT. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725463&pc=FTR | 9617229-2021-55590 | 2018-01-01 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | INSPIRA TEXTURED SILICONE GEL FILLED BREAST IMPLANT | Device Appears to Trigger Rejection; Material Rupture | Capsular Contracture | P020056 |  | 1 | Event Description: PATIENT REPORTED LEFT SIDE RUPTURE. HEALTHCARE PROFESSIONAL ADDITIONAL REPORTED "DEFORMITY", "SLIGHT DISCOMFORT", ¿THIS IS THE CAPSULE, [ILLEGIBLE] BUT IT IS CONTRACTED, IT IS COMPRESSED¿ BAKER GRADE UNKNOWN. DEVICE HAS BEEN EXPLANTED AND REPLACED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. THE EVENT OF CAPSULAR CONTRACTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: RUPTURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725464&pc=FWM | 9617229-2021-54578 | 2021-10-06 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FWM | STYLE 68 SALINE FILLED BREAST IMPLANT | Fluid Leak | Failure of Implant | P990074 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED "DEFLATION" AGAINST UNKNOWN SIDE. THE DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: DEFLATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725465&pc=FTR | 9617229-2021-54782 | 2021-10-01 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | STYLE 10 SILICONE GEL FILLED BREAST IMPLANT | Material Rupture | Failure of Implant | P020056 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED LEFT SIDE RUPTURE . DEVICE HAS BEEN EXPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: RUPTURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725466&pc=OLO | 3010266064-2021-00736 | 2021-10-08 | Malfunction | BLUE BELT TECHNOLOGIES | 2021-10-30 | OLO | NAVIO BONE SCREW DRIVER | Solder Joint Fracture | No Clinical Signs, Symptoms or Conditions | K191223 |  | 1 | Event Description: IT WAS REPORTED THAT DURING SET UP FOR A NAVIO ASSISTED TKA SURGERY, THE NAVIO BONE SCREW DRIVER WAS BROKEN AND WOULD NOT HOLD THE NAVIO PINS PROPERLY. THE PROCEDURE WAS COMPLETED, WITHOUT DELAY, USING A S+N BACK-UP DEVICE. PATIENT WAS NOT HARMED AS CONSEQUENCE OF THIS PROBLEM. Manufacturer Narrative: INTERNAL COMPLAINT REFERENCE CASE (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725467&pc=FTR | 9617229-2021-54855 | 2021-07-15 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | STYLE 115 SILICONE GEL FILLED BREAST IMPLANT | Material Rupture | Failure of Implant | P020056 |  | 1 | Event Description: PATIENT REPORTED RIGHT SIDE "RUPTURE". PATIENT ADDITIONALLY REPORTED "LUMP"; THIS IS NOT DEVICE RELATED. THE DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: RUPTURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725468&pc=FTR | 9617229-2021-54548 | 2021-10-06 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | INSPIRA SMOOTH SILICONE GEL FILLED BREAST IMPLANT | Material Rupture | Failure of Implant | P020056 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED RIGHT SIDE RUPTURE. DEVICE EXPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: RUPTURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725469&pc=FWM | 9617229-2021-54619 | 2021-10-03 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FWM | STYLE 68 SALINE FILLED BREAST IMPLANT | Fluid Leak | Failure of Implant | P990074 |  | 1 | Event Description: PATIENT REPORTED A LEFT SIDE DEFLATION. HEALTHCARE PROFESSIONAL CONFIRMED DEFLATION. THE DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: DEFLATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725470&pc=NAD | 8043484-2021-01918 | 2021-07-15 | Injury | SMITH & NEPHEW MEDICAL LTD. | 2021-10-30 | NAD | REPLICARE THIN DRESS 3.5X5.5IN | Adverse Event Without Identified Device or Use Problem | Skin Erosion |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING WOUND TREATMENT WITH REPLICARE THIN DRESS 3.5X5.5IN, ON FOLLOW UP VISIT, SUBSTANTIAL EXCORIATION WAS NOTED ON ONE PATIENT. THE USAGE OF THE DRESSING WAS SUSPENDED IMMEDIATELY. THE TREATMENT WAS COMPLETED USING A VAC THERAPY, AND THE PATIENT RECEIVED STEROIDS AND ANTIFUNGAL ORAL MEDICATIONS TO TREAT THE EXCORIATION. NO OTHER COMPLICATIONS WERE REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725471&pc=NAD | 8043484-2021-01919 | 2021-07-15 | Injury | SMITH & NEPHEW MEDICAL LTD. | 2021-10-30 | NAD | REPLICARE THIN DRESS 3.5X5.5IN | Adverse Event Without Identified Device or Use Problem | Skin Erosion |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING WOUND TREATMENT WITH REPLICARE THIN DRESS 3.5X5.5IN, ON FOLLOW UP VISIT, SUBSTANTIAL EXCORIATION WAS NOTED ON ONE PATIENT. THE USAGE OF THE DRESSING WAS SUSPENDED IMMEDIATELY. THE TREATMENT WAS COMPLETED USING A VAC THERAPY, AND THE PATIENT RECEIVED STEROIDS AND ANTIFUNGAL ORAL MEDICATIONS TO TREAT THE EXCORIATION. NO OTHER COMPLICATIONS WERE REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725472&pc=QBJ | 3004753838-2021-252748 | 2021-10-11 | Malfunction | DEXCOM, INC. | 2021-10-30 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT SENSOR STOPPED WORKING OCCURRED. THE SENSOR WAS INSERTED INTO THE ABDOMEN ON (B)(6) 2021. DATA WAS EVALUATED AND THE ALLEGATION WAS CONFIRMED AS SIGNAL LOSS GREATER THAN AN HOUR WAS CONFIRMED. THE PROBABLE CAUSE WAS DETERMINED TO BE SIGNAL LOSS. THE PROBABLE CAUSE OF THE SIGNAL LOSS COULD NOT BE DETERMINED. THE REPORTED EVENT OF A SENSOR STOPPED WORKING IS REPORTABLE BASED ON THE FINDING OF SIGNAL LOSS GREATER THAN AN HOUR. NO INJURY OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725473&pc=NAY | 2955842-2021-11395 | 2021-06-29 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-30 | NAY | ENDOWRIST | Material Fragmentation; Arcing | No Clinical Signs, Symptoms or Conditions | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725474&pc=FTR | 9617229-2021-54768 | 2020-10-06 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | INSPIRA SMOOTH SILICONE GEL FILLED BREAST IMPLANT | Device Appears to Trigger Rejection | Capsular Contracture | P020056 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED LEFT SIDE CAPSULAR CONTRACTURE, BAKER GRADE IV. DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. THE EVENT OF CAPSULAR CONTRACTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. REASON FOR REOPERATION: CAPSULAR CONTRACTURE, BAKER GRADE IV.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725475&pc=FWM | 9617229-2021-54761 |  | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FWM | STYLE 68 SALINE FILLED BREAST IMPLANT | Fluid Leak; Device Appears to Trigger Rejection | Capsular Contracture | P990074 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED VIA NBIR RIGHT SIDE "CAPSULAR CONTRACTURE BAKER GRADE III", "SUSPECTED/ACTUAL RUPTURE/DEFLATION", AND "CHANGE SHAPE/SIZE/STYLE" . DEVICE HAS BEEN EXPLANTED. Manufacturer Narrative: THE EVENT OF CAPSULAR CONTRACTURE IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. NO CONTACT INFORMATION WAS PROVIDED FOR THE INITIAL REPORTER, THEREFORE ADDITIONAL EVENT, PRODUCT, AND/OR PATIENT DETAILS ARE NOT ATTAINABLE. REASON FOR REOPERATION: "CAPSULAR CONTRACTURE BAKER GRADE III" AND "SUSPECTED/ACTUAL RUPTURE/DEFLATION".
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725476&pc=FTR | 9617229-2021-54960 | 2021-06-15 | Injury | ALLERGAN (COSTA RICA) | 2021-10-30 | FTR | INSPIRA SMOOTH SILICONE GEL FILLED BREAST IMPLANT | Device Appears to Trigger Rejection | Capsular Contracture | P020056 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED LEFT SIDE CAPSULAR CONTRACTURE, BAKER GRADE III. DEVICE REMAINS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED. FURTHER INFORMATION FROM THE REPORTER REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THE EVENT OF CAPSULAR CONTRACTURE BAKER GRADE III IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ALLERGAN IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. THE REASON FOR REOPERATION IS CAPSULAR CONTRACTURE BAKER GRADE III. REOPERATION HAS NOT BEEN SCHEDULED AT THIS TIME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725477&pc=NAY | 2955842-2021-11396 | 2021-09-13 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-30 | NAY | ENDOWRIST;DAVINCI SI | Material Split, Cut or Torn | No Clinical Signs, Symptoms or Conditions | K050369 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725478&pc=QFG | 3013756811-2021-116281 | 2021-10-08 | Malfunction | TANDEM DIABETES CARE | 2021-10-30 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE PUMP STOPS DELIVERING INSULIN ON A DAILY BASIS. THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER'S BLOOD GLUCOSE LEVEL. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO FOLLOW UP BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725479&pc=NAY | 2955842-2021-11397 | 2021-09-03 | Injury | INTUITIVE SURGICAL, INC | 2021-10-30 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Necrosis | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725480&pc=NAY | 2955842-2021-11398 | 2021-08-19 | Injury | INTUITIVE SURGICAL, INC | 2021-10-30 | NAY | ENDOWRIST | Failure to Deliver Energy | Hemorrhage/Bleeding | K131861 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725481&pc=NAY | 2955842-2021-11399 | 2021-09-14 | Malfunction | INTUITIVE SURGICAL, INC | 2021-10-30 | NAY | ENDOWRIST | Material Fragmentation | No Clinical Signs, Symptoms or Conditions | K180351 |  | 1
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725482&pc=QFG | 3013756811-2021-115748 | 2021-10-08 | Malfunction | TANDEM DIABETES CARE | 2021-10-30 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Incorrect Measurement; Battery Problem | No Clinical Signs, Symptoms or Conditions | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE PUMP BATTERY GAUGE WAS FLUCTUATING AND THAT THE PUMP COULD NOT BE CHARGED. CUSTOMER¿S BLOOD GLUCOSE LEVEL WAS 125 MG/DL. THE CUSTOMER REVERTED TO MULTIPLE DOSE INJECTIONS FOR INSULIN THERAPY AND A REPLACEMENT PUMP WAS SENT. Manufacturer Narrative: THE DEVICE HAS BEEN RECEIVED FOR EVALUATION; HOWEVER, DEVICE EVALUATION IS NOT YET COMPLETE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE EVALUATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725483&pc=QFG | 3013756811-2021-116882 | 2021-10-12 | Malfunction | TANDEM DIABETES CARE | 2021-10-30 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE "PUMP WASN'T WORKING PROPERLY". THERE WAS NO REPORTED ADVERSE IMPACT TO THE CUSTOMER'S BLOOD GLUCOSE LEVEL. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO FOLLOW UP BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. DEVICE NOT RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725484&pc=LZG | 3004464228-2021-19602 | 2021-09-10 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Premature Activation | No Clinical Signs, Symptoms or Conditions | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE NEEDLE MECHANISM HAD FULLY DEPLOYED BEFORE THE POD WAS ABLE TO BE USED. THE POD WAS DISCARDED. Manufacturer Narrative: ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725485&pc=LZG | 3004464228-2021-19603 | 2021-10-27 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Premature Activation | No Clinical Signs, Symptoms or Conditions | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE NEEDLE MECHANISM HAD FULLY DEPLOYED BEFORE THE POD WAS ABLE TO BE USED. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725486&pc=LZG | 3004464228-2021-19604 | 2021-10-26 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD INSULIN MANAGEMENT SYSTEM | Premature Activation | No Clinical Signs, Symptoms or Conditions | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE NEEDLE MECHANISM HAD FULLY DEPLOYED BEFORE THE POD WAS ABLE TO BE USED. THE POD WAS DISCARDED. Manufacturer Narrative: ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725487&pc=LZG | 3004464228-2021-19605 | 2021-10-26 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Failure to Fire | Hyperglycemia | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT FOUND CANNULA HAD NOT DEPLOYED AS INTENDED. THE CANNULA WAS NOT VISIBLE AND THE PINK SLIDE DID NOT MOVE FORWARD, INDICATING A NEEDLE MECHANISM FAILURE. THE PATIENT'S BLOOD GLUCOSE (BG) VALUES REACHED 332 MG/DL, WHILE WEARING THE POD BETWEEN 4 AND 24 HOURS ON THE ARM. FOR TREATMENT, THE PATIENT CHANGED OUT THE POD, DRANK WATER AND GAVE CORRECTION BOLUSES. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725490&pc=BSZ | 2112667-2021-02402 | 2021-10-04 | Malfunction | DATEX-OHMEDA, INC. | 2021-10-30 | BSZ | AISYS CS2 | Calibration Problem | No Clinical Signs, Symptoms or Conditions | K172045 |  | 1 | Event Description: THE HOSPITAL REPORTED A MALFUNCTION RESULTING IN THE LOSS OF MECHANICAL VENTILATION. THERE WAS NO REPORT OF PATIENT INJURY. Manufacturer Narrative: A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE SYSTEM AND CONFIRMED THE REPORTED ISSUE. THE DEVICE WAS CALIBRATED TO CORRECT THE REPORTED ISSUE. PATIENT INFORMATION COULD NOT BE OBTAINED DUE TO COUNTRY PRIVACY LAWS. UNIQUE IDENTIFIER: (B)(4). THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. LEGAL MANUFACTURER: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725492&pc=LZG | 3004464228-2021-19606 | 2021-10-28 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Failure to Fire | No Clinical Signs, Symptoms or Conditions | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT FOUND CANNULA HAD NOT DEPLOYED AS INTENDED. THE CANNULA WAS NOT VISIBLE AND THE PINK SLIDE DID NOT MOVE FORWARD, INDICATING A NEEDLE MECHANISM FAILURE. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725493&pc=LZG | 3004464228-2021-19607 | 2021-10-28 | Malfunction | INSULET CORPORATION | 2021-10-30 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Failure to Fire | No Clinical Signs, Symptoms or Conditions | K192659 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT FOUND CANNULA HAD NOT DEPLOYED AS INTENDED. THE CANNULA WAS NOT VISIBLE AND THE PINK SLIDE DID NOT MOVE FORWARD, INDICATING A NEEDLE MECHANISM FAILURE. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO CONFIRM THE REPORTED NEEDLE MECHANISM FAILURE OR TO DETERMINE ITS ROOT CAUSE. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING CONFIRMING THAT THE DEVICE DEPLOYED THE CANNULA CORRECTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725494&pc=LKK | 3004209178-2021-16169 | 2021-08-26 | Malfunction | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2021-10-30 | LKK | SYNCHROMED II | Migration or Expulsion of Device | Pain; Insufficient Information | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (HEALTHCARE PROVIDER, CLINICAL STUDY) REGARDING A PATIENT WHO WAS RECEIVING FENTANYL (500 MCG AT 105.2295 MCG/DAY) AND BUPIVACAINE (5 MG AT 1.0523 MG/DAY) VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT ON (B)(6) 2021 THE PATIENT CALLED IN REPORTING THAT SINCE RF PROCEDURE, THE PATIENT HAD PAIN AT THE PUMP SITE AND THE PUMP APPEARS TO BE PROTRUDING AND MOVED WITHIN THE POCKET. IT WAS NOTED THAT PATIENT STILL REPORTS RELIEF WITH PUMP AND NO EROSION. ON (B)(6) 2021 THE PUMP SITE NERVE BLOCK WAS PERFORMED. IT WAS NOTED THAT IF PAIN PERSISTS, IT WAS RECOMMENDED MOVING LOCATIONS. THE OUTCOME WAS NOTED AS ONGOING. THE CLINICAL DIAGNOSIS WAS PAIN AT THE PUMP SITE. THE ETIOLOGY OF THE EVENT INDICATED THE RELATIONSHIP OF THE EVENT TO THE DEVICE OR THERAPY WAS RELATED AND INDICATED THE RELATIONSHIP OF THE EVENT TO THE IMPLANT PROCEDURE WAS NOT RELATED. THE INCISIONAL SITE/DEVICE TRACT WAS PUMP POCKET. THE EVENT DATE WAS (B)(6) 2021. Manufacturer Narrative: IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725495&pc=EFB | 9614977-2021-00055 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E PRO 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Overheating of Device | Burn(s); No Clinical Signs, Symptoms or Conditions | K972436 |  | 13 | Event Description: THIS REPORT SUMMARIZES 13 MALFUNCTION EVENTS WHERE A MIDWEST E PRO HANDPIECE OVERHEATED. IN 5 EVENTS, THE PATIENT'S EXPERIENCED MINOR BURNS THAT DID NOT REQUIRE INTERVENTION. IN 8 EVENTS, NO INJURY RESULTED. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 13 OF 13 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 13 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICES DID NOT HEAT UP DURING EVALUATION. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725496&pc=EFB | 9614977-2021-00057 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 3 | Event Description: THIS REPORT SUMMARIZES 3 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 3 OF 3 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. EVALUATION OF 2 DEVICES FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICES WERE CLEANED OF DEBRIS AND THE TURBINES WERE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725497&pc=EFB | 9614977-2021-00058 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725498&pc=EFB | 9614977-2021-00053 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION PRO HANDPIECE | Overheating of Device | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST TRADITION PRO HANDPIECE OVERHEATED. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF ONE DEVICE FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICE DID NOT HEAT UP DURING EVALUATION. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725499&pc=EFB | 9614977-2021-00054 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E MINI 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Overheating of Device | Burn(s); No Clinical Signs, Symptoms or Conditions | K972436 |  | 12 | Event Description: THIS REPORT SUMMARIZES 12 MALFUNCTION EVENTS WHERE A MIDWEST E MINI HANDPIECE OVERHEATED. IN 2 EVENTS, THE PATIENT'S EXPERIENCED MINOR BURNS THAT DID NOT REQUIRE INTERVENTION. IN 10 EVENTS, NO INJURY RESULTED. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 12 OF 12 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 8 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICES DID NOT HEAT UP DURING EVALUATION. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 4 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICES DID HEAT UP DURING EVALUATION. THE DEVICES HAD DEBRIS BUILD-UP. THIS DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725500&pc=EFB | 9614977-2021-00056 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E PLUS 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Overheating of Device | Burn(s); Insufficient Information; No Clinical Signs, Symptoms or Conditions | K972436 |  | 55 | Event Description: THIS REPORT SUMMARIZES 55 MALFUNCTION EVENTS WHERE A MIDWEST E PLUS HANDPIECE OVERHEATED. IN 2 EVENTS, THE OUTCOME IS UNKNOWN. IN 12 EVENTS, THE PATIENT'S EXPERIENCED MINOR BURNS THAT DID NOT REQUIRE INTERVENTION. IN 41 EVENTS, NO INJURY RESULTED. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 53 OF 55 DEVICES WERE RETURNED FOR EVALUATION. WE ARE AWAITING THE RETURN OF ONE DEVICE. ONE DEVICE WILL NOT BE RETURNED FOR EVALUATION. EVALUATION OF 41 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICES DID NOT HEAT UP DURING EVALUATION. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 9 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICES DID HEAT UP DURING EVALUATION. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 3 DEVICES FOUND EXCESSIVE WEAR DUE TO A LACK OF PROPER MAINTENANCE AND LUBRICATION. THE DEVICES HAD DEBRIS BUILD-UP. THE DEVICES DID NOT HEAT UP DURING EVALUATION. THE DEVICES WAS REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725501&pc=EFB | 9614977-2021-00059 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725502&pc=EFB | 9614977-2021-00062 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725503&pc=EFB | 9614977-2021-00061 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION PRO HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 2 | Event Description: THIS REPORT SUMMARIZES 2 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION PRO HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 2 OF 2 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725504&pc=EFB | 9614977-2021-00063 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 2 | Event Description: THIS REPORT SUMMARIZES 2 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 2 OF 2 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 2 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725505&pc=EFB | 9614977-2021-00065 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION PRO HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 2 | Event Description: THIS REPORT SUMMARIZES 2 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION PRO HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 2 OF 2 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 2 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725506&pc=EFB | 9614977-2021-00064 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 4 | Event Description: THIS REPORT SUMMARIZES 4 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 3 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725507&pc=EFB | 9614977-2021-00066 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS PLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 3 | Event Description: THIS REPORT SUMMARIZES 3 MALFUNCTION EVENTS WHERE A MIDWEST STYLUS PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 3 OF 3 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 3 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725508&pc=EFB | 9614977-2021-00060 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST TRADITION HANDPIECE | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 23 | Event Description: THIS REPORT SUMMARIZES 23 MALFUNCTION EVENTS WHERE A MIDWEST TRADITION HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 22 OF 23 DEVICES WERE RETURNED FOR EVALUATION. WE ARE AWAITING THE RETURN OF 1 DEVICE. EVALUATION OF 21 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMER. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725509&pc=EFB | 9614977-2021-00070 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS PLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST STYLUS PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725510&pc=EFB | 9614977-2021-00069 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K003518 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST STYLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725511&pc=EFB | 9614977-2021-00068 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST PHOENIX | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 5 | Event Description: THIS REPORT SUMMARIZES 5 MALFUNCTION EVENTS WHERE A MIDWEST PHOENIX HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 4 OF 5 DEVICES WERE RETURNED FOR EVALUATION. 1 DEVICE WILL NOT BE RETURNED FOR EVALUATION. EVALUATION OF 3 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE AND A LACK OF LUBRICATION. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725512&pc=EFB | 9614977-2021-00073 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS PLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 3 | Event Description: THIS REPORT SUMMARIZES 3 MALFUNCTION EVENTS WHERE A MIDWEST STYLUS PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 3 OF 3 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 3 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725513&pc=EFB | 9614977-2021-00072 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS PLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST STYLUS PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND NORMAL CHUCK WEAR AND THE TURBINE NEEDED TO BE REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725514&pc=EFB | 9614977-2021-00071 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS PLUS | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 9 | Event Description: THIS REPORT SUMMARIZES 9 MALFUNCTION EVENTS WHERE A MIDWEST STYLUS PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 9 OF 9 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 10 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725515&pc=EFB | 9614977-2021-00067 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST PHOENIX | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST PHOENIX HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725516&pc=EFB | 9614977-2021-00077 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS ATC MINI | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K003518 |  | 1 | Event Description: THIS REPORT SUMMARIZES 1 MALFUNCTION EVENT WHERE A MIDWEST STYLUS ATC MINI HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED. Manufacturer Narrative: 1 OF 1 DEVICE WAS RETURNED FOR EVALUATION. EVALUATION OF 1 DEVICE FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICE WAS CLEANED OF DEBRIS AND THE TURBINE WAS REPLACED. THE DEVICE WAS REPAIRED AND RETURNED TO THE CUSTOMER. THIS EVENT IS BEING SUBMITTED AS PART OF VMSR. ONLY ONE EVENT WAS RECEIVED FOR THIS DEVICE DURING THIS QUARTER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725517&pc=EFB | 9614977-2021-00076 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E PLUS 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions |  |  | 11 | Event Description: THIS REPORT SUMMARIZES 11 MALFUNCTION EVENTS WHERE A MIDWEST E PLUS HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 10 OF 11 DEVICES WERE RETURNED FOR EVALUATION. 1 DEVICE WILL NOT BE RETURNED FOR EVALUATION. EVALUATION OF 10 DEVICES FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICES WERE CLEANED OF DEBRIS AND THE TURBINES WERE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725518&pc=EFB | 9614977-2021-00079 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST PHOENIX | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K131319 |  | 4 | Event Description: THIS REPORT SUMMARIZES 4 MALFUNCTION EVENTS WHERE A MIDWEST PHOENIX HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 4 OF 4 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 4 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725519&pc=EFB | 9614977-2021-00075 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E PRO 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K972436 |  | 2 | Event Description: THIS REPORT SUMMARIZES 2 MALFUNCTION EVENTS WHERE A MIDWEST E PRO HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 2 OF 2 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 2 DEVICES FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICES WERE CLEANED OF DEBRIS AND THE TURBINES WERE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725520&pc=EFB | 9614977-2021-00078 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST STYLUS ATC MINI | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K003518 |  | 4 | Event Description: THIS REPORT SUMMARIZES 4 MALFUNCTION EVENTS WHERE A MIDWEST STYLUS ATC MINI HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: ADDITIONAL SERIAL NUMBERS INCLUDED IN THIS REPORT: (B)(4). 4 OF 4 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 4 DEVICES FOUND NORMAL CHUCK WEAR AND THE TURBINES NEEDED TO BE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725521&pc=EFB | 9614977-2021-00074 |  | Malfunction | SIRONA DENTAL SYSTEMS GMBH | 2021-10-30 | EFB | MIDWEST E MINI 1:5 HIGH SPEED CONTRA ANGLE ATTACHMENT | Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | K972436 |  | 4 | Event Description: THIS REPORT SUMMARIZES 4 MALFUNCTION EVENTS WHERE A MIDWEST E MINI HANDPIECE WOULD NOT HOLD BURS. NO INJURY RESULTED IN ANY OF THE EVENTS. Manufacturer Narrative: 4 OF 4 DEVICES WERE RETURNED FOR EVALUATION. EVALUATION OF 4 DEVICES FOUND CHUCK WEAR AND LACK OF MAINTENANCE. THE DEVICES WERE CLEANED OF DEBRIS AND THE TURBINES WERE REPLACED. THE DEVICES WERE REPAIRED AND RETURNED TO THE CUSTOMERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725522&pc=LKK | 3004209178-2021-16170 | 2021-08-13 | Injury | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2021-10-30 | LKK | SYNCHROMED II | Obstruction of Flow | Muscle Weakness; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (HEALTHCARE PROVIDER, CLINICAL STUDY) REGARDING A PATIENT WHO WAS RECEIVING COMPOUNDED BACLOFEN ( AT ), FENTANYL ( AT ), AND BUPIVACAINE ( AT ) VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT ON (B)(6) 2021 THE PATIENT CONTACTED THE CLINIC REPORTING WORSENING LOWER EXTREMITY WEAKNESS RESULTING IN TO FALLS. ON (B)(6) 2021THE CATHETER WAS EXPLANTED AND REPLACED BY AN OUTSIDE PROVIDER WHOSE OPERATIVE NOTE INDICATED INTRATHECAL (IT) OPIOID WITHDRAWAL AND CATHETER OBSTRUCTION. THE OUTCOME OF THE EVENT RESOLVED WITHOUT SEQUELAE ON (B)(6) 2021. THE DEVICE DIAGNOSIS WAS CATHETER OCCLUSION AND THE CLINICAL DIAGNOSIS WAS LOWER EXTREMITY WEAKNESS. THE ETIOLOGYOF THE EVENT INDICATED THE RELATIONSHIP OF THE EVENT TO THE DEVICE OR THERAPY WAS RELATED AND INDICATED THE RELATIONSHIP OF THE EVENT TO THE IMPLANT PROCEDURE WAS NOT RELATED. THE EVENT DATE WAS (B)(6) 2021. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: PRODUCT ID :8780, SERIAL# (B)(4), IMPLANTED: (B)(6)2016, EXPLANTED: (B)(6) 2021, PRODUCT TYPE:CATHETER. PRODUCT ID :8782, SERIAL# (B)(4), IMPLANTED: (B)(6)2017, EXPLANTED: (B)(6) 2021, PRODUCT TYPE: CATHETER. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 8780, SERIAL/LOT #: (B)(4), UBD: 16-AUG-2018, UDI#: (B)(4) ; PRODUCT ID: 8782, SERIAL/LOT #: (B)(4), UBD: 21-JAN-2018, UDI#: (B)(4). IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725523&pc=OZD | 1220648-2021-01115 | 2021-10-02 | Injury | ABIOMED INC. | 2021-10-30 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: AN IMPELLA CP WAS PLACED FOR HEMODYNAMIC SUPPORT FOR AN OBESE PATIENT (BMI OF 40) SUFFERING FROM CARDIOGENIC SHOCK. THE CP WAS PLACED VIA THE FEMORAL ARTERY. THE PATIENT DEVELOPED A HEMATOMA AT THE ACCESS SITE, AND SUCTION WAS ALARMING ON THE CONSOLE. THE TEAM TREATED THE BLEED WITH INFUSION OF RED BLOOD CELLS, APPROXIMATELY 8 UNITS. THE THOUGHT WAS TO REPLACE THE PUMP AND INSERT A LONGER SHEATH, THAT MAY HAVE HAD A BETTER ABILITY TO REACH THE ARTERIOTOMY AND REDUCE THE ACCESS SITE BLEED, BUT THE PATIENT EXPIRED. THE PATIENT'S OBESITY WAS THOUGHT TO BE THE CAUSE OF THE BLEED AS PER THE PHYSICIAN, AS THE TISSUE TRACK WAS TOO LONG TO ACHIEVE SUCCESSFUL HEMOSTASIS. Manufacturer Narrative: THE INVESTIGATION INTO THE ACCESS SITE BLEEDING IS ON-GOING AT THIS TIME. NO IMPELLA PRODUCT HAS BEEN RETURNED AND CLINICAL DETAILS ARE BEING CLARIFIED. UPON COMPLETION OF THE INVESTIGATION, A FINAL REPORT WILL BE FILED.
The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance | detect potential device-related safety issues | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers.
Although MDRs are a valuable source of information | this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | untimely | unverified | or biased data. In addition | the incidence or prevalence of an event cannot be determined from this reporting system alone due to potential under-reporting of events and lack of information about frequency of device use. Because of this | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
MAUDE data does not represent all known safety information for a reported medical device and should be interpreted in the context of other available information when making device-related or treatment decisions. Variations in trade | product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | importer | distributor | manufacturer | or medical personnel caused or contributed to the event. Certain types of report information are protected from public disclosure under the Freedom of Information Act (FOIA).
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.
Device user facilities include hospitals | outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA | if the manufacturer is not known.